**AMENDMENT No. 7 TO PROCUREMENT DOCUMENTS**

**Issued on** **13th of October 2020**

**FOR** **THE PROCUREMENT OF EQUIPMENT FOR MOTHER AND CHILD INSTITUTE „DR VUKAN ČUPIĆ”**

**NO. IOP/38-2019/UHI**

In accordance with the Clause 8. Part 1. Bidding Procedures, Section I. Instructions to Bidders, Contents of Bidding Documents, Amendment of Bidding Documents of the Procurement Documents, Public Investment Management Office, No. 11 Nemanjina street, Republic of Serbia, as the Purchaser, hereby notifies all persons concerned for Procurement of equipment for Mother and Child Institute „Dr Vukan Čupić”, No. IOP/38-2019/UHI, that there has been an amendment made in the Procurement Documents.

The following provisions of Tender Documents are hereby replaced as follows:

1. **Tender Documents are changed in Technical Specification, Lot 1 - Angio Operating Theatre, Line item No. 1.3.1 - Cardiovascular ultrasound machine, ID 2.15, which reads as follows:**

Ultrasound system can be upgraded with fusion imaging which enables display of the same anatomical section on ultrasound live image as well as on last exam image performed on CT or MR systems. Image Fusion functionality should be available on 2D transthoracic matrix cardiac probe

**Is replaced and reads as follows:**

Ultrasound system can be upgraded with fusion imaging which enables display of the same anatomical section on ultrasound live image as well as on last exam image performed on CT or MR systems

1. **Tender Documents are changed in Technical Specification, Lot 4 - Radiology, Line item No. 4.2 - Premium ultrasound machine for radiology department, ID 21, which reads as follows:**

Linear Transducer: Frequency range: min. 4 – 18 MHz±2MHz, number of elements min.280, min., FOV max.40mm+/-10mm

**Is replaced and reads as follows:**

Linear Transducer: Frequency range: min. 4 – 18 MHz±3MHz, number of elements min.280, min., FOV max.50mm+/-10mm

1. **Tender Documents are changed in Technical Specification, Lot 4 - Radiology, Line item No. 4.6 - PACS, ID 16, which reads as follows:**

at least 21", LED, Maximum luminance at least 800cd/m2, DICOM calibrated luminance of at least 500cd/m2, contrast ratio of at least 1400:1, Response time: maximum 10ms

**Is replaced and reads as follows:**

at least 21", LED, Maximum luminance at least 800cd/m2, DICOM calibrated luminance of at least 500cd/m2, contrast ratio of at least 1400:1, Response time: maximum 12ms

In all other aspects, Tender documents for the Procurement of equipment for Mother and Child Institute „Dr Vukan Čupić”, No. IOP/38-2019/UHI, remains unchanged.

This Amendment to Procurement Documents is a constituent part of the Procurement of equipment for Mother and Child Institute „Dr Vukan Čupić”, No. IOP/38-2019/UHI and will be posted on the Purchaser’s web site: <http://www.obnova.gov.rs/english/public-procurement>

Updated version of technical specifications is given in excel file and is constituent part of this amendment.

In order to submit a responsive bid, bidders are expected to prepare their offers in accordance with this amendment.

Public Procurement Committee