**PROCUREMENT OF EQUIPMENT FOR BIOSENSE INSTITUTE IN NOVI SAD, SERBIA**

**(PROCUREMENT NO. IOP/41-2019/RD)**

**CLARIFICATION NO. 14**

Issued on April 16, 2020

Regarding the list of questions that the Purchaser, Public Investment Management Office Belgrade, No. 11 Nemanjina street, have received from the potential bidders, concerning the procurement procedure: Procurement of Equipment for BioSense Institute in Novi Sad, Serbia no. IOP/41-2019/RD, we give you the following answers:

**Question 1:**

question for id number 7.18 Ultra low temperature freezer

the freezer that is specified in the tender documentation is not available anymore on market, the producer stopped production of this small model. because of that, is it allowed to offer freezer with next characteristics:

RELIABLE ULT CHEST FREEZER

Reaching -80°C in only 85 minutes

Temperature range (˚C) -40 / -86

Capacity (L) 71

Cooling technology single system / static Controller model / probe type G214 / PT100

Lock on handle for optimal security of the valued samples.

Refrigerant Nature R

Interior cabinet material painted steel Exterior cabinet material painted steel

Lock with key yes

Display type 4,3'' with touch buttons

Alarm display as text - not codes yes

Visual / acoustic alarm yes / yes

Power failure alarm yes

Adjustable high / low temperature alarm yes / yes

Open door alarm yes

Probe failure alarm yes

Low battery alarm yes

Compressor failure alarm yes

Battery backup for alarms (approx. hrs.) 72

**Answer 1:**

After carefully reviewing the technical specification listed, the Beneficiary concluded that the freezer with the listed characteristics meets technical requirements and therefore it is acceptable to offer.

**Question 2:**

* Position 7.2. ADVANCED STEREOMICROSCOPE line 11 EXTERNAL STABLE STAND

REQUESTED:

• External stable stand for using for bigger specimens

• Vertical Cross Travel: minimum 229 mm

• Horizontal Cross Travel: minimum 260 mm

QUESTION: Could be the EXTERNAL STABLE STAND offered with the following characteristics:

• External stable stand for using for bigger specimens

• Vertical Cross Travel: minimum 245mm

• Horizontal Cross Travel: minimum 272mm

* Position 7.2. ADVANCED STEREOMICROSCOPE line 13 POSSIBLE CONTRAST METHODS OR UPGRADES:

QUESTION: Is the Teaching head necessary as well

* Position 7.3. STEREOMICROSCOPE line 1 MAIN BODY

REQUESTED: Optical System: Parallel-optics type (zooming type)

QUESTION: - Is it possible to offer the stereomicroscope of the same characteristics with Greenough optics type (zooming type) instead of required

**Answer 2:**

* Position 7.2. ADVANCED STEREOMICROSCOPE line 11 EXTERNAL STABLE STAND

Since the offered stable stand has a bigger vertical and horizontal cross travel range, it meets technical requirements and therefore it is acceptable to offer.

* Position 7.2. ADVANCED STEREOMICROSCOPE line 13 POSSIBLE CONTRAST METHODS OR UPGRADES

Since this option falls into the “Possible contrast methods or upgrades” it is not mandatory, therefore it is not necessary to offer.

* Position 7.3. STEREOMICROSCOPE line 1 MAIN BODY

Yes, it is possible to offer the stereomicroscope with Greenough optic type as long as it meets specifications listed for the Zooming Bodies with the respect to minimum zoom ratio, zoom range, etc.

Public Procurement Committee