## Summary
Bidder: |  | Offered price:
ID | Equipment name | QTY | Manufacturer/Model | Total Price per line item
1.1. | Radiofrequency Ablation System | 2
1.2. | Atheroctomy System | 1
1.3. | Application software packages for "pace mapping", automatic CT segmentation and interpretation of tachycardia | 1
1.4. | Infusion Pump | 2
1.5. | Trolley for Instruments | 4
1.6. | Ultrasound cardiovascular premium 4D system of the latest generation with 4D TEE probe | 1
1.7. | Heart Lung Machine | 1
1.8. | Intensive care bed | 4
1.9. | Patient bed | 10
1.10. | Anasthesia Machine | 1
1.11. | Intensive Care Monitor | 4
1.12. | Contrast Injecton System | 1
Total estimated value

## 1.RF Ablation
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Radiofrequency Ablation System |  | 2 |  | 0
| 1 | An integrated radio frequency ablation system consisting of the following components: pump, generator, remote control system.
| 2 | The tube system must go through the sensor for bubble detection
| 3 | Active monitoring of the quantity of liquid in the bag with irrigated solution and alert notification when 100ml is left in the bag
| 4 | Maximum adjustable power, min. 100W,
| 5 | Maximum adjustable time, min. 999s
| 6 | Graphic display for monitoring of power, temperature, flow and impedance in real time
| 7 | Automatic flow adjustment, depending on the selected power
| 8 | Touchscreen
| 9 | Option to connect to the temperature probe in the esophagus

## 2.Atheroctomy
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Atheroctomy System |  | 1 |  | 0
| 1 | Console with minimum one peristaltic pump
| 2 | Rotational atherectomy for peripheral arteries
| 3 | Ability to be used below the knee
| 4 | Catheter length minimum 120 cm
| 5 | Compatible with 0.014” wire
| 6 | Mounting on standard IV pole

## 3.Software
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Application software packages for "pace mapping", automatic CT segmentation and interpretation of tachycardia |  | 1 |  | 0
| 1 | Application software package for "pace mapping" for adequate ECG localization during arythmia and during of mapping of heart chambers
| 2 | Application software package for automatic CT segmentation with:
| 2.1. | Segementation time not longer than 20 sec
| 2.2. | Automatic segmentation for all 4 cardiac chambers
| 2.3. | Visualization of papillary muscles and left ventricular trabeculation
| 2.4. | The possibility of measuring the wall thickness of the left ventricle
| 3 | Application software packages for interpretation of tachycardia
| 3.1. | Estimation of the path of tachycardia in real time
| 3.2. | Precise determination of the area of interest during catheter ablation of arrhythmias.

## 4.Inf. Pump
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Infusion Pump |  | 2 |  | 0
| 1 | Volumetric infusion pump (linear, peristatic)
| 2 | Intended use: receiving IV citostatic, nutritients
| 3 | High precision ±5%
| 4 | Flow rate min 0,1-1200ml/h, possibility to change flow without infusion interrupt. Automatic calculation of speed according to volume and infusion time.
| 5 | Infusion volume 0,1-9999 , step 0,01 ml/h or 0,1 ml/h
| 6 | Infusion time 1 min - 99h 59 min.
| 7 | Dose units: g, mg, ug, ng, IU, hour, min., kg
| 8 | Manual and automatic bolus
| 9 | Bolus flow min. 0,2-1200 ml/h
| 10 | Drop sensor
| 11 | Air bubble detector: min. 50-800µl
| 12 | KVO rate 0.1-5.0 ml/h
| 13 | Possibility to set oclusion pressure in min.10 steps.
| 14 | Automatic pressure lowering during occlusion
| 15 | Displaying pressure in IV line
| 16 | Data show on LCD display. Displaying speed, volume and infusion time
| 17 | Possibility to show additional informations on display: meny for programming functions and alarms, additional parameters and settings (pressure limit, battery status, bolus dose calculation etc.)
| 18 | Using standard IV sets from different manufacturers
| 19 | Working modes: Flow rate, Volume+Time
| 20 | Possibility to work on IV stand or stand-alone
| 21 | Possibility to connect to docking station. Connectors: USB or RS232, and nurse call
| 22 | Lock function
| 23 | Memory for more then 1000 events
| 24 | Alarms: Visual alarm, red colored, visible from large distance + alarm message on display, sound alarm with possibility to adjust sound volume.
| 25 | Powered by AC source and battery. Battery work min 6 hours with adjusted flow 25ml/h.
| 26 | Weight up to 2,5 kg with battery and charging system.

## 5.Trolley
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Trolley |  | 4 |  | 0
| 1 | Stainless steel tubular frame
| 2 | 2 stainless steel shelves with raised edges.
| 3 | Swivelling castors Ø 125 mm approximatively, lockable;
| 4 | PVC corner bumpers
| 5 | External dimmensions: min. 800x500x950h mm +/-5%
| 6 | Permitted load of 30 kg per shelf

## 6.US
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Ultrasound cardiovascular premium 4D system of the latest generation with 4D TEE probe |  | 1 |  | 0
| 1 | Digital color doppler 4D ultrasound system, not portable, with wheels. Front and rear handles for easy transport.
| 2 | Advanced beamforming technology, software based, C-Sound or similar with dynamic range minimum 270Db.
| 3 | Frame rate more than 1000fps
| 4 | Four active probe ports, not including pencil ports.
| 5 | 22" wide screen High-Definition (HD) OLED display with monitor on an articulating arm. Resolution 1920x1080 px
| 6 | Integrated Ultra HD LCD touch-screen panel, diagonal at least 12”.
| 7 | Floating operating panel, with integrated alphanumeric backlight keyboard, adjustable in three dimensions: Height, Rotation, Extension.
| 8 | Integrated HDD
| 9 | Maximum depth range in B-mode not less then 30 cm – probe specific
| 10 | Width range in 2D mode 10-120°
| 11 | Transducer Types:
| 11.1. | Convex array probes
| 11.2. | Micro convex array probes
| 11.3. | Linear array probes
| 11.4. | Sector phased array probes
| 11.5. | Sector phased array TEE probes
| 11.6. | Sector matrix phased array probes
| 11.7. | 4D transthoracic matrix phased array
| 11.8. | 2D TEE probes
| 11.9. | 4D TEE volume matrix array probes
| 12 | Operating mode
| 12.1 | B-Mode ( 2D mode )
| 12.2 | M-Mode
| 12.3 | Anatomical M-mode
| 12.4 | Curved anatomical M-mode
| 12.5 | Color Flow Mode (CFM)
| 12.6 | 2D Virtual Apex or similar
| 12.7 | Color Angio Power Doppler Imaging (PDI)
| 12.8 | PW Doppler with High PRF
| 12.9 | TVI, TDI, Tissue Tracking
| 12.1 | CWD
| 12.11 | 4D Tissue
| 12.12 | 4D Color flow
| 13 | Stress echo
| 14 | An advanced image processing technique to remove speckle in real-time examining. As UD clarity and UD speckle reduce imaging or similar.
| 15 | Multiple-angle compound imaging in real time which reduce angular dependence of border or edge as compared to no-compound imaging as Compound Imaging, Real time sono CT, SieClear or similar.
| 16 | Automatic Optimization of B-mode image and spectral image with one key touch as TEQ, Iscan, AO or similar
| 17 | Automatic Spectrum Optimization – provides a single push, automatic, real-time optimization of PW or CW spectrum scale and baseline display. ASO or similar
| 18 | CW Doppler Triplex mode (simultaneous display of 2D, Color Flow and CW Doppler image in real time )
| 19 | Visualization providing enhanced display of anatomical structures using advanced shadowing techniques in combination with depth colormaps as HD live or similar
| 20 | Software which, depending on the given parameters, allows a volume rate of min. 1500 on 4D TEE probes in single beat acquisition, without reducing image quality Vmax or similar.
| 21 | Simultaneous visualization of Bi-plane acquisition and Tri-plane acquisition. Bi-plane, Tri-plane or similar.
| 22 | Simultaneous display of three independent random slices through the 4D volumen (tissue and color) as FlexiSlice or similar(Qlab?
| 23 | Simultaneus display of 5,7,9 or 12 combined of short-axis and long-axis standard slices extracted from the 4D volume data (tissue and/or color) available in live and replay as Multi-slice or similar.
| 24 | Transducers:
| 24.1 | 2D Sector matrix cardiac probe ( matrix technology provides probe elements in several rows in matrix order) Frequency range from 1.4 to 4.6 MHz (+/- 0,5 MHz).Depth of field: minimum 30 cm. Sector width: minimum120 °
| 24.2 | 4D Volume Transesophageal (TEE) matrix array probe transducer for real-time three-dimensional imaging (matrix technology provides probe elements in severel rows in matrix order). Frequency range from 3 MHz to 8 MHz (+/- 0,5 MHz) or wider
| 24.3 | Pencil Probe
| 25 | Automated EF measurement tool based on 2D-speckle tracking algorithm.
| 26 | ECG cable
| 27 | Thermal B&W printer

## 7.Heart Lung
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Heart Lung Machine |  | 1 |  | 0
| 1. | Mobile console - stand
The console should be mobile, transparent and flexible in terms of orientation of other system elements according to user requirements and should have:
| 1.1. | Mobile base with at least 4 roler pumps with an electronics housing
| 1.2. | Possibility of integrating of the centrifugal pump
| 1.3. | Power supply elements with cables and battery
| 1.4. | Possibility of automatic switching of the machine to the battery / accumulator power
| 1.5. | Not less than two infusion sets with telescopic height adjustment
| 1.6. | Shelf
| 1.7. | LED lamp with flexible holder
| 1.8. | Front or side access to all connections, cables and battery
| 1.9. | Other elements necessary for the functioning of the machine
| 2. | Roller pumps
The machine must contain at least three large roller pumps (min 150 mm) and two additional for delivery of cardioplegic solution which may be of smaller diameter, but with independent adjustment and control. All roller pumps should have transparent displays that show their condition and actions.Each pump should have:
| 2.1. | Possibility of rotating the head by 180 degrees
| 2.2. | Manual actuation lever in the event of a standstill
| 2.3. | Adapters for different measuring lines (3/8 ", 1/4", 3/16 ")
| 2.4. | Possibility of rotation in both directions
| 2.5. | Continuous and pulse mode
| 2.6. | Possibility of automatic stopping by alarm
| 2.7. | Possibility of stopping the pump operation when opening the roller pump cover
| 2.8. | Ability of follow up the operation of the main pump by the other pumps
| 2.9. | Other elements necessary for the functioning of the machine
| 3. | Integrated centrifugal pump
The centrifugal pump should be an integral part of the machine. It should have a console that is transparent and easy to operate and a precise sensor for flow measurement. It consists of:
| 3.1. | Integrated console
| 3.2. | Driving unit
| 3.3. | Flow sensor
| 3.4. | Manual actuators
| 3.5. | Centrifugal pumps
| 3.6. | Other elements necessary for the functioning of the pump
| 4. | Control unit for control of system elements
The control unit should provide easy control of the system elements. It is should be compact and transparent .It should provide an overview of the following parameters:
| 4.1. | Minimum three time durations
| 4.2. | Minimum two pressures with the option for upgrade of the module for displaying two more pressures
| 4.3. | Minimum two temperatures
| 4.4. | Control of cardioplegic parameters
| 4.5. | Blood level in the oxygenator
| 4.6. | Presence of gas bubbles in the lines
| 4.7. | Visual and audible alarm
| 5. | Pressure measuring module
It must provide precise measurement of at least two pressures, with easy calibration and limiting of individual displayed quantities ​​which are connected with the alarm system in terms of controlling of pump flow according to the given limits. Module contains:
| 5.1. | Pressure measurement module
| 5.2. | Reusable transducers
| 5.3. | Transducer holders
| 5.4. | Possibility of regulating the pressure by changing the pump speed
| 5.5. | Possibility of stopping the pump depending on the preset pressure limit
| 5.6. | Cables and other elements necessary for the functioning of the module
| 6. | Temperature measurement module with temperature probes
The module should provide precise measurement and individual display of up to four temperatures ​​with the possibility of limiting and alarming of the limit values. It should include:
| 6.1. | Temperature measurement module
| 6.2. | Probe for measuring of temperature in esophagus (YSI 400)
| 6.3. | Four probes for measuring the temperature in the oxigenator (YSI 400)
| 6.4. | Minimum 4 temperature channels
| 7. | Module for measuring time duration
The module must include the ability to measure at least three time durations.
| 8. | Module for protection against air embolism
This module includes two models of protection: blood level control in the oxygenator and the presence of gas bubbles in the system lines. Both systems must be able to control the flow of the main roller pump according to the indication obtained from the sensor. Module contains:
| 8.1. | Blood level measurement module in an oxygenator with an audio-visual alarm
| 8.2. | Possibility of stopping the pump
| 8.3. | Module for indication of gas bubbles in lines with audio-visual alarm
| 8.4. | Possibility of at least two types of bubble registration
| 8.5. | Ability to detect bubbles
| 8.6. | Level sensor and other elements necessary for the functioning of the module
| 9. | Gas mixer
The gas mixer should provide precise mixing of oxygen and air and adequate delivery of such formed gas mixture with a concentration range of oxygen at least from 21% to 100%.. It consists of:
| 9.1. | Gas mixer with two tubes for oxygen and air
| 9.2. | The holder
| 9.3. | Cables, connectors and other elements for the necessary functioning
| 10 | Heater-cooler unit
It should enable rapid, precise, independent and effective control of the patient's temperature and cardioplegia solution. Device should be reliable and easy to handle. It consists of:
| 10.1. | Cooling and heating device
| 10.2. | Minimum 2 water tanks
| 10.3. | Minimum 2 separate water circulation systems with separate pumps
| 10.4. | Minimum 2 separate pairs of hose connections (Hansen connectors)
| 10.5. | Thermometer with display of current and given water temperature
| 10.6. | Mattress for heating or cooling the patient
| 10.7. | Control unit detachable from the body of the tank with the possibility of positioning it in suitable locations for the operator, including the possibility of mounting on the machine
| 10.8. | Color LCD touch monitor, within the control panel
| 10.9. | Temperature adjustment, at least 0.10C
| 11. | Module for cardioplegia
It should provide precise delivery of cardioplegic solution and consists of:
| 11.1. | Time and pressure display module and the possibility of bubble detection in the lines
| 11.2. | Possibility of controlling at least two roller pumps
| 11.3. | Audio-visual alarm

## 8.ICU bed
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Intensive care bed |  | 4 |  | 0
| 1 | ICU bed with minimum 5 motors
| 2 | Total bed dimensions 1050mm ±20mm x 2200mm ±40mm
| 3 | Safe work load (SWL) min.260kg
| 5 | Movable sides of the head and front. Head side optionally possible take off or put in a horizontal position.
| 6 | Cassette carrier for X-ray beneath the backrest
| 7 | Lying plate: made of segments that are easily removed for cleaning and desinfection
| 8 | Infusion stand with four hooks made of stainless steel or iron, the carrier for oxygen bottle made of stainless steel or in the basis of the bed.
| 9 | Bed frame is made of steel profiles painted with powder-colored polyester resistant to desinfectant
| 10 | Lying plate dimensions: 870 mm x 2000 mm ±20mm
| 11 | The total length of the plate with the including extension is 2300 mm ±20mm. Extension of plate is integrated in the bed frame
| 12 | Castors are duble, minimum Ø150mm. The foot brake position blocks all 4 castors or blocking the rotation of one castor
| 13 | Regulation of height with electric motor in min. range: 460mm - 800mm
| 14 | Electrical automatic Tr ≥ 15 °/ ATr ≥ 19 °
| 15 | Tilt angle of back rest ≥ 70 °.
| 16 | Tilt angle of the thigh rest ≥ 60 °
| 17 | Possibility of lateral movement of the lying plate ≥ 25 °
| 18 | Autoregression of the back rest part
| 19 | Quick positioning function of the patient in a shock position
| 20 | Function of cardio chair
| 21 | Nursing control panel for all bed functions (up-down, autocontour of the back part, Trendelenburg, Reverse Trendelenburg, lateral) and with special buttons for electrical CRP, Trendelenburg and battery charge indicator.
| 22 | Hand control unit is portable and it is possible to hang it on the sides of the bed
| 23 | Auto-stop automatic stop function when hovering downboard on the obstacle from the bottom
| 24 | Easily visible and approachable manual CPR commands
| 25 | Possibility of manually lowering the backrest and upper legs
| 25 | Movable split safety sides made of metal, with rounded edges made of technoploymer
| 26 | Battery for starting all electrical functions when the bed is disconnected from the mains supply
| 27 | Alarm connected to the braking system. It is activated when the brakes are off and the cable plugged into the socket
| 28 | Rails for mounting accessories on the sides, below the laying plate
| 29 | Rubber bumpers at the corners of the bed
| 30 | Decontamination degree of protection IPX6W
| 31 | Anti-Decubitus mattress with imprecable cover, which is in accordance with the bed dimmensions .

## 9.Bed
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Bed |  | 10 |  | 0
| 1 | Bed with minimum 4 motors
| 2 | Total bed dimensions 1000mm ±20mm x 2180mm ±40mm
| 3 | Safe work load (SWL) min. 250kg
| 4 | Movable sides of the head and front with security lock.
| 5 | Lying plate: made of minimum 4 segments that are easily removed for cleaning and desinfection
| 6 | Infusion stand with four hooks made of stainless steel or iron, the carrier for oxygen bottle made of stainless steel or in the basis of the bed.
| 7 | Bed frame is made of steel profiles painted with powder-colored polyester resistant to desinfectant
| 8 | Lying plate dimensions: 870 mm x 2000 mm ±30mm
| 9 | The total length of the plate with the including extension is 2300 mm ±20mm
| 10 | Castors are minimum Ø150mm. The foot brake position blocks all 4 castors or blocking the rotation
| 11 | Regulation of height with electric motor in min. range: 380mm - 800mm
| 12 | Electrical automatic Tr ≥ 16 °/ ATr ≥ 16 °
| 13 | Tilt angle of backrest ≥ 70 °.
| 14 | Tilt angle of the thighrest ≥ 42 °
| 15 | Autoregression of the back part
| 16 | Quick positioning function of the patient in a shock position
| 17 | Function of cardio chair
| 18 | Hand control unit is portable and it is possible to hold on the sides of the bed
| 19 | Easily visible and approachable manual CPR commands
| 20 | One-piece side pannels covering at least ¾ of the lying slabs, with a locking system in the top position
| 21 | Battery for starting all electrical functions when the bed is disconnected from the mains supply
| 22 | Alarm connected to the braking system. It is activated when the brakes are off and the cable plugged into the socket
| 23 | Rubber bumpers at the corners of the bed
| 24 | Nursing control panel for all bed functions (up-down, autocontour of the back part, Trendelenburg, Reverse Trendelenburg, lateral) and with special buttons for electrical CRP, Trendelenburg and battery charge indicator.
| 25 | Anti-Decubitus mattress with imprecable cover, which is in accordance with the bed dimmensions .

## 10.Anesthesia
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Anasthesia Machine |  | 1 |  | 0
| 1 | The device is with its own mobile platform with the wheels that can be locked
| 2 | For working with adults and children
| 3 | Circular system of anesthesia
| 4 | Possibility of performing anesthesia of low, i/e/ minimal flow
| 5 | Aspirator
| 6 | Auxiliary module for oxygen
| 7 | Regulated gas surplus outlet
| 8 | Autonomous operation (no power supply) - battery power over 30 minutes
| 9 | Outlet of excess anesthetic gases
| 10 | Electronic control of the supply of fresh gases (O2, N2O, air)
| 11 | Electronic regulation of the concentration of O2 in the mixture of O2-N2O and O2-air
| 12 | Possibility of instant delivery of O2 high flow
| 13 | Possibility of delivering O2 for emergency ventilation using a special system
| 14 | Display a mixture of gases and their concentration on the screen
| 15 | System for electronically driven and controlled direct injection of volatile anesthetics (evaporator)
| 16 | Possibility of installing 2 systems for electronically driven and controlled direct injection of volatile anesthetics (evaporator)
| 17 | System for electronically driven and controlled direct injection of volatile anesthetics (evaporator) for sevoflurane
| 18 | Possibility of manual and automatic (mechanical ventilation)
| 19 | Manual ventilation system has an APL valve
| 20 | Ventilator: electronically or pneumatically driven
| 21 | Complementary with absorber, reusable canister for CO2 absorber for a volume of 0.7l or greater.
| 22 | Mechanical ventilation modes: controlled by volume, pressure controlled, supported by pressure, synchronized regimes controlled by pressure and volume
| 23 | Synchronization - trigger by pressure or by flow or by both
| 24 | Fresh gas flow, 0.3 l - 15 l / min or greater range
| 25 | Breathing volume: 50 to 1500 ml or more
| 26 | Breathing frequency (mechanical ventilation): 3-80 preminute or a larger range
| 27 | Inspiration/expiration ratio: 2: 1 - 1: 4 or greater range
| 28 | Possibility of using PEEP
| 29 | Integrated pressure monitoring, breathing volume (during inhalation and exhalation), display of pressure-volume loop and flow-volume loop
| 30 | Possibility of computing the compliance (static or dynamic, or both of them)
| 31 | Integrated gas analyzer with display (on the monitor): concentrations (partial pressure) of CO2 during the breathing cycle - capnometry with capnography; concentration of O2 in the mixture of gases during inhalation and exhalation; concentration of N2O in the mixture of gases during inhalation and exhalation
| 32 | Calculation and percentage of MAC
| 33 | Automated software system for anesthesia of low or minimum flow
| 34 | Automatically identification of the gases in the mixture
| 35 | Automated complete check up by turning device on.
| 36 | Automated test of compliance and leakage
| 37 | Automated test system for electronically driven and controlled direct injection of volatile anesthetics (evaporator)
| 38 | Graphic and numerical trends (possibility of storing and exporting data)
| 39 | Alarms: of minute volume, of leakage in the system, of high pressure in the respiratory system, of high permanent pressure in the manual mode, of PEEP, of breathing frequency, of apnea, of inhalation and exhaustion O2, CO2 and of anesthetic gases, technical malfunction
| 40 | Different levels of alarm (priority)
| 41 | Necessary accessories which are delivered and provided with the device: set of hoses, balloon, hose fixing holder, aspiration system, hoses for connecting the device to the central gas supply system (for O2, air and N2O, length of min 5)

## 11.Monitor
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Intensive Care Monitor |  | 4 |  | 0
| 1 | Stationary modular monitor for patients of all ages ( adult, children, neonates ). Modules are inserted directly to monitor housing as plug-and-play modules.
| 2 | Monitor has color display, diameter 19 inches, with high resolution ( 1280x1024 ) touchscreen
| 3 | At least 12 curves can be displayed on screen
| 4 | Adjustable audio/visual alarm system for all parameters
| 5 | Trend memory: at least 48 hours of graphical/tabular trends
| 6 | Possibility for networking ( LAN, WLAN ) and connecting to central monitoring system.
| 7 | Integrated USB ports, minimum 4 for connecting keyboard, mouse, printer, bar-code reader or other peripherals.
| 8 | Monitor can provide following measurements
- 3/5 ECG
- ST segmenth analysis
- Continuous QT/QTc
- HR
- Detection of minimum 20 types of arrhythmia
- Arrhythmia analysis
- Respiration
- NIBP – non-invasive blood pressure
- SpO2 - saturation
- Temperature
- IBP – invasive blood pressure – 2 channels
- EtCO2 - capnography
| 9 | ST templates can be memorized and shown on screen for comparing with real time ST segments.
| 10 | Displaying perfusion index (PI)
| 11 | Monitor should be upgradable with the following options:
| 12 | С.О. – Cardiac Output
| 13 | Hemodynamic parameters - PICCO
| 14 | Measuring ScvO2
| 15 | Accessories which should be delivered togetherwith monitor
| 16 | ЕCG cable 5-lead ( 1 pc )
| 17 | Tube and cuff for blood pressure (1 pc )
| 18 | Reusable cable and sensor for SpО2, (1 pc )
| 19 | Temperature probe (1 pc )
| 20 | Accessories for invasive blood pressure (2 cables and 4 transducers)
| 21 | Accessories for capnography

## 12.Injector
Bidder:
LOT 1 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Contrast Injection System |  | 1 |  | 0
| 1 | High pressure injector for contrast media
| 2 | Integrated transducer or possibility to connect to transducer
| 3 | Syringe volume minimum 100 ml
| 4 | Possibility of injecting contrast media and saline