## Price Schedule
Bidder: |  |  |  | Offered price (currency)
ID | Equipment name | QTY | Technical Specification Offered(model) | Unit price | Total price per line item
1.1 | Fiberoptic bronchoscope | 1 |  |  | 0
1.2 | Patient monitor 1 | 24 |  |  | 0
1.3 | Patient monitor 2 | 24 |  |  | 0
1.4 | Central monitoring station for patient vital signs monitors | 4 |  |  | 0
1.5 | Operating lamp | 1 |  |  | 0
1.6 | Mobile operating lamp | 1 |  |  | 0
1.7 | Ultrasound 1 | 1 |  |  | 0
1.8 | PREMIUM ULTRASOUND COLOR DOPPLER DEVICE | 1 |  |  | 0
1.9 | Defibrillator | 1 |  |  | 0
1.1 | Laparoscopic instruments set | 1 |  |  | 0
1.11 | Video gastroscope | 1 |  |  | 0
1.12 | Intraaortic balloon Pump | 1 |  |  | 0
1.13 | Mobile radiography digital system with wireless FPD (Flat Panel Detector) | 1 |  |  | 0
1.14 | Endoscopic video tower with duodenoscope | 1 |  |  | 0
1.15 | Rigid Macintosh Laryngoscope Set | 3 |  |  | 0
1.16 | Laryngoscope set for a difficult intubation | 3 |  |  | 0
0

## Specifications
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 1.1 | Fiberoptic bronchoscope |  | 1 |  | 0
| 1 | Field of view 120°,
| 2 | Depth of view 3-50 mm,
| 3 | Insertion tube outer diameter 6.0 mm,
| 4 | Distal end outer diameter 5.9 mm,
| 5 | Instrument channel inner diameter 3.0 mm,
| 6 | Working length 600 mm,
| 7 | Bending up/down 180°/130°
| 8 | Set of instruments consisting of:
| 9 | Single-Use Alligator-Cup Biopsy Forceps
| 10 | Ribbed sheat to aid insertion; Fenestrated cups for larger biopsies; Unique SwingJaw mechanism to aid tangential biopsiesAlligator Jaw to prevent slippage; cups with superior cutting performance for fewer artifacts and less bleeding; min.working channel 2,0 mm; working length 1150mm
| 11 | Single-Use Oval-Cup Biopsy Forceps
| 12 | Ribbed sheat to aid insertion; Fenestrated cups for larger biopsies; Unique SwingJaw mechanism to aid tangential biopsies; cups with superior cutting performance for fewer artifacts and less bleeding; min.working channel 2,0 mm; working length 1150mm
| 13 | Reusable Round-Cup Biopsy Forceps
| 14 | For routine biopsies and transbronchial lung biopsies; Fenestrated cups for larger biopsies; Minimal tissue damage; Rotatable version for accurate positioning ; min.working channel 2,0 mm; working length 1050mm, rotatable
| 15 | Reusable Round-Cup Biopsy Forceps
| 16 | For routine biopsies and transbronchial lung biopsies; Fenestrated cups for larger biopsies; Minimal tissue damage; min.working channel 2,6 mm; working length 1050mm
| 17 | Reusable Round-Cup Biopsy Forceps
| 18 | For routine biopsies and transbronchial lung biopsies; Fenestrated cups for larger biopsies; Minimal tissue damage; min.working channel 2,8 mm; working length 1050mm
| 19 | -Reusable Grasping Baskets
| 20 | For retrieval of smooth, round objects, to be used with FG Handle MA-479
| 21 | -Reusable Grasping Baskets handle
| 22 | -Reusable Alligator-Jaw Grasping Forceps
| 23 | For holding or removing objects; Aligator jaws prevent slipping; min.working channel 2,8 mm; working length 1650mm, opening with 7,5 mm
2 | 1.2 | Patient monitor 1 |  | 24 |  | 0
| 1 | TFT LCD color display, min. 10", touchscreen
| 2 | Simultaneous display of min. 5 waveforms and 10 numeric parameters
| 3 | Connections for monitoring of ECG, Respiration, SpO2, NIBP, Temperature
| 4 | Connection for monitoring of IBP
| 5 | Connection for intubated and nonintubated capnography
| 6 | Pacemaker signal detection and rejection
| 7 | Automatic ST segment level detection, alarming and storing
| 8 | Automatic arrhythmia detection, alarming and storing: min. 10s per episode, total min. 72h
| 9 | Graphic and numerical parameters trend, min. 72h
| 10 | ECG full disclosure min. 72h
| 11 | Visual and sound alarm, 3 priority levels, adjustable limits, storage for min. 72h
| 12 | Continuous noninvasive or minimally invasive monitoring of estimated hemodynamic parameters CO and SV
| 13 | ECG sync sound with pitch variable according to current SpO2 or IBP value
| 14 | Choice for enlarged numerical display mode, readable from long distance
| 15 | Power supply AC 220V/50Hz with battery backup min. 3h
| 16 | Upgradeable with integrated thermal recorder, min. 3 channels
| 17 | Defibrillation proof patient connectors
| 18 | Holder for fixation to patient bed and medical rail
| 19 | Upgradeable for connection to central monitoring system, interbed monitoring and wireless telemetry
| 20 | Onscreen operation guide with illustrations
| 21 | Maximal weight: 5 kg
| 22 | To be delivered with all accessories needed, including:
| 23 | NIBP cuff, 2 pcs
| 24 | ECG cable with 3 leads, 1 pc
| 25 | Reusable waterproof SpO2 probe, 1 pc
| 26 | Temperature probe, 1 pc
| 27 | Capnography sensor with 50 airway adapters, 1 set
| 28 | Accessories for continous noninvasive or minimally invasive monitoring of CO and SV, for 100 patients
| 29 | Battery, 1 pc
3 | 1.3 | Patient monitor 2 |  | 24 |  | 0
| 1 | TFT LCD color display, min. 15", touchscreen
| 2 | Connection for optional secondary display of any size
| 3 | Simultaneous display of min. 12 waveforms and corresponding numeric parameters
| 4 | Connections for monitoring of ECG, Respiration, SpO2, NIBP, Temperature
| 5 | Upgradeable for 12-ch ECG monitoring and ECG interpretation with min. 100 findings
| 6 | Connection for monitoring of three simultaneous IBP
| 7 | Expandable with min. 4 additional simultaneous IBP connections
| 8 | Connection for intubated and non-intubated capnography and End-tidal CVP
| 9 | Additional NIBP measurements triggered by sudden changes in system pressure or pulse wave propagation time or any similar technique
| 10 | Pacemaker signal detection and rejection
| 11 | Automatic ST segment level detection, alarming and storing for min. 72h
| 12 | Automatic detection and alarming for min. 20 arrhythmia types with storing
| 13 | Graphic and numerical parameters trend, min. 72h
| 14 | Short OCRG trend on display
| 29 | ECG full disclosure min. 72h
| 30 | Visual and sound alarm, 3 priority levels, adjustable limits, storage for min. 72h
| 33 | Alarm lamp on top of the unit, with three colors, visible from all directions
| 34 | Continuous noninvasive monitoring of estimated or minimally invasive hemodynamic parameters CO and SV
| 35 | Software for hemodynamic management and fluid menagement
| 36 | ECG sync sound with pitch variable according to current SpO2 or IBP value
| 37 | Choice for enlarged numerical display mode, readable from long distance
| 38 | Power supply AC 220V/50Hz with battery backup min. 1h
| 39 | Defibrillation proof patient connectors
| 40 | Active connection to central monitoring system and interbed monitoring
| 41 | Upgradeable for wireless telemetry
| 42 | Upgradeable to sidestream monitoring of CO2, N2O, O2 and anesthetic agents with automatic ID detection, with calculation of MAC value and with simultaneous spirometry
| 43 | Upgradeable with min. 4-ch EEG, BIS, TOF, Flow/Paw and thermodilution CO monitoring
| 44 | Upgradeable with integrated 3-ch thermal recorder
| 45 | To be delivered with all accessories needed, including:
| 46 | NIBP cuff, 2 pcs
| 47 | ECG cable with 3 leads, 1 pc
| 48 | Reusable waterproof SpO2 probe, 1 pc
| 49 | Accessories for continous noninvasive or minimally invasive monitoring of CO and SV, for 100 patients
| 50 | Battery, 1 pc
| 51 | Holder for fixation to medical rail, 1 pc
4 | 1.4 | Central monitoring station for patient vital signs monitors |  | 4 |  | 0
| 1 | PC based central station for simultaneous monitoring of vital signs from min. 16 patients
| 2 | Two TFT LCD color displays, min 22" each
| 3 | Monitoring of waveforms and numerical parameters and storing for all beds
| 4 | Upgradeable for up to min. 48 monitors
| 5 | Capable to display min 12 simultaneous waveforms for selected patient bed
| 6 | Numeric and graphical trend min. 120h, for each patient
| 7 | Memory for storing min. 750 arrhythmia episodes and min 7200 ST episodes
| 8 | Memory for storing min. 120h of full disclosure, min. 8 waveforms per bed
| 9 | Displaying of 12-ch ECG for selected patient, if available at bedside monitor
| 10 | Displaying of patient graphic and visual alarms, three priority levels
| 11 | Adjustment of alarm limits
| 12 | Memory for storing min. 1000 alarm episodes per bed
| 13 | Upgradeable for 12 lead ECG interpretation and 3-ch thermal recorder
| 14 | Dedicated insulation units (transformers) for each patient monitor
| 15 | Patient reports printed on laser printer A4
| 16 | Unconditional power supply (UPS)
| 17 | Communication via LAN, network is provided by hospital
| 18 | Must be compatible with patient monitors offered in other lots
5 | 1.5 | Operating lamp |  | 1 |  | 0
| 1 | Ceiling anchorage with two arms, hard construction, diameter 110mm or more;
| 2 | Two arms with the spring and two cardanic arms for suspension of the lamp head and the satellite;
| 3 | Rotatable main lamp and the satellite along vertical axis in circle of 360° without the blocker;
| 4 | Technical features for the compliance of the main lamp:
| 5 | LED light technology with provision to change hues of the white color;
| 6 | Shape of the lamp head without sharp edges, easy to wipe and maintain;
| 7 | Light intensity at the distance of 1m 160 000 Lux;
| 8 | Centrally positioned light segment improving the depth illumination;
| 9 | Illumination depth 1500mm or more;
| 10 | Color rendering index Ra≥ 95;
| 11 | Light color temperature changeable between 3750K and 4750K;
| 12 | Light field diameter adjustable in the range from 20 to 32cm and wider;
| 13 | Provision for individual replacement of the light elements;
| 14 | Optical system for increasing contrast and reduction of shadows in the light field or equivalent
| 29 | Light mode for endoscopic procedures;
| 30 | Light head diameter ≤750mm;
| 33 | Lifetime of LEDs ≥50 000 work hours;
| 34 | Light intensity control in the range of 5-100% and wider;
| 35 | Temperature increase in the area around the surgeon’s head ≤ 1 °C;
| 36 | Total power consumption of the lamp head ≤80W;
| 37 | Additional sterilizing hand;
| 38 | Two hand grips on the side of the lamp head for easier handling.
| 39 | Technical features for the compliance of the satellite lamp:
| 40 | LED light technology with provision to change hues of the white color;
| 41 | Shape of the lamp head without sharp edges, easy to wipe and maintain;
| 42 | Light intensity at the distance of 1m 140 000 Lux;
| 43 | Centrally positioned light segment improving the depth illumination;
| 44 | Illumination depth 1200mm or more;
| 45 | Color rendering index Ra≥ 95;
| 46 | Light color temperature changeable between 3750K and 4750K;
| 47 | Light field diameter adjustable in the range from 18 to 28cm and wider;
| 48 | Provision for individual replacement of the light elements;
| 49 | Optical system for increasing contrast and reduction of shadows in the light field or equivalent
| 50 | Light mode for endoscopic procedures;
| 51 | Light head diameter ≤600mm;
| 52 | Lifetime of LEDs ≥50 000 work hours;
| 53 | Light intensity control in the range of 5-100%;
| 54 | Temperature increase in the area around the surgeon’s head ≤ 1 °C;
| 55 | Total power consumption of the lamp head ≤60W;
| 56 | Additional sterilizing hand;
| 57 | Two hand grips on the side of the lamp head for easier handling.
6 | 1.6 | Mobile operating lamp |  | 1 |  | 0
| 1 | Mobile LED lamp with focus adjustment;
| 2 | Mobile stand with four castors, at least two castors with the brake;
| 3 | Adjustable height of the lamp on the stand in the range of 1080mm and higher;
| 4 | Single color LED light technology;
| 5 | Light intensity at the distance of 1m 130.000 Lux or more;
| 6 | Illumination depth d20 1500mm or more;
| 7 | Color rendering index Ra≥ 95;
| 8 | Light color temperature fixed 4300K
| 9 | Light field diameter adjustable in the range from 17 to 22cm and wider;
| 10 | Provision for replacement of individual light segments;
| 11 | Shadows control
| 12 | Lamp head diameter ≤400mm;
| 13 | Lifetime of LEDs ≥50 000 work hours;
| 14 | Light intensity control in the range of 50-100%;
| 29 | Temperature increase in the area near the surgeon’s head ≤ 1 °C;
| 30 | Total power consumption of the lamp head ≤35W
7 | 1.7 | Ultrasound 1 |  | 1 |  | 0
| 1 | Digital color doppler ultrasound system with advanced signal processing technology
| 2 | High resolution LCD/LED monitor, screen size of min 21”, on articulating arm
| 3 | 4 wheels for system easy manevaurability
| 4 | Minimum 5 active transducer ports plus pencil probe connector. Each offered probe can be connected on each transducer’s port
| 5 | Display depth min 30 cm
| 6 | Touch screen diagonal size min 13”
| 7 | Min 4 USB ports
| 8 | Archiving on integrated SSD (min 500 GB), DVD and USB
| 9 | QWERTY keyboard with backlight functionality
| 10 | THI (Tissue Harmonic Imaging) technology with selection of min 2 frequencies
| 11 | CINE memory with minimum 1200 frames or 500 MB
| 12 | Operating modes: 2D mode, Dual mode, Triplex mode, Dual Live mode, M mode, Color Doppler mode, Anatomical M mode, PW doppler, Power doppler, Advanced technology for slow flow detection withut aliasing
| 13 | Minimum 490 fps in 2D mode
| 14 | System dynamic range minimum 250 dB
| 29 | Image zoom before and after freezing
| 30 | System supports intraoperative probes
| 33 | Convex, abdominal, multifrequency probe, frequency range of min 2.0-5.0 MHz with minimum 128 elements
| 34 | Linear multifrequency probe, frequency range of min 3.0-11.0 MHz with minimum 192 elements
| 35 | System supports following advanced technology transducers: matrix, single crystal or hanafy
| 36 | Black/white digital video printer
8 | 1.8 | PREMIUM ULTRASOUND COLOR DOPPLER DEVICE |  | 1 |  | 0
| 1 | Integrated color LCD or TFT display of high resolution 21”, on artificial arm
| 2 | Operating panel is height adjustable with background illumination of buttons
| 3 | Control LCD touch screen (diagonal minimal 10”) integrated in operating panel for faster access to device functions and measurements.
| 4 | Integrated full QWERTY keyboard illuminated or with background illumination which can be retracted from operating panel for less space consumption
| 5 | High device maneuverability with 4 rotating wheels not portable in concept
| 6 | Minimum 3 active probe connectors (not counting connector for pencil probes) and every offered probe can be connected to any connector, electronic probe selection
| 7 | Every offered probe must have choice of minimum 3 fundamental operating frequencies
| 8 | Modes of acquisition: B mode, M mode, color M mode, color doppler, PW doppler with HPRF, "Power" doppler, "Tissue Doppler Imaging", CW Doppler
| 9 | Maximal scan range in B mode, at least 40 cm
| 10 | Simultaneous real time display of B mode image and Color Doppler image
| 11 | At least 192 hardware transmitters
| 12 | "Tissue Harmonic Imaging" based on the following latest techniques: pulse subtraction, pulse phase inversion, wide band pulse inversion coded harmonics etc. with minimum 3 operating frequencies in all the tendered probes
| 13 | Spatial image compositing in order to reduce the noise in real time, Real time sonoCT, CROSSX Beam, SieClear, Aplipure or equivalent.
| 14 | One of the techniques of automatic image optimization in B-mode at a single push of a button: TEQ, iScan, QuickScan, ATO.. or equivalent
| 29 | Automatic baseline optimization and speed range in PW/CW mode at a single push of a button.
| 30 | One of the advanced color doppler techniques for better display of slow flows without overflow: B-flow, Advanced Dynamic Flow, Clarify or equivalent. Directional “Power” doppler is not acknowledged.
| 33 | Automatic trace in PW doppler mode with measurement
| 34 | Included option of Trapezoid scan
| 35 | Automatic measurement of intima-media complex
| 36 | Standard set of measurements in B-mode and Doppler as well as the set of cardiology measurement including: LV mass, PISA and MPI
| 37 | Possibility to upgrade the device for scan with contrasts during cardiological examinations
| 38 | “Speckle Reduction“ option
| 39 | Possibility to upgrade the device with the program for compression elastography
| 40 | The device must have the technique to improve the display of needle during the biopsy (technique name must be indicated)
| 41 | Possibility to enlarge the image without losing the diagnostic quality and increase in image granulation (“HD Zoom“or equivalent)
| 42 | Possibility to attach “Single Crystal“ probes to the device
| 43 | Possibility to upgrade the device with the program for “Shear Wave“ elastography
| 44 | Possibility to activate on the device the Fusion Imaging (live display of the same anatomic region live through ultrasound and with the previously performed CT or MR recording)
| 45 | The device features software for stress for echo
| 46 | The device has the display of ECG signal on the screen
| 47 | The device features the software package for the analysis or myocardial movement through “speckle tracking“ technology the following features
| 48 | The technique enables the quantification of myocardial wall movement
| 49 | Calculation also includes EF
| 50 | The long axis can calculate “strain“ “strain rate“ “disinchrony imaging“ for longitudinal and transversal “strain“
| 51 | The short axis can calculate “strain“ “strain rate“ “disinchrony imaging“ for radial and circumferential “strain“
| 52 | Based on the measurement is created the “bull-eye“ display chart of regional myocardial contractiveness
| 53 | Possibility to display the curve of global “strain“ and left chamber volume
| 54 | Memory of minimum 512 MB
| 55 | Prospective and retrospective clip recording
| 56 | Storing images in DICOM format on HDD, DVD/CD, USB Flash memory
| 57 | Export of images in JPEG format and cine clips in AVI (MPEG-4) format
| 58 | Network connection minimum 1000 Mbps
| 59 | The device has minimum 3 USB ports
| 60 | The device can be upgraded with the following DICOM functions: Store, Print, Query/Retrieve, Verification, Modality Worklist Management, MPPS, Structured reporting, DICOM Multi Frame, DICOM Media Storage
| 61 | Possibility for the user to draw up personalized reports on the device with possibility to export them in PDF format and possibility to connect the device with the office printer for printing reports
| 62 | Possibility to connect convex and linear and TEE probes
| 63 | Sector cardiological probe in the frequency range from 2,0-4,0 MHz or greater, visual field width minimum 90⁰
| 64 | Black-white thermal printer
9 | 1.9 | Defibrillator |  | 1 |  | 0
| 1 | Portable biphasic defibrillator for external defibrillation, synchronous cardioversion and AED
| 2 | Color TFT LCD display, min. 6.5", capable for min. 3 waveforms and 5 numeric parameters
| 3 | Defibrillator needs to have integrated carrying handle
| 4 | External defibrillation paddles with adult and pediatric size
| 5 | Charge and shock buttons needs to be integrated in external paddles
| 6 | Visual indicator for skin impedance
| 7 | Power supply: AC 220V/50Hz and backup battery for continuous ECG monitoring min.2.5h
| 8 | Battery capacity sufficient for min. 100 defibrillations with 200J
| 9 | Battery charging time: max. 4h
| 10 | Defibrillation energy selectable: from 2J to 200J or higher, biphasic
| 11 | Energy charging time from 0 to 200J: max 6s, powered by AC or powered by battery
| 12 | Fast ECG base line recovery after defibrillation: max. 7s
| 13 | Automatic detection of R peak during synchronized cardio version
| 14 | Automatic internal energy discharge after predefined time
| 29 | Visual and sound indication of energy charged
| 30 | Adjustable alarm limits for HR, SpO2, EtCO2 and apnea time
| 33 | Automatic detection of pacemaker signal
| 34 | Automatic arrhythmia detection with sound and visual alarms
| 35 | Automatic defibrillation reports, with storing and recording from memory.
| 36 | Defibrillation report with ECG waveform
| 37 | Automatic arrhythmia storing, with numeric parameters and printing reports
| 38 | Capable for storing trends min. 24h
| 39 | Integrated 2-ch thermal recorder
| 40 | Operational temperature: 0 to +40°C or better
| 41 | Dust and moisture protection: min. IP34
| 42 | Weight including battery and paddles: max 7.5kg+/- 10%
| 43 | Upgradeable for monitoring of NIBP, SpO2 and EtCO2 (intubated and nonintubated)
| 44 | Upgradeable with bed mount hooks
| 45 | Upgradeable with internal defibrillation paddles
| 46 | Upgradeable for Bluetooth or USB communication
| 47 | Dedicated defibrillation trolley, with castors with brakes, included Reanimation set
| 48 | To be delivered with all accessories needed, including:
| 49 | ECG cable with 3 leads, 1 pc
| 50 | Battery, 1 pc
| 51 | Thermal paper, 10 pcs
10 | 1.1 | Laparoscopic instruments set |  | 1 |  | 0
| 1 | Spare handle for HIQ + laparoscopic instrument,autoclavable
| 2 | Spare valve tube for HIQ + handle, suction 5mm,autoclavable
| 3 | Light guide cable, 4.25mm,autoclavable
| 4 | Monopolar HF cable, 3.5mm,autoclavable
| 5 | Jaws for grasping forceps, Johann, 5x330mm,autoclavable
| 6 | Jaws for dissector, Maryland, 5x330 mm, autoclavable
| 7 | Dissection forceps, 10x330mm, autoclavable
| 8 | Handle for laparoscopic instruments, Ergo, monopolar, autoclavable
| 9 | Spare valve tube, for HIQ+ handle, suction 5mm, autoclavable
| 10 | Sealing cap for troacar, 10mm, autoclavable
| 11 | Sealing cap for troacar, 5mm, autoclavable
| 12 | Valve, transparent for 3.5 and 5.5 mm troacars, autoclavable
| 13 | Laparoscopic Scissors, Metzenbaum, HIQ+ 5x330mm, autoclavable
| 14 | Jaws for grasping forceps, 10x330mm, autoclavable
| 29 | Handle for laparoscopic instrument, Ergo, with ratchet, autoclavable
| 30 | Compatible with existing equipment (Olympus manufacturer)
11 | 1.11 | Video gastroscope |  | 1 |  | 0
| 1 | Field of view 140°,
| 2 | Depth of view 3-100 mm,
| 3 | Insertion tube outer diameter 9.2 mm,
| 4 | Distal end outer diameter 9.2 mm,
| 5 | Instrument channel inner diameter 2.8 mm,
| 6 | Working length 1030 mm,
| 7 | Bending up/down 210°/90°, left/right 100°/100°
| 8 | Compatible with existing equipment (Olympus manufacturer)
12 | 1.12 | Intraaortic balloon Pump |  | 1 |  | 0
| 1 | Intraaortic balloon Pump with fiberoptic technology.
| 2 | Automatic Lead & Trigger Selection.
| 3 | Automatic Inflate & Deflate Timing.
| 4 | User Options to Fine-Tune Deflation Timing within Auto Mode.
| 5 | Battery time for one battery 180 min.
| 6 | Doppler fixed on the IABP to control the arterial perfusion.
| 7 | Automatic condensate removal module.
| 8 | Zeroing possible after insertion of the IAB.
| 9 | In autofill mode the IABP automatically purge and fill the IABP when local atmospheric pressure decreases or increases.
| 10 | Automatic in vivo recalibration every 2 hours or sooner should patients or environmental conditions change.
| 11 | Fiber-optic pressure signal transmission to a patient monitor.
| 12 | An electrically isolated, low-level pressure output from the pump to send the fiber-optic arterial pressure waveform directly to a patient monitor by simply attaching an appropriate cable.
13 | 1.13 | Mobile radiography digital system with wireless FPD (Flat Panel Detector) |  | 1 |  | 0
| 1 | HV GENERATOR
| 2 | Power min. 35 kW
| 3 | Voltage in range of min 40-133 kV
| 4 | Maximum current value, min. 450 mA
| 5 | mAs values in range of min 0.4-320 mAs
| 6 | Minimum exposure time 1 ms
| 7 | Number of anatomical program min 100
| 8 | Exposure with battery power
| 9 | Battery power indicator
| 10 | 220V input for battery charging
| 11 | Hand switch exposure
| 12 | Remote controller for exposure activation (infra red or equivalent)
| 13 | Exposure possible when system is connected on 220V (when batteries are empty and system is on charging)
| 14 | X-RAY TUBE
| 15 | Rotational anode
| 16 | Anode heat capacity, min. 120 kHU
| 17 | Single focus max 0.8 mm or dual focus, small max 0.8 mm, big min 1.2 mm
| 18 | Manual or automatic collimation
| 19 | Collimator rotation min +/- 90º
| 20 | WIRELESS PORTABLE FLAT PANEL DETECTOR
| 21 | Wireless (WLAN) communication between detector and acquisition workstation intagrated in the system
| 22 | Scintilator material: CsI (cezijum-jodid)
| 23 | Detector dimensions min: 35х43 cm
| 24 | Acquisition depth min. 16 bits
| 25 | Pixel size max. 150 μm
| 26 | Thickness max 20 mm
| 27 | Detector weight: max 3 kg
| 28 | Min two batteries for normal operation of flat panel detector
| 29 | Battery charger with min. 2 slots for simultaneous charging
| 30 | Stationary grid (for scattered radiation)
| 31 | ACQUISITION WORKSTATION INTEGRATED IN SYSTEM
| 32 | LCD touch screen monitor diagonal size min. 17"
| 33 | Disk storage capacity of min 10.000 images
| 34 | Image preview in max 3 seconds after exposure
| 35 | Gray scale inversion
| 36 | Zoom
| 37 | Image Rotation
| 38 | L/R markers
| 39 | Annotations
| 40 | Wireless interface for exporting images in DICOM format
| 41 | DICOM functionality min: Send, Storage Commitment Print, MWM
| 42 | Image archiving on CD/DVD and USB in DICOM format
| 43 | Exposure Index (EI) display
| 44 | MECHANICAL PARAMETERS
| 45 | Motorized back/forward movement on batteries
| 46 | Possibility of selection different speed movement
| 47 | Buttons for fine movement of the system back/forward on front side of collimator or on arm, next to X-ray tube, for easy and faster tube positioning, without need to go in fromt of the system
| 48 | Distance from focal spot to floor in vertical direction, in range of min 500 – 2100 mm
| 49 | Stand rotation min +/- 90°
| 50 | X-ray tube rotation around arm min +/- 150º
| 51 | X-ray tube rotation around vertical axes min 90°/-20°
| 52 | System weight max 400 kg
| 53 | System width max 600 mm
| 54 | System height in transport position, max 1600 mm
14 | 1.14 | Endoscopic video tower with duodenoscope |  | 1 |  | 0
| 1 | Endoscopic video tower with duodenoscope
| 2 | Video system center with integrated light source:
| 3 | LED light source
| 4 | Analogue HDTV signal output: RGB (1080/60I: NTSC)/(1080/50I: PAL) or YPbPr (1080/60I: NTSC)/(1080/50I: PAL) output
| 5 | Analogue SDTV signal output: VBS composite (480/60I: NTSC)/(576/50I: PAL), Y/C (480/60I: NTSC)/(576/50I: PAL) and RGB (480/60I: NTSC)/(576/50I: PAL);
| 6 | Digital signal output: HD-SDI (SMTPE 292M), SD-SDI (SMPTE 259M) and DVI (WUXGA, 1080p or SXGA).
| 7 | White balance adjustment
| 8 | Automatic brightness adjustment: Automatic exposure in 13 steps
| 9 | Colour tone adjustment: red adjustment: ±8 steps, blue adjustment: ±8 steps and chroma adjustment: ±8 steps.
| 10 | Enhancement of the visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light ( not by softwatre enhacement) .
| 11 | Automatic gain control (AGC)
| 12 | Iris control: Peak and average
| 13 | Contrast: · N (Normal): H (High): L (Low):
| 14 | Electronic zoom function 1x; 1,2x and 1,5x
| 15 | Displaying the endoscope information : Displays the position that a diameter of the instrument channel and an accessory can be seen, the setting of endoscope’s remote switches and the serial number, the diameter of the instrument channel,the diameter of the distal end, and the diameter of the insertion tube of the endoscope.
| 16 | Freeze function
| 17 | Pre-freeze function - selects the clearest still image automatically.
| 18 | Air-feeding : Diaphragm type pump - 3-level (OFF, low, high)
| 19 | Registration of patient data: 50 patients
| 20 | Memorization of user settings: 20 users
| 21 | 2GB Portable memory to store endoscopic images
| 22 | Compatible with existing equipment (OLYMPUS EXERA II Gastroscopes)
| 23 | Mobile workstation:
| 24 | Mobile workstation for endoscopic equipment
| 25 | Gas assisted LCD Monitor Arm
| 26 | Antistatic castors (two with brakes)
| 27 | Scope hanger
| 28 | Main ON/OFF switch
| 29 | Separation transformer
| 30 | Duodenoscope:
| 31 | Dual locking mechanism - center and side.
| 32 | Field of view 100°/ 5°retro
| 33 | Depth of field 5-60mm
| 34 | Enhancement of the visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light
| 35 | Insertion tube Outer diameter: 11.3mm
| 36 | Distal end outer diameter: 13.7mm
| 37 | Working length: 1240mm
| 38 | Instrument channel 4.2mm
| 39 | Angulation range UP 120° / DOWN 90°
| 40 | LEFT 90° / RIGHT 110°
| 41 | Distal end cap fixed
| 42 | Suction pump:
| 43 | Nominal vacuum: 95 kPa
| 44 | pre-set flow speed settings 40l/min, 50l/min and 60l/min
| 45 | Additional vacuum regulation knob
| 46 | Vacuum gauge
| 47 | Thermal cut-out protection
| 48 | Suction jar
| 49 | Tubing
| 50 | Disposable bacteria filter 10 pcs
| 51 | Leakage tester
| 52 | LCD Monitor:
| 53 | 27’’ Full HD LCD Monitor
| 54 | Aspect 16:9
| 55 | Luminance 300 cd/m2
| 56 | Contrast Ratio 1000 : 1
| 57 | Colors: Approx.16.7 million colors
| 58 | Viewing Angle) 89°/89°/89°/89° (up/down/left/right)
| 59 | Inputs:
| 60 | Composite Input BNC (x1)
| 61 | Y/C Input Mini-DIN 4-pin (x1)
| 62 | RGB, Component Input RGB: Via HD-15 connector (D-sub 15-pin)
| 63 | DVI-D Input DVI-D (x1)
| 64 | SDI Input BNC (x1)
| 65 | HD/SD-SDI
| 66 | HD15 Input D-sub 15-pin (x1)
| 67 | Output
| 68 | DVI-D Output
15 | 1.15 | Rigid Macintosh Laryngoscope Set |  | 3 |  | 0
Macintosh Laryngoscope Blades:
Paediatric 1 Length: 95mm, Distal width: 13mm,
Macintosh 2 Length: 115mm, Distal width: 13mm,
Macintosh 3, Length: 135mm, Distal width: 15mm,
Macintosh 4, Length: 155mm, Distal width: 15mm
Minimum 6200 individual micro-fibers for improved light transmission
Chrome-plated, stainless, high-grade steel
Minimum guarantee for all blades with HiLite fiber bundles 5 years
Minimum 35000 standard autoclave cycles of 134°C/4min. with more than 1000 lux.
Rechargeable Handle with Li-ion Battery and table charger
Laryngoscope Handle compatible with laryngoscope blades to the ISO 7376 (green standard)
LED 3.5V Li-ion illumination to define optimal light intensity
Fade-Out Feature
Reference Colour Temperature: 5,000K
Charge Status Indicator in the bottom insert.
Minimum 12 hours operation time
Zipper case
16 | 1.16 | Laryngoscope set for a difficult intubation |  | 3 |  | 0
Macintosh Laryngoscope Blade with a tip which is adjustable through 70°
Minimum 6200 individual micro-fibers for improved light transmission
Chrome-plated, stainless, high-grade steel
Swivelling blade tip
Minimum 35000 standard autoclave cycles of 134°C/4min. with more than 1000 lux.
LED Blade sizes: 3 and 4, Length: 115mm/137mm, Distal width: 3.2mm/13.2mm
Rechargeable Handle with Li-ion Battery and table charger
Laryngoscope Handle compatible with laryngoscope blades to the ISO 7376 (green standard)
LED 3.5V Li-ion illumination to define optimal light intensity
Fade-Out Feature
Reference Colour Temperature: 5,000K
Charge Status Indicator in the bottom insert.
Minimum 12 hours operation time
Hard case