## PriceSchedule
Bidder: |  |  |  | Offered price:
5_Radiology_Imaging
Item No. | Equipment name | Qty | Model, Manufacturer and Country of Origine | Unit price | Total Price per line item
1 | Ceiling mounted x-ray device | 1 |  |  | 0
2 | Mobile x-ray device | 1 |  |  | 0
3 | Digital mammography system | 1 |  |  | 0
4 | Mobile digital X-ray with C-arm | 1 |  |  | 0
5 | Premium Ultrasound system with 3 probes | 1 |  |  | 0
6 | Ultrasound system with 3 probes | 1 |  |  | 0
7 | Ultrasound system for gynecology | 1 |  |  | 0
8 | Portable ultrasound system with 2 probes for ICU | 1 |  |  | 0
9 | Ultrasound device for Neurology | 1 |  |  | 0
10 | Ultrasound system for Emergency room | 1 |  |  | 0
11 | Ultrasound system for pediatric cardiology | 1 |  |  | 0
12 | Ultrasound system for adult cardiology | 1 |  |  | 0
Total price per Lot |  |  |  |  | 0

## TechSpecification
Bidder: | [5 empty] | Total DAP price:
5_Radiology_Imaging
Item No. | ID | Technical Specification Requested | Technical Specification Offered | Qty | Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Ceiling mounted x-ray device |  | 1 |  | 0
| 1.1 | Patient table
| 1.2 | Table top motorized height adjustment from floor min. 55-80 cm
| 1.3 | Table top size min. 230x80 cm
| 1.4 | Tabletop transversal movement range min. 25 cm
| 1.5 | Tabletop longitudinal movement range min. 65 cm
| 1.6 | Max. patient weight during hegiht adjusment or floating table top min.300±10kg
| 1.7 | Automatic detector charging in bucky
| 1.8 | Automated detector longitudinal tracking to align with tube (including tube longitudinal and angulation)
| 1.9 | Auto-tracking during table height: X-ray tube follows table height adjustment to keep fixed SID
| 1.10 | Motorized height adjustment foot controlled
| 1.11 | Compression belt and table hand grips
| 1.12 | Removable grid focus min. 100 cm
| 1.13 | Lateral detector holder
| 1.14 | Ceiling mounted tube suspension
| 1.15 | Lateral, trasversal travel range min. 190 cm
| 1.16 | Vertical motorized travel range min. 160 cm
| 1.17 | Longitudinal travel range min. 300 cm
| 1.18 | Tube angulation range min ± 120°
| 1.19 | Column rotation range min. ± 180°
| 1.20 | Automated vertical tracking to align with Wallstand detector. Automated vertical tracking to maintain table SID
| 1.21 | LCD Touch Screen user interface
| 1.22 | User Interface: Display of SID, Tube Angle, Column Rotation, Patient Name and Date of Birth.
| 1.23 | User Interface adjustable parameters: at least kV, mAs
| 1.24 | Bucky wallstand
| 1.25 | Floor mounted fixed wall stand
| 1.26 | Tilting range min. +90°/-20°
| 1.27 | Min. height from floor to center of detector max. 45 cm
| 1.28 | Vertical travel range min. 145 cm
| 1.29 | Auto tracking with X-ray tube vertical movement
| 1.30 | Motorized and manual detector housing vertical movement
| 1.31 | Compatible with 43x43cm and 43x35cm detectors
| 1.32 | Removable grid for focus min. 130 cm
| 1.33 | Electromagnetic brakes
| 1.34 | Vertical motion control by foot switch
| 1.35 | Generator
| 1.36 | Nominal output min. 65 kW
| 1.37 | kV range min. 40-150 kV
| 1.38 | mAs range min. 0,5 – 600 mAs
| 1.39 | mA range: min. 20-800 mA
| 1.40 | Integrated AEC
| 1.41 | Predefined APR
| 1.42 | Tube and collimator
| 1.43 | Anode rotation speed min. 9000 rpm
| 1.44 | kV range min. 40 – 150 kV
| 1.45 | Focal spot nominal value: small (IEC 60336) max. 0,6 mm
| 1.46 | Focal spot nominal value: large (IEC 60336) min. 1,2 mm
| 1.47 | Anode heat storage capacity min. 300 kHU
| 1.48 | Inegrated DAP
| 1.49 | Rotation min. ± 45°
| 1.50 | Collimator blades are set manually and automatic by APR
| 1.51 | LED filed light
| 1.52 | Wireless flat detectors
| 1.53 | Number of wiresless detectors min. 2 pcs
| 1.54 | Detector sizes first: min 35x43 cm, second min. 43x43 cm
| 1.55 | Detectors are exchangable between wallstand and patient table
| 1.56 | Wireless detector thickness max. 19 mm
| 1.57 | In-bin charging, the detectors automatically charge when they are in the bucky housing
| 1.58 | Pixel size max. 125 µm
| 1.59 | Acquisition depth min. 16 bits
| 1.60 | Cesium Iodide (CSI) scintillator
| 1.64 | DQE @ 0lp/mm min. 60% (RQA 5)
| 1.65 | Ingress protection min IPX4
| 1.66 | Acquisition Workstation and Image processing
| 1.67 | Image storage capacity min. 15.000 images
| 1.68 | Image preview times max. 2 sec
| 1.69 | Full DICOM 3.0 compatibility
| 1.70 | DICOM Dose Structure Report
| 1.71 | Monitor size min 24”
| 1.72 | Integrated control station for controlling generator and detector functions, managing a work list and viewing images.
| 1.73 | Post processing: Windowing, rotate, zoom, annotate, mirror, etc.
| 1.74 | Automatic parameters settings according HIS/RIS protocol selection
| 1.75 | Automatic FOV according protocoll selected
| 1.76 | Noise reduction
| 1.77 | UPS
| 1.78 | Other:
| 1.79 | Power distribution box
2 |  | Mobile x-ray device |  | 1 |  | 0
| 2.1 | Main system
| 2.2 | Motorized forward and reverse direction
| 2.3 | Weight max 450 kg
| 2.4 | Width max 60 cm
| 2.5 | Focal spot range min. 68 – 200 cm
| 2.6 | Colum rotation min. ±180°
| 2.7 | Tube rotation around horizontal axis min. ±165°
| 2.8 | Tube tilting min. +90° / -10°
| 2.9 | Generator
| 2.10 | High frequncy generator min. 60 kHz
| 2.11 | Nominal power min. 30 kW
| 2.12 | kV range min 50-125 kV
| 2.13 | Max mA value min. 300 mA
| 2.14 | mAs range min 0.5-320 mAs
| 2.15 | X-Ray tube and collimator
| 2.16 | Two focal spots: max. 0.7 and min. 1.2 mm
| 2.17 | Anode heat capacity min. 120 kHU
| 2.18 | Collimator rotation min. ± 90°
| 2.19 | DAP
| 2.20 | Flat panel detector
| 2.21 | Detectors and mobile xray unit are from the same manufacturer
| 2.22 | CsI wireless detector
| 2.23 | Detector size min 35 x 43 cm
| 2.24 | Pixel size max 150 microns
| 2.25 | Removable snap-on anti scatter grid or virtual grid
| 2.26 | Detector and grid stored in the bin (only in case when phyisical grid is offered)
| 2.27 | Automatic in bin charging or external charger with additional battery
| 2.28 | User interface and processing
| 2.29 | LCD touchscreen user interface diagonal size min 15”
| 2.30 | Brightness and Contrast adjustment
| 2.31 | Image invert
| 2.32 | Zoom, Rotate, Markers, annotation
| 2.33 | Image storage capacity min 3.500 images
| 2.34 | Advanced Noise Reduction
| 2.35 | Repeat/Reject analisys
| 2.36 | Wireless DICOM 3.0 services: Storage, Storage Commitment, MWL, Print, MPPS
| 2.37 | Images export to CD/DVD or USB media for storage
3 |  | Digital mammography system |  | 1 |  | 0
| 3.1 | Generator
| 3.2 | High frequency generator
| 3.3 | kW min. 5 kW
| 3.4 | kV range min. 25-45 kV, with 1 kV steps
| 3.5 | Max mAs min. 500 mAs
| 3.6 | Tube and gantry
| 3.7 | Anode rotation speed min. 8000 rpm
| 3.8 | Anode heat capacity min. 300 kHU
| 3.9 | Anode material: Molybdenum (Mo) or Rhodium (Rh) or Tungsten (W)
| 3.10 | Focal spot size (IEC 60336) max 0.1 mm and 0.3 mm
| 3.11 | Automatix exposure mode KV, mAs, filter
| 3.12 | Specific paddle and AEC for implants
| 3.13 | Automatic collimation set to paddle size
| 3.14 | Motorized gantry vertical movement min. 70-140 cm
| 3.15 | Gantry isocentric motorized rotation min. -140/+180 degrees
| 3.16 | SID min. 65 cm
| 3.17 | 2 set of footswitches on both side of the gantry for column height and compression adjustments
| 3.18 | Compression paddles min. 5 different shapes
| 3.19 | Motorized and manual compression with numeric display
| 3.20 | Auto decompression after exposure
| 3.21 | Magnification factors 1.8
| 3.22 | Faceshield
| 3.23 | Acquisition workstation
| 3.24 | Workstation with integrated radiation shield
| 3.25 | Enter data manually or via DICOM worklist
| 3.26 | Repeat/Reject analysis
| 3.27 | Special processing algorithm for implant imaging
| 3.28 | Possibility of individual work under the user's credentials (name and password)
| 3.29 | PC memory min. 8 GB RAM
| 3.30 | HDD min. 1 TB
| 3.31 | LCD monitor min. 1MP min. 19”
| 3.32 | DICOM 3.0 (Send / Receive, Query / Retrieve, Storage Commitment, Print, Worklist, MPPS)
| 3.33 | UPS
| 3.34 | Flat panel digital detector
| 3.35 | Detector size min. 24 x 29 cm
| 3.36 | CsI or aSe
| 3.37 | Pixel size (2D and 3D mode) max. 100 µm
| 3.39 | Anti-scatter grid
| 3.40 | Tomosynthesis
| 3.41 | Max scaning range min. 25°
| 3.42 | Scan time max 10 sec
| 3.43 | Projections min 9
| 3.44 | Synthetic 2D image
| 3.45 | Distance between reconstructed planes max. 1 mm
| 3.46 | Diagnostic workstation
| 3.47 | Two medical TFT monitors, min. 21-inch, resolution min. 5MP or one min. 10MP monitor approved for mammography diagnostics.
| 3.48 | Workstation with dedicated software for 2D and tomosynthesis images
| 3.49 | Dedicated tools to review tomosynthesis data sets: cine-loop, bookmarks, breast localizer, breast height ruler
| 3.50 | LCD monitor for patient administration min. 19”
| 3.51 | Memory min. 16 GB RAM
| 3.52 | HDD min. 2 TB
| 3.53 | Special programmable mammography keypad
| 3.54 | Database search by multiple criteria, patient name, patient ID, type of modality, study date, etc.
| 3.55 | Image processing tools: contrast, brightness, inversion, zoom, rotation, annotations, measurements
| 3.56 | Support for other modalities, US, MRI, etc.
| 3.57 | Hanging Protocols
| 3.58 | Automatic pre-fetching of prior exams with a visual indication of the pre-fetching status
| 3.59 | All images can be zoomed simultaneously and panned with auto-alignment
| 3.60 | Displays plane of point selected in corresponding image view
| 3.61 | Possibility to displays CAD markers for micro-calcs and masses
| 3.62 | Upgradability:
| 3.63 | Contrast Enhanced Mammography
| 3.64 | Tomo guided biopsy
| 3.65 | Contrast Enhanced Mammography guided biopsy
| 3.66 | Other
| 3.67 | Power distribution box
4 |  | Mobile digital X-ray with C-arm |  | 1 |  | 0
| 4.1 | Mechanical parameters
| 4.2 | C-arm free space min. 75 cm
| 4.3 | C-arm depth min. 65 cm
| 4.4 | C-arm orbital rotation min. 150º
| 4.5 | C-arm angulation min. ± 200°
| 4.6 | C-arm vertical motorized movement min. 40 cm
| 4.7 | C-arm horizontal movement min. 20 cm
| 4.8 | C-arm/ wig-wag movement min. +/-10º
| 4.9 | Touch screen user interface diagonal size min. 10 inch
| 4.11 | Generator
| 4.12 | KV range min. 40-110 kV
| 4.13 | High frequency generator
| 4.14 | Monobloc design
| 4.15 | Pulsed fluoro mode min. 20 mA
| 4.16 | Digital Spot Mode min. 20 mA
| 4.17 | Generator power min 2.4 kW
| 4.18 | Max Pulse frequency min. 8 pps
| 4.19 | X-ray tube
| 4.20 | Stationary -anode
| 4.21 | Single or dual focus
| 4.22 | Small focus size max 0.6 mm
| 4.23 | Anode heat capacity min. 70 kHU
| 4.24 | Anode Cooling Rate min. 35 kHU/min
| 4.25 | Collimator
| 4.26 | Slot collimator with preview
| 4.27 | Virtual collimation without fluoroscopy
| 4.28 | Flat Detector
| 4.29 | CMOS min. 30fps full frame rate
| 4.30 | Size min. 20x20 cm
| 4.31 | Number of field of views min. 3
| 4.32 | DQE (typical) min. 70% at 0lp/mm
| 4.33 | Pixel size max. 140 microns
| 4.34 | Laser aimer integrated in flat panel
| 4.35 | Monitor(s) integrated on C-arm
| 4.36 | High resolution LCD touch screen monitor size min. 27” with dual screen view or two LCD touch screen monitors min. 19” inch each
| 4.37 | Image matrix min. 1k
| 4.38 | Max. brightness of monitor(s) min. 550 cd/m2
| 4.39 | 100% retention of Field of View without image cropping, distortion, or cut-off – must be maintained in at least 0/90/180/270 degree rotated angular position of the detector
| 4.40 | Image Display and porcessing
| 4.41 | Digital image rotation
| 4.42 | Left / Right Image Mirroring
| 4.43 | Up / Down Image Mirroring
| 4.44 | Measurements
| 4.45 | Current kV and mA values on the screen
| 4.46 | Automatic kV and mA adjustment to optimize image quality
| 4.47 | Automatic picture swap between left and right monitor
| 4.48 | Last image hold (LIH)
| 4.49 | Automatically adjusts brightness and contrast when metal object is present in the imaging field
| 4.50 | Wireless DICOM Storage commitment, Store and Print
| 4.51 | Export images to USB with DICOM image viewer
| 4.52 | Storage min. 100.000 images
5 |  | Premium Ultrasound system with 3 probes |  | 1 |  | 0
| 5.1 | Ultrasound system, not portable, with wheels and front and rear handles
| 5.2 | Four active probe ports, not counting pencil probe connector
| 5.3 | OLED or LCD monitor min. 22'', wide screen, high-resolution (min. 1920 x 1080) with translation (independent of console), swivel and fold down
| 5.4 | Wide, high-resolution, color touch LCD screen, min. 10,” with interactive menu, brightness adjustment and user-configurable layout
| 5.5 | Operating keyboard adjustable in height and rotation
| 5.6 | Backlit alphanumeric keyboard
| 5.7 | Integrated SSD with disk encryption (encrypts patient data archive partition)
| 5.8 | Integrated DVD multi drive
| 5.9 | Integrated speakers
| 5.10 | Integrated gel warmer
| 5.11 | Maximum depth of field min. 30 cm (probe dependent)
| 5.12 | Coded harmonic imaging
| 5.13 | Supported transducer types:
| 5.14 | Convex array
| 5.15 | Micro convex array
| 5.16 | Bi-plane micro convex probe
| 5.17 | Linear array
| 5.18 | Sector phased array
| 5.19 | Matrix array
| 5.20 | Intraoperative probes
| 5.21 | Biopsy convex probe
| 5.22 | Volume probes (convex array, micro convex array, linear array)
| 5.23 | TEE probe
| 5.24 | Split crystal probes
| 5.25 | Operating modes:
| 5.26 | B-Mode
| 5.27 | M-Mode
| 5.28 | Anatomical M-Mode
| 5.29 | Curved Anatomical M-Mode
| 5.30 | Color Flow Mode (CFM)
| 5.31 | Power Doppler Imaging (PDI)
| 5.32 | PW Doppler
| 5.33 | 3D acquisition with 2D probes including acquisition of color data as Advanced 3D or corresponding
| 5.34 | Convex field of view available on linear and sector probes
| 5.35 | Speckle reduction imaging which provides multiple levels of speckle reduction
| 5.36 | Variable angle spatial compounding
| 5.37 | Automatic optimization of B-Mode image to improve axial and lateral gain uniformity, overall gain level and contrast resolution
| 5.38 | Auto spectral optimization – adjusts baseline, invert, PRF (on live image) and angle correction
| 5.39 | Method for showing hemodynamics in grayscale using B-Mode technology (not the Doppler technology) which is based on reflection of blood cells (direct, real-time visualization of blood flow echoes)
| 5.40 | On-board reporting package for report writing
| 5.41 | DICOM 3.0
| 5.42 | CINE memory min. 700 MB
| 5.43 | Probes:
| 5.44 | Convex probe Bandwidth: 1 – 6 MHz or wider Number of elements: min. 192 Field of view (FOV): min. 70°
| 5.45 | Linear probe Bandwidth: 2 – 11 MHz or wider Number of elements: min. 256 Field of view (FOV): 50 mm (+/- 5 mm)
| 5.46 | Linear probe Bandwidth: 2 – 8 MHz or wider Number of elements: min. 192 Field of view (FOV): 45 mm (+/- 5 mm)
| 5.47 | Multiple USB 2.0 and USB 3.0 ports
| 5.48 | Ethernet
| 5.49 | DVI-D or HDMI connector
| 5.50 | Digital BW thermal printer
6 |  | Ultrasound system with 3 probes |  | 1 |  | 0
| 6.1 | Ultrasound system, not portable, with wheels and front and rear handles
| 6.3 | Four active probe ports, not counting pencil probe connector
| 6.4 | Minimum 23'' Bezel-less LCD LED backlight monitor
| 6.5 | Articulating monitor arm
| 6.6 | Advanced user interface with high resolution min. 10'' wide LCD touch screen
| 6.7 | Operator Keyboard swivel-adjustable and height-adjustable
| 6.8 | Integrated solid state drive min. 500 GB
| 6.9 | Speckle reduction imaging which provides multiple levels of speckle reduction
| 6.10 | Spatial compounding with minimum 9 angles
| 6.11 | Automatic optimization
| 6.12 | Digital TGC and digital A/N keyboard
| 6.13 | Auto TGC
| 6.14 | Maximum frame rate minimum 3200 F/s (depends on probes and modes)
| 6.15 | Maximum depth of field minimum 45 cm (probe dependent)
| 6.16 | Convex field of view on linear probes a
| 6.17 | Max Read Zoom to 20x: baseline shift on “freeze” or recall
| 6.18 | Continuous Wave Doppler
| 6.19 | The ability to import CT/MRI images for comparison on a split screen with active ultrasound image
| 6.20 | Harmonic imaging
| 6.21 | Transducer types:
| 6.22 | Convex array
| 6.23 | Linear array
| 6.24 | Linear with 4 configurable buttons
| 6.25 | Microconvex array
| 6.26 | Intraoperative
| 6.27 | Sector phased array
| 6.28 | Operating modes:
| 6.29 | B-Mode
| 6.30 | M-Mode
| 6.31 | Anatomical M-Mode
| 6.32 | Curved anatomical M-Mode
| 6.33 | M-Color Flow Mode
| 6.34 | PW Doppler with high PRF
| 6.35 | Color Flow Mode (CFM)
| 6.36 | Power Doppler Imaging (PDI)
| 6.37 | Ability to display ultrasound image at the same time on monitor and on the touch panel
| 6.38 | Hardware and software options which can be added later:
| 6.39 | Method for showing flow in grayscale using B-Mode technology (not the doppler technology), which is based on reflection of blood cells, such as B-Flow, SieFlow or corresponding
| 6.40 | Shear Wave Elastography
| 6.41 | DICOM 3.0
| 6.42 | Ethernet network connection
| 6.43 | Inputs and outputs: HDMI out, Composite video out, S-video out, USB
| 6.44 | Probes:
| 6.45 | Convex array probe Bandwidth: 1 – 6 MHz or wider FOV: minimum 70° Number of elements: minimum 192
| 6.46 | Linear array probe Bandwidth: 2 – 11 MHz or wider FOV: 50 mm (+/- 5 mm) Number of elements: minimum 256
| 6.47 | Sector array probe Bandwidth: 1 – 5 MHz or wider FOV: minimum 120° Number of elements: minimum 64
| 6.48 | Digital BW thermal printer
7 |  | Ultrasound system for gynecology |  | 1 |  | 0
| 7.1 | Ultrasound system, not portable, with wheels and front and rear handles
| 7.2 | Minimum 3 active probe ports
| 7.4 | Dynamic range min. 265 Db
| 7.5 | 23'' High Resolution LED Monitor in a Articulating Monitor Arm with DVI interface. Tilt/Rotate Adjustable. Resolution min 1920 x 1080
| 7.6 | Integrated HDD min. 500 GB
| 7.7 | Widescreen ultrasound image mode
| 7.8 | Control panel with integrated alphanumeric keyboard, backlit, height-adjustable and rotation
| 7.9 | Integrated DVD+R(W)/CD-R(W) drive
| 7.10 | Maximum depth of field min. 40 cm, probe dependent
| 7.11 | Maximum Frame rate min. 1200 fps, probe and mode dependent
| 7.12 | Coded Harmonic Imaging with Pulse Inversion
| 7.13 | Technology which improves image resolution and penetration in the far field which allows to use a higher frequency on technically difficult-to-scan patients
| 7.14 | Technology which utilizes two different transmit frequencies (a low frequency to increase the penetration and higher frequency to keep a high resolution) and two different focal zones in the 2D image for reducing speckle and artifacts in the 2D image to facilitate the examination of difficult-to-scan patients
| 7.15 | Speckle reduction imaging which provides multiple levels (min. 5) of speckle reduction
| 7.16 | Method of spatial compouding of the image from several different angles
| 7.17 | Automatic tissue optimization.
| 7.18 | Auto TGC
| 7.19 | Method to construct and view a static 2D image which is wider than the field of view of a given transducer
| 7.20 | Method which provides to adjust scanning plane of volume probes in 2D mode using the button on the control panel without moving the probe , in order to provide more comfortable transvaginal exam
| 7.21 | Write Zoom up to 8x and read zoom with HD-Zoom functionality up to 22x .
| 7.22 | Semiautomatic determination of nuchal and intracranial translucency (NT and IT)
| 7.23 | Semiautomatic Fetal Biometry for BPD, AC, HC, HL, FL.
| 7.24 | At least 4 programmable keys for direct access to commonly used probes, programs and modes as Probe Favorites or corresponding
| 7.25 | Transducer type:
| 7.26 | Convex array
| 7.27 | Linear array
| 7.28 | Matrix linear with more than 1000 elements
| 7.29 | Microconvex array
| 7.30 | Sector array
| 7.31 | Volume Probe (4D) Convex array
| 7.32 | Volume Probe (4D) Microconvex array
| 7.33 | Operating modes:
| 7.34 | B-Mode
| 7.35 | M-Mode
| 7.36 | Color Flow Doppler (CFM) mod
| 7.37 | Anatomical M-Mode
| 7.38 | Power Doppler Mode
| 7.39 | PW-Doppler with HPRF
| 7.40 | High definition Power Doppler as HD-Flow or corresponding
| 7.41 | Tissue Doppler – TD
| 7.42 | Volume mod (3D/4D)
| 7.43 | Mandatory functions in 3D / 4D mode:
| 7.44 | Static 3D
| 7.45 | 4D Realtime
| 7.46 | Method which during 3D/4D acquisition automatically determines the position and shape of the render start line and optimally adjusts it for the best adaptation to the render object
| 7.47 | Hardware and software options which can be added later:
| 7.48 | Battery that allows you to scan without mains power and move the system without turning it off
| 7.49 | The ability to obtain any slice in actual and projected view from 3D/ 4D acquisition by cutting with straight, curved and freeform line or freeform line from point to point
| 7.50 | Archiving and connections:
| 7.51 | Ethernet network connection
| 7.52 | DICOM 3.0
| 7.53 | Min. 5 USB ports
| 7.54 | Export data in formats: Raw Data, BMP, TIFF, JPEG; AVI, MP4, DICOM Files
| 7.55 | Export data in formats compatible with 3D printers (STL, OBJ, PLY, 3MF, XYZ)
| 7.56 | Probes:
| 7.57 | 4D (Volume) Convex probe Bandwidth: 2 – 8 MHz or wider Number of elements: min. 192 Maximum FOV (wide): 2D min. 90° volume scan min. 85°x 90°
| 7.58 | Wideband Micro-Convex Volume (3D/4D) Probe: Bandwidth: 4,0 – 9,0 MHz (+/- 0,5 MHz) Number of Elements: min 192 Maximum FOV (wide): 2D min. 180° volume scan min. 180°x 120°
| 7.59 | Digital BW thermal printer
8 |  | Portable ultrasound system with 2 probes for ICU |  | 1 |  | 0
| 8.1 | Laptop style portable ultrasound system
| 8.2 | System weight maximum 5 kg (with battery)
| 8.3 | Cart for the system, height adjustable, with min. 3 probe ports
| 8.4 | Beamformer of advanced architecture
| 8.5 | Systemic dynamic range min. 265 dB
| 8.6 | High-resolution (min.1920 x 1080 pixels) LCD monitor min.15'' with brightness, contrast and color temperature adjustment
| 8.7 | Integrated SSD min. 256 GB
| 8.8 | Min. 8 TGC pods
| 8.9 | Max. number of focal points min. 8, probe and application dependent
| 8.10 | Coded harmonic imaging
| 8.11 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 8.12 | Method of spatial compouding of the image from several different angles
| 8.13 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 8.14 | Full alphanumeric keypad covered with washable protection film
| 8.15 | Maximum depth of field min. 30 cm, probe dependent
| 8.16 | Technique which provides a convex field of view on linear and sector transducers
| 8.17 | HD write zoom and read zoom min. 50X
| 8.18 | Real-time automatic Doppler calculations
| 8.19 | Voice comments recording
| 8.20 | Transducer type:
| 8.21 | Convex array
| 8.22 | Linear array
| 8.23 | Microconvex array
| 8.24 | Sector phased array
| 8.25 | Volume Probe (4D)
| 8.26 | Operating modes:
| 8.27 | B-Mode
| 8.28 | M-Mode
| 8.29 | Color M-Mode
| 8.30 | PW Doppler with high PRF
| 8.31 | Color Flow mode (CFM)
| 8.32 | Power Doppler Imaging (PDI) and Directional PDI
| 8.33 | Inputs and outputs: min. 3 x USB, HDMI output (1920 x 1080), Ethernet (RJ45)
| 8.34 | Storage formats: - DICOM – compressed/uncompressed, single/multi-frame, with/without raw data - Export JPEG, WMV (MPEG 4) and AVI formats
| 8.35 | DICOM 3.0 connectivity
| 8.36 | B/W thermal printer
| 8.37 | Probes:
| 8.38 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 8.39 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
9 |  | Ultrasound device for Neurology |  | 1 |  | 0
| 9.1 | Ultrasound system, not portable, with wheels and front and rear handles
| 9.2 | Beamformer of advanced architecture
| 9.3 | Min 3 active probe connectors not counting the pensil probe connector
| 9.4 | Systemic dynamic range min. 265 dB
| 9.5 | LED monitor min. 21" high resolution min. 1920 x 1080
| 9.6 | Height adjustable operator keyboard
| 9.7 | Integrated HDD (500 GB)
| 9.8 | Minimum 8 TGC pods
| 9.9 | Max. number of focal points min. 8, probe and application dependent
| 9.10 | Coded harmonic imaging
| 9.11 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction as SRI-HD or corresponding
| 9.12 | Method of spatial compouding of the image from several different angles
| 9.13 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 9.14 | Full alphanumeric keypad covered with washable protection film
| 9.15 | Maximum depth of field min. 30 cm, probe dependent
| 9.16 | Technique which provides a convex field of view on linear and sector transducers
| 9.17 | HD write zoom - read zoom min. 50X
| 9.18 | Voice comments recording
| 9.19 | Transducer type:
| 9.20 | Convex array
| 9.21 | Linear array
| 9.22 | Microconvex array
| 9.23 | Sector phased array
| 9.24 | Volume Probe (4D)
| 9.25 | Operating modes:
| 9.26 | B-Mode
| 9.27 | M-Mode
| 9.28 | Color M-Mode
| 9.29 | PW Doppler with high PRF
| 9.30 | Color Flow mode (CFM)
| 9.31 | Power Doppler Imaging (PDI) and Directional PDI
| 9.32 | Hardware and software options which can be added later:
| 9.33 | Lithium-ion battery that allows the device to operate without mains power min. 15 minutes
| 9.34 | Method for showing hemodynamics in grayscale using B-Mode technology (not the Doppler technology) which is based on reflection of blood cells (direct, real-time visualization of blood flow echoes) as B-flow or corresponding
| 9.35 | Inputs and outputs:HDMI output (1920 x 1080 resolution), VGA output (1920 x 1080 resolution), CVBS output, S-Video output, 4 x USB, Ethernet (RJ45)
| 9.36 | B/W thermal printer
| 9.37 | Probes:
| 9.38 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
| 9.39 | B/W thermal printer
| 9.40 | Probes:
| 9.41 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
10 |  | Ultrasound system for Emergency room |  | 1 |  | 0
| 10.1 | Ultrasound system, not portable, with wheels and front and rear handles
| 10.2 | Beamformer of advanced architecture
| 10.3 | Min 3 active probe connectors not counting the pensil probe connector
| 10.4 | Systemic dynamic range min. 265 dB
| 10.5 | Frequency range 2 - 13 MHz or wider
| 10.6 | LED monitor min. 21" high resolution min. 1920 x 1080
| 10.7 | Height adjustable operator keyboard
| 10.8 | Integrated HDD (500 GB)
| 10.9 | Minimum 8 TGC pods
| 10.10 | Max. number of focal points min. 8, probe and application dependent
| 10.11 | Coded harmonic imaging
| 10.12 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 10.13 | Method of spatial compouding of the image from several different angles
| 10.14 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 10.15 | Full alphanumeric keypad covered with washable protection film
| 10.16 | Maximum depth of field min. 30 cm, probe dependent
| 10.17 | Technique which provides a convex field of view on linear and sector transducers
| 10.18 | HD write zoom - read zoom min. 50X
| 10.19 | Voice comments recording
| 10.20 | Transducer type:
| 10.21 | Convex array
| 10.22 | Linear array
| 10.23 | Microconvex array
| 10.24 | Sector phased array
| 10.25 | Volume Probe (4D)
| 10.26 | Operating modes:
| 10.27 | B-Mode
| 10.28 | M-Mode
| 10.29 | Color M-Mode
| 10.30 | PW Doppler with high PRF
| 10.31 | Color Flow mode (CFM)
| 10.32 | Power Doppler Imaging (PDI) and Directional PDI
| 10.33 | Inputs and outputs:HDMI output (1920 x 1080 resolution), VGA output (1920 x 1080 resolution), CVBS output, S-Video output, 4 x USB, Ethernet (RJ45)
| 10.34 | B/W thermal printer
| 10.35 | Probes:
| 10.36 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 10.37 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
11 |  | Ultrasound system for pediatric cardiology |  | 1 |  | 0
| 11.1 | Ultrasound system, not portable, with wheels and front and rear handles
| 11.2 | Minimum 4 active probe ports
| 11.3 | Composite dynamic range min 400 dB
| 11.4 | Max. frame rate 1000 fps, probe and mode dependent
| 11.5 | Minimum 21" wide-screen, High-Definition (HD), flicker-free LCD display. Resolutions min. 1920 x 1080, with adjustable position and tilt and digital brightness adjustment. Articulating monitor arm.
| 11.6 | Minimum 10" ultra-high-resolution, wide-screen format, color, multi-touch LCD screen with Interactive user-configurable dynamic software menu
| 11.7 | Alphanumeric keyboard
| 11.8 | Integrated HDD min. 500 GB
| 11.9 | Six TGC pods, automatically calculated lateral gain
| 11.10 | Maximum depth of field min. 30 cm, probe dependent
| 11.11 | Harmonic Imaging
| 11.12 | An advanced image processing technique to help reduce speckle in real time examining the relative difference between neighboring pixel values and determining whether the grayscale variations have a sharp difference, follow a trend, or are random in nature
| 11.13 | Spatial compounding of the image from several different angles
| 11.14 | Automatic tissue optimization
| 11.15 | The system provides an extended field of view on sector probes in order to better visualize structures near the probe (near field),
| 11.16 | Auto-spectral optimize –adjustments baseline, PRF (on live image) and angle correction
| 11.17 | Battery standby
| 11.18 | Transducer type:
| 11.19 | Sector phased array
| 11.20 | Convex array
| 11.21 | Linear array
| 11.22 | CW Pencil probe
| 11.23 | TEE probe
| 11.24 | Intraoperative probe
| 11.25 | Operating modes:
| 11.26 | 2D tissue (B-mod)
| 11.27 | Tissue M-mode
| 11.28 | Color M-mode
| 11.29 | Tissue velocity M-mode
| 11.30 | 2D color flow, CFM
| 11.31 | Anatomical M-mode
| 11.32 | PW Doppler with HPRF with maximum sample volume depth of min.30 cm
| 11.33 | CW Doppler, CW
| 11.34 | 2D angio flow (Power Doppler)
| 11.35 | Tissue velocity imaging (TVI)
| 11.36 | Tissue Doppler imaging (TDI)
| 11.37 | Hardware and software options which can be added later:
| 11.38 | A function that enables automatic Doppler cardiac measurements
| 11.39 | A digital imaging technique that provides real-time visualization of vascular hemodynamics by directly visualizing blood reflectors and presenting this information in a grayscale display as a B-flow, Sieflow or corresponding
| 11.40 | Parametric imaging tool which gives quantitative data for global and segmental strain which allows comprehensive assessment at a glance by combining three apical longitudinal views into one comprehensive bulls-eye view, with integrated auto EF calculation.
| 11.41 | Archiving and connections:
| 11.42 | CINE memory minimum 500 MB
| 11.43 | Export formats: JPEG, MPEG, AVI
| 11.44 | Export of alphanumeric data in MS Excel format
| 11.45 | Ethernet network connection (RJ45)
| 11.46 | VGA video output (min. 1080 resolution), HDMI video output, audio stereo output, min. 4 USB
| 11.47 | Possibility of connection to a PC with installed software for review and advanced postprocessing analysis of data from the ultrasound device in the original format (raw data)
| 11.48 | B/W thermal printer
| 11.49 | Probes:
| 11.50 | 2D Sector pediatric probe Frequency Range: 2,0 – 7,0 MHz or wider Maximum FOV: min. 120°
| 11.51 | 2D Sector probe Frequency Range: 1,5 – 4,0 MHz or wider Maximum FOV: min. 120°
12 |  | Ultrasound system for adult cardiology |  | 1 |  | 0
| 12.1 | Ultrasound system, not portable, with wheels and front and rear handles
| 12.2 | Minimum 4 active probe ports
| 12.3 | Composite dynamic range min 400 dB
| 12.4 | Max. frame rate 1000 fps, probe and mode dependent
| 12.5 | Minimum 21" wide-screen, High-Definition (HD), flicker-free LCD display. Resolutions min. 1920 x 1080, with adjustable position and tilt and digital brightness adjustment. Articulating monitor arm.
| 12.6 | Minimum 10" ultra-high-resolution, wide-screen format, color, multi-touch LCD screen with Interactive user-configurable dynamic software menu
| 12.7 | Alphanumeric keyboard
| 12.8 | Integrated HDD min. 500 GB
| 12.9 | Six TGC pods, automatically calculated lateral gain
| 12.10 | Maximum depth of field min. 30 cm, probe dependent
| 12.11 | Harmonic Imaging
| 12.12 | An advanced image processing technique to help reduce speckle in real time examining the relative difference between neighboring pixel values and determining whether the grayscale variations have a sharp difference, follow a trend, or are random in nature.
| 12.13 | Spatial compounding of the image from several different angles
| 12.14 | Automatic tissue optimization. Automatic Optimization or corresponding
| 12.15 | The system provides an extended field of view on sector probes in order to better visualize structures near the probe (near field), Virtual Apex or coresponding
| 12.16 | Auto-spectral optimize –adjustments baseline, PRF (on live image) and angle correction
| 12.17 | Battery standby
| 12.18 | Transducer type:
| 12.19 | Sector phased array
| 12.20 | Convex array
| 12.21 | Linear array
| 12.22 | CW Pencil probe
| 12.23 | TEE probe
| 12.24 | Intraoperative probe
| 12.25 | Operating modes:
| 12.26 | 2D tissue (B-mod)
| 12.27 | Tissue M-mode
| 12.28 | Color M-mode
| 12.29 | Tissue velocity M-mode
| 12.30 | 2D color flow, CFM
| 12.31 | Anatomical M-mode
| 12.32 | PW Doppler with HPRF with maximum sample volume depth of min.30 cm
| 12.33 | CW Doppler, CW
| 12.34 | 2D angio flow (Power Doppler)
| 12.35 | Tissue velocity imaging (TVI)
| 12.36 | Tissue Doppler imaging (TDI)
| 12.37 | Archiving and connections:
| 12.38 | CINE memory minimum 500 MB
| 12.39 | Export formats: JPEG, MPEG, AVI
| 12.40 | Export of alphanumeric data in MS Excel format
| 12.41 | Ethernet network connection (RJ45)
| 12.42 | VGA video output (min. 1080 resolution), HDMI video output, audio stereo output, min. 4 USB
| 12.43 | Possibility of connection to a PC with installed software for review and advanced postprocessing analysis of data from the ultrasound device in the original format (raw data)
| 12.44 | B/W thermal printer
| 12.45 | Probes:
| 12.46 | 2D Sector probe Frequency Range: 1,5 – 4,0 MHz or wider Maximum FOV: min. 120°