## Price schedule
Bidder:
LOT 4 |  | Radiology equipment |  |  | Date:
Total price:
Line item No. | ID | Equipment name | Technical Specification Offered (model) | QTY | Unit price | Total Price per line item
1 |  | CT Scanner |  | 1 |  | 0
2 |  | Digital mamography system |  | 1 |  | 0
3 |  | Mobile digital X Ray unit with C-Arm and FPD |  | 1 |  | 0
4 |  | PACS- PICTURE ARCHIVING AND COMMUNICATION SYSTEM |  | 1 |  | 0
5 |  | Digital X-ray machine (ceiling, 2 flat panels) |  | 1 |  | 0
6 |  | Ultrasound system for the gastro department |  | 1 |  | 0
7 |  | Ultrasound system for use in Radiology |  | 2 |  | 0
8 |  | Ultrasound system for use in Ginecolgy |  | 1 |  | 0
9 |  | Portable ultrasound system for use on intensive care |  | 1 |  | 0
10 |  | Portable ultrasound system for use in Ginecolgy |  | 1 |  | 0
11 |  | Portable ultrasound system for use in General surgery |  | 1 |  | 0
12 |  | Portable ultrasound system for use in Cardiology |  | 1 |  | 0
13 |  | Portable ultrasound system for use in Orthopedics |  | 1 |  | 0
TOTAL Radiology equipmet | [5 empty] | 0

## Specifications
Bidder:
LOT 4 |  | Radiology equipment |  |  | Date:
Total price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | CT Scanner |  | 1 |  | 0
| 1 | HIGH VOLTAGE GENERATOR
| 2 | Max Generator hardware power: at least 72 kW
| 3 | mA range: at least 20-460 mA or more (real, not equivalent values), with 5mA increments
| 4 | kV range for radiography: 80 - 140 kV or wider
| 5 | X-RAY TUBE
| 6 | Number of tube focuses, at least two
| 7 | Real heat capacity of the tube anode (not equivalent value), minimum 7 MHU or X-ray tube cooling rate: min 2.7 MHU/min
| 8 | X-ray tube anode cooling rate: min 1 MHU/min
| 9 | RADIOGRAPHIC PATIENT TABLE
| 10 | Patient table load capacity: at least 200 kg
| 11 | Horizontal scanning range min. 200cm
| 12 | Coronal imaging head holder.
| 13 | Hand holder for easier patient positioning
| 14 | GANTRY
| 15 | Aperture min. 70cm
| 16 | Shortest time for one full scan rotation: max. 0,35sec not only for Cardiac
| 17 | DETECTOR SYSTEM
| 18 | At least 64 detector rows
| 19 | At least 128 reconstruction layers (slice) per rotation
| 20 | Minimum layer (slice) thickness: max. 0.625mm
| 21 | Spatial resolution: max 0.31mm
| 22 | High Contrast Resolution not less than 12 lp/cm (X-Y-Plane at 50%MTF)
| 23 | ACQUISITION PARAMETERS / ACQUISITION WORKSTATION - OPERATING CONSOLE
| 24 | Field of view for scanning min. 50cm
| 25 | Pitch factor of 0.20-1.50, selectable and applicable for scanning with full FOV of min. 50cm
| 26 | Trigger for acquisition when bolus arrives to ROI
| 27 | Real Time MPR reconstruction
| 28 | CT brain 4D perfusion software package
| 29 | Iterative reconstruction technology that allows dose reductions with the same diagnostic image quality (iDOSE, SAFIRE, ASIR, AIDR 3D or equivalent)
| 30 | CT viewer with tools for fast 3D postprocessing (3D volume rendering, MIP, MPR, VRT)
| 31 | Advanced non-invasive CT angiography (CTA) analysis tools with automatic bone removal
| 32 | System for correction of artifacts from the presence of metals in the body (not Dual Energy)
| 33 | Cardiology study package (at least ECG gated calcium scoring, ECG waves display on console and gantry)
| 34 | DICOM protocol for network print, sent, query/retrieve
| 35 | DICOM Media study archive on CD/DVD with DICOM viewer
| 36 | Injector class IV synchronization interface
| 37 | Image Reconstruction speed: min 55 images per second
| 38 | Technology applicable for Low Dose Screening of lung cancer and high-risk patients
| 39 | Extended range for 4D perfusion scans: min 140mm coverage
| 40 | System must have package for acquisition synchronized with ECG signal (hardware and software)
| 41 | ECG Monitor synchronized with CT system for cardiac gating
| 42 | High performance console host computer with minimum 36GB RAM and minimum 8 CPU cores
| 43 | Diagnostic Server Workstation characteristics:
| 44 | Client-server architecture with three (3) client workplaces
| 45 | Simultaneous operation of at least 3 clients/workplaces for advanced post processing.
| 46 | Client computers with monitors (2 monitors per workplace) with screen diagonal of at least 24 inches
| 47 | Dynamic licenses distribution between clients (floatable licenses)
| 48 | Server image database storage, at least 2 TB
| 49 | Remote server access with FULL ability to analyze and review studies
| 50 | Server must be equipped with the following permanent (unlimited time of use) licenses:
| 51 | - 3D anatomy analysis and display tools (3D volume rendering, MIP, MPR, VRT), at least three licenses
| 52 | - Advanced non-invasive CT angiography (CTA) analysis tools with automatic bone removal, at least one license
| 53 | - Software for assessment and automatic visualization and monitoring of pulmonary nodules and diagnosis -CT LNA (lung nodule analysis), at least one license
| 54 | - Advanced tools for analysis of lung disease (e.g. chronic obstructive pulmonary disease), at least one license
| 55 | - CT brain 4D perfusion software package, at least one license
| 56 | - Oncology package for automatic multimodal tumor monitoring, automatic calculation following WHO criteria, RECIST 1.1., at least one license
| 57 | - Cardiology study package ( Scoring calcifications of coronary blood vessels, cardiac calcium plaque burden quantification and coronary artery disease risk assesment) at least one license.
| 58 | - Advanced tools for analysis of neurological exams (CT brain perfusion, stroke assessment and analysis of aneurysms), at least one license
| 59 | CT INJECTOR
| 60 | Dual head CT injector, with 1 heads for contrast media and 1 head for saline solution. Possesion of air monitoring with sensors or LED lights for easy visual detection of air bubble presence.
| 61 | Personal protective equipment:
| 62 | Lead apron, one-piece |  | 6
| 63 | Double-sided protection
| 64 | The front of the apron is equivalent to 0.5 mm lead
| 65 | The back of the apron is equivalent to 0.25 mm lead
| 66 | The supplier is obliged to harmonize the size of the offered protective equipment with the end user before delivery
| 67 | Lead collar |  | 6
| 68 | Thyroid protection collar
| 69 | Easy fastening with buckles with the possibility of adjusting the girth
| 70 | Lead equivalent 0.5 mm
| 71 | The supplier is obliged to provide the size of the offered protective equipment before delivery agree with the end user
| 72 | Protective goggles |  | 4
| 73 | Goggles that protect the eyes from radiation from all angles
| 74 | Front side equivalent to 0.5 mm lead
| 75 | Sides equivalent to 0.5 mm lead
| 76 | Anti-reflective coating
| 77 | The supplier is obliged to provide the size of the offered protective equipment before delivery agree with the end user
| 78 | Electronic dosimeter |  | 2
| 79 | For staff exposed to X radiation
| 80 | Audible, visual and vibrating alarms to warn users of immediate overdose
| 81 | Software that allows to record and monitor user exposure to radiation
| 82 | Built-in rechargeable battery, charging via USB
| 83 | Energy range from 20 keV to 10 MeV
2 |  | Digital mamography system |  | 1 |  | 0
| 1 | Generator
| 2 | High frequency generator
| 3 | kW min. 5 kW
| 4 | kV range: min. 23-49 kV, with 1 kV steps
| 5 | Max mAs: Min. 500 mAs
| 6 | Tube and gantry
| 7 | Anode heat capacity: min. 300 kHU
| 8 | Focal spot size (IEC 60336) Max 0.1 mm and 0.3 mm
| 9 | AOP - KV , mAs , filter
| 10 | Specific paddle and AEC for implants
| 11 | Automatic collimation set to paddle size
| 12 | Motorized gantry vertical movement min. 70,5-140 cm
| 13 | Gantry motorized rotation min. -140/+180 degrees
| 14 | SID min. 65 cm
| 15 | Bucky thickness from chest wall side max. 85 mm
| 16 | 2 set of footswitches on both side of the gantry
| 17 | Compression paddles min. 5 different shapes
| 18 | Motorized and manual compression with numeric display
| 19 | Auto decompression after exposure
| 20 | Magnification factors 1.8
| 21 | Faceshield
| 22 | Acquisition workstation
| 23 | Workstation with shielding
| 24 | Enter data manually or via DICOM worklist
| 25 | Repeat/Reject analysis
| 26 | Special processing algorithm for implant imaging
| 27 | Possibility of individual work under the user's credentials (name and password)
| 28 | Touch screen console
| 29 | PC memory min. 8 GB RAM
| 30 | HDD min. 1 TB
| 31 | 3MP LCD monitor min. 21”
| 32 | DICOM 3.0 (Send / Receive, Query / Retrieve, Storage Commitment, Print, Worklist, MPPS)
| 33 | UPS
| 34 | Flat panel digital detector
| 35 | Detector size min. 24 x 29 cm
| 36 | CsI or aSe
| 37 | Pixel size (2D and 3D mode) max. 100 µm
| 38 | DQE@ 0,5 lp/mm min. 62%.
| 39 | Diagnostic workstation
| 40 | Two medical TFT monitors, min. 21-inch, resolution min. 5MP or one min. 10MP monitor approved for diagnostics.
| 41 | Workstation with dedicated software for 2D
| 42 | LCD monitor for patient administration min. 19”
| 43 | Memory min 16 GB RAM
| 44 | SSD min. 512GB and additional HDD min. 1. TB or HDD min. 2TB
| 45 | Special programmable mammography keypad
| 46 | Database search by multiple criteria, patient name, patient ID, type of modality, study date, etc.
| 47 | Image processing tools: contrast, brightness, inversion, zoom, rotation, annotations, measurements
| 48 | Support for other modalities, US, MRI
| 49 | Hanging Protocols
| 50 | Worklist
| 51 | Clinical education
| 52 | Access to online educational material
| 53 | Upgradability
| 54 | Tomosynthesis
| 55 | Imaging technique that allows to use mammogram with iodinated contrast agent, contrast enhanced mammography
| 56 | Tomo guided biopsy
| 57 | Upgradeability to contrast-enhanced stereo biopsy
3 |  | Mobile digital X Ray unit with C-Arm and FPD |  | 1 |  | 0
| 1 | C-ARM
| 2 | SID min. 102cm
| 3 | Orbital rotation min. 160°
| 4 | Axial rotation min. ± 220°
| 5 | Swiveling ± 10°
| 6 | Foot switch pedal and hand switch for exposition.
| 7 | Touch screen user interface integrated on C-arm.
| 8 | Live image display on “Touch screen” user interface integrated on C arm
| 9 | X RAY GENERATOR
| 10 | Maximum output power min. 2.3 kW
| 11 | kV range for pulsed fluoroscopy and digital radiography min. 40-120 kV
| 12 | mA range for pulsed fluoroscopy min. 3mA - 20 mA
| 13 | mA range for digital radiography up to 24 mA or more
| 14 | Programmable anatomical programs
| 15 | Automatic adjustment of current and voltage during fluoroscopy
| 16 | Automatic correction of metal artifacts
| 17 | X RAY TUBE AND COLIMATION
| 18 | Single or dual focus
| 19 | X ray tube heating capacity min. 1 MHU
| 20 | Collimator rotation ± 90°, iris and asymetric slot collimator
| 21 | Virtual collimation without radiation
| 22 | DAP (Dose Area Product)
| 23 | OPERATING MODES
| 24 | Pulsed fluoroscopy
| 25 | Digital radiography
| 26 | Program for: automatic motion detection, automatic dose reduction, automatic metal correction, automatic pulse rate adjustment in order to prevent tube overheating
| 27 | IMAGE RECEIVER
| 28 | FPD - Flat Panel Detector
| 29 | Active detector size min. 20cm x 20cm
| 30 | Detector matrix min. 2.000 x 2.000 pixels
| 31 | Grayscale min. 16 bit
| 32 | Pixel size max. 100 µm
| 33 | Integrated laser positioning device on FPD
| 34 | Minimum 3 fields of view
| 35 | DIGITAL MODULE
| 36 | 19” duo faltscreen monitor, with monitor support arm integrated on C arm
| 37 | LIH (Last Image Hold)
| 38 | Image archive on integrated HDD min. 100.000 images
| 39 | Image archive on USB in min. DICOM, JPEG or TIFF and AVI format, with DICOM viewer software
| 40 | Zoom, grayscale inversion, noise reduction filter, edge enhancement, Digital image rotation, measurement function
| 41 | Min. 16 image mosaic display
| 42 | DICOM Package min.: Storage (incl. Comm.), Worklist (incl. MPPS), Query/Retrieve, DICOM Viewer Software with auto run for USB.
| 43 | ELECTRICAL REQUIREMENTS
| 44 | Single phase plug 220-240V/50-60 Hz
4 |  | PACS- PICTURE ARCHIVING AND COMMUNICATION SYSTEM |  | 1 |  | 0
| 1 | IMAGE VIEWER FUNCTIONS
| 2 | DIAGNOSTIC CLIENT FOR IMAGE REVIEW (MIN. 2 CONCURENT USERS)
| 3 | Configurable search/sort criteria ( like patient name,patient ID, modality,date,etc. )
| 4 | On worklists, jokers are supported as criteria for search in all fields( like „Neuro *“, „* Thoraks“)
| 5 | Search is enabled by several criteria simultaneously
| 6 | User can create and modify his own worklists
| 7 | Defined worklist can be controlled centrally and assigned to different groups and/or users
| 8 | User interface can be customized to an user and remembered as such
| 9 | Display of different informations ( patient name, patient ID, modality, date, description, etc.. ) is configurable in a work list
| 10 | Sorting in a worklist can be done alphabetically or numerically
| 11 | List can be sorted on at least three levels: study date ,study period , modality
| 12 | Quick swithch to a next study is possible by pressing one button only
| 13 | Defining atributs-diagnostically relevant images
| 14 | Defining atributs-Images which are to be saved
| 15 | Defining atributs-new imported studies
| 16 | Defining atributs-processed studies
| 17 | Defining atributs-completed finding
| 18 | The ability to display patient's history/study
| 19 | Displaying of current and previous studies (via Drag&Drop) in any desired combination
| 20 | Automatic layout of image series by using predefined hanging protocols
| 21 | Hanging protocols can be defined per user/examination type by the user via user interface
| 22 | The Loading of hanging protocols and the arrangement of the series can be controled by any DICOM header entries
| 23 | For different exam type several hanging protocols can be linked to one chain in a freely definable order and can be enabled via key button
| 24 | Automatically load appropriate preliminary examinations via Hanging protocol
| 25 | Customizing multiple Hanging protocols simultaneously
| 26 | Simultaneously Display of all DICOM-compliant images (e.g., CT, MR, CR, DX, XA, US, NM, OT, SR, PT, PDF)
| 27 | Cine and video display
| 28 | List of the last displayed studies
| 29 | Display of all available series in one view
| 30 | Different schedules are supported 1x1,1x2,2x2,2x3,3x2,3x3,1x4,2x4
| 31 | Simple window size adjustment
| 32 | Regular checking if new patient data has arrived
| 33 | Cutline can be displayed on the correlating levels
| 34 | The user-specific tool bar is configurable
| 35 | Tool pallete for each modality can be configured
| 36 | Quick access to the most important tools is enabled via shortcut
| 37 | Interactive Window level & Zoom & Pan & navigation function on the mouse without changing the mouse button setting
| 38 | Recording of settings for L/W preset
| 39 | Recording L/W level for specific modality
| 40 | Recording L/W level for specific exam ( soft tissuses,bones,etc…)
| 41 | The ability to remember personalized settings of L/W
| 42 | Automatic adjustment of L/W
| 43 | L/W reset
| 44 | Interactive zoom and pan zoom
| 45 | Continuous zoom
| 46 | Simultaneous summation of a group of displayed images
| 47 | Free image rotation ,mirroring and flipping
| 48 | Edge enhacement filter
| 49 | Gray scale inversion
| 50 | Automatic cine mod control
| 51 | Cine mode can be controlled manually by moving the mouse (wheel) with a variable frame rate or in one step
| 52 | L/W adjustment in cine mod
| 53 | View the image in full screen at the touch of a button or by double-clicking
| 54 | Distance measurement
| 55 | Lateral pelvic tilt
| 56 | Automatic density measurement based on cursor tip position
| 57 | Density measurement within a ROI (circle and asymmetric)
| 58 | Manual calibration of pixel size
| 59 | Download image to windows clipboard using copy and paste
| 60 | Export of the images into different data formats (at least: jpg, dicom, bmp, avi) possible
| 61 | Normal paper print and DICOM Print (Basic Greyscale Print Management)
| 62 | Storage of image states as additional series (screenshot)
| 63 | Simultaneous, anatomic synchronized browsing through multiple series on multiple monitors (also on different layer thickness and different start positions)
| 64 | Displaying any patient CD / DVD with DICOM images on the same display
| 65 | Multiplanar reconstruction MPR
| 66 | Maximum Intensity Projection (MIP)
| 67 | Minimum Intensity Projection (MIN)
| 68 | Storing of MPR as a new series within the PACS
| 69 | Integration of 3D functions within defined display frames
| 70 | Cursor position synchronization in all series including 3D, VRTs, MPR
| 71 | Cursor mode for slices with correlation position notification at all levels (3D cursor)
| 72 | Spine marking
| 73 | Three and four points angle measurement
| 74 | Goniometry
| 75 | Temporary "frozen" view ( no screenshot)
| 76 | Display dimming
| 77 | Time navigation in dynamic studies (CT / MRI) with one mouse / button click
| 78 | Time navigation of dynamic studies configurable according to different parameters
| 79 | Perfusion curve measurment
| 80 | ADVANCED DIAGNOSTIC CLIENT - ADDITIONAL FUNCTIONALITIES IN RELATION TO THE DIAGNOSTIC CLIENT (MIN. 2 CONCURENT USERS)
| 81 | Volume display (VTR) available
| 82 | Archiving of VRT as a new 360°-view series in the PACS
| 83 | Integration of Volume rendering within the Hanging protocol function
| 84 | Integration of vessels analysis functions into Hanging Protocol
| 85 | Cutting function (scalpel) within VRT available
| 86 | Displaying of synchronized 3D reconstructions (VRTs) on min. two sets of data is possible
| 87 | Curved MPR is availible
| 88 | Manual registration of images of multiple studies of different recording modalities to allow comparison of reports for the same body part
| 89 | Image fuzion ( eg. PET/CT )
| 90 | OS AND HARDWARE REQUIREMENTS
| 91 | At least 3 monitors are supported by the diagnostic workstation
| 92 | Multiple monitor formats / resolutions are supported
| 93 | The PACS software is hardware (manufacturer) independent
| 94 | The PACS software is independent of the Windows version
| 95 | 3D reconstructions can be rendered by use of GPU (graphic board)
| 96 | WEBVIEWER- AVAILABLE FOR AT LEAST 10 CONCURRENT USERS
| 97 | BASIC FUNCTIONS
| 98 | Zero Footprint viewer (no installation necessary)
| 99 | Viewer based on HTML5
| 100 | Web image distribution is realized via standardized interface functions (like Ethernet TCP/IP, HTML, HTTPs, JPEG/Wavelet compression)
| 101 | Web image distribution is browser independent
| 102 | Image distribution via Web client is available by using predefined compression rates to reach higher transfer speed
| 103 | It is possible to reload the images in high quality for diagnistic
| 104 | All requests for a web server are recorded in accordance with data protection
| 105 | WEBVIEWER - SEARCH CRITERIA
| 106 | Configurable search criteria for the worklist (like patient name, patient ID, modality, date, requesting department) are available
| 107 | Wild-card character will be supported for all search-criteria fields (like "Neuro*", "*department*")
| 108 | Previous studies can be displayed anytime
| 109 | Selection of multiple studies of one patient is possible
| 110 | The administrator must be able to manage worksheets centrally
| 111 | Sorting functions for the worklist
| 112 | Worklist specific configuration for the users
| 113 | WEBVIEWER - DISPLAY FUNCTION
| 114 | Existing overview (thumbnails) of selected images, series and studies
| 115 | Image formation via Drag&Drop is possible
| 116 | Displaying of old and new images and previous images
| 117 | The cross-section of the cut line can be displayed at the appropriate levels
| 118 | Two series of images on the same monitor
| 119 | Cinemod
| 120 | Display Multiframes and video (e.g. MP4)
| 121 | Displaying of black/white images
| 122 | Displaying of color images
| 123 | Displaying of .pdf documents
| 124 | Displaying of ECG image data (not in .pdf format!)
| 125 | Displaying of written reports from RIS
| 126 | DICOM Send availible
| 127 | WEBVIEWER - TOOLS
| 128 | Fast access to the most important tools (scroll, W/L, PAN, Zoom)
| 129 | Interactive Window/Level/ Zoom/PAN setup
| 130 | Point-to-point distance measurement
| 131 | Angle calculation
| 132 | Inverting of images is possible
| 133 | ECG measurement
| 134 | ROI marking
| 135 | Consultation workflow (Request and Reply)
| 136 | WEBVIEWER - EXPORT
| 137 | Transfer the images to the Windows clipboard using Copy & Paste
| 138 | DICOM export
| 139 | Printing on paper
| 140 | PACS BASIC CONFIGURATION
| 141 | BASIC
| 142 | Highly available PACS Server
| 143 | Possibility to connect at least 5 DICOM modalities
| 144 | Server can be restarted at any time without loss of data
| 145 | The PACS servers do not need dongles or licensing hardware
| 146 | The offered PACS is modular, scalable and the individual processes can be distributed to server systems
| 147 | The PACS has a DICOM compliant data management
| 148 | The PACS is multi-client capable. All data will be centralized archived. The access is done in client-based areas and is related to the user-rights of the client
| 149 | Parallel processing of several on-demand jobs is possible
| 150 | A study can be performed simultaneously at several workstations
| 151 | At the same time several studies can be considered performant in one workplace
| 152 | The response time for loading images and series "On demand" allows fast processing (single images <2s, series <10s)
| 153 | The PACS communicates easily with existing modalities, workstations, readers and dry imagers via DICOM
| 154 | Central administration of individual workplace configurations, server and workflow settings is possible
| 155 | The connection of modalities by administrators is possible after training without additional service
| 156 | Admin functions can be divided by permissions (e.g., image mapping by MTRA's)
| 157 | Easily report problems at the touch of a button
| 158 | PACS STORAGE
| 159 | PACS includes storage management (internal such as Raid5 or external such as SAN, NAS, Fast LTA, ...)
| 160 | The option of a modular extension of the storage capacity (online availability of image data) has to be supported
| 161 | Any kind of storage media (such as DAS, SAN, NAS, SDD-HD, USB-Disk) have to be supported
| 162 | Storage media can be partitioned by the PACS internally
| 163 | Study-level related attributes can be set (such as verified, external images, teaching archive, temporarily)
| 164 | Configurable prefetching of Studies
| 165 | Data have to be stored by "first-in-first-out" regulations
| 166 | Images from all storage locations can be displayed simultaneously on all workstations (diagnostic workstation, review workstation, web workstation).
| 167 | The initial images from the study cannot be changed. Subsequent changes to the DICOM header, notes, etc. They are enrolled in the DICOM studio
| 168 | At the administrator level, studies, series and images can be shared, summarized, "redistributed" and manually deleted.
| 169 | DATABASE
| 170 | Database is taking care of the storage-location of the image data as well as of all stored parameters of the studies within the PACS
| 171 | Storage of all demographic patient data
| 172 | Storage of all examination-related patient data
| 173 | Storage of all annotations
| 174 | Storage of all Window/Level changes
| 175 | Storage of all comments
| 176 | Storage of all user-rights and user-profiles
| 177 | Linux OS is supported
| 178 | Reading-access on the used database have to be possible by using standard SQL tools
| 179 | IMAGE COMPRESSION
| 180 | A lossless data compression is available
| 181 | A lossless data compression for the centralized archiving of image data has to be configurable
| 182 | Compression rates are configurable
| 183 | On-demand data transfer and data decompression in the diagnostic station
| 184 | DICOM
| 185 | Supported DICOM standard for communication within PACS and with related systems and modalities
| 186 | DICOM-Arhiva
| 187 | DICOM-Storage Commitment
| 188 | DICOM-Query & Retrieve
| 189 | DICOM-Print
| 190 | DICOM-Worklist management
| 191 | Data sent by MPPS (such as dose protocol) must be stored in the PACS in connection with the review
| 192 | The DICOM nodes are freely defined and configured by the administrator.
| 193 | It is possible to supplement the missing DICOM fields with information from the HL7 RIS data
| 194 | NON DICOM FILES
| 195 | No DICOM objects are supported
| 196 | DICOM-Worklist Management for Non-DICOM Converter
| 197 | Missing DICOM information can be filled in with HL7-RIS data information
| 198 | QUALITY ASSURENCE
| 199 | PATIENT DATA / IMAGE DATA ADMINISTRATION
| 200 | Structured quality assurance - harmonization of incoming image studies in the archive with orders from the patient administration system. Orders can be accepted via the HL7 Gateway and the DICOM worksheet protocol.
| 201 | Application administration (e.g., patient data change, study assignment) is possible with particular authorization
| 202 | Studies, series, images can be splitted on the admin level
| 203 | Studies, series, images can be merged on the admin level
| 204 | Studies, series, images information can be edit on the admin level
| 205 | Studies, series, images can be manually deleted at the Admin level
| 206 | INTEGRATION
| 207 | HL7 protocols, IHE-compatible interface for all leading information systems - 1 license. Through this module, data communication with clinical information systems is realized. Areas of application include order and report communication, as well as data consistency protection. The following messages are supported: HL7 ADT, HL7 ORM, HL7 ORU, HL7 MDM.
| 208 | The system supports communication with min. with one independent PACS solution. The protocols allow connection to independent archiving systems compliant with DICOM. This allows you to import data from independent PACS solutions or connect such solutions online as departments.
| 209 | PACS can capture image data from different image processing system manufacturers and integrate it as digital data into PACS.
| 210 | Connect-a-thon integration profiles tested
| 211 | One-time registration is supported for HIS and PACS
| 212 | The HIS client can connect to a PACS diagnostic workstation or review workstation, using a shared keyboard and mouse
| 213 | Patient and examination data will be automatically submitted from HIS to PACS (no additional choice)
| 214 | IMAGE TRANSFER/EXPORT/IMPORT (MIN. 1 LICENCE)
| 215 | Imported studies (on CD / DVD) can be temporarily saved in the PACS archive and deleted after a certain period of time
| 216 | Patient studies received by image transmission can be stored temporarily in the PACS archive and will be deleted after a defined period of time
| 217 | After the import of patient studies a prefix can be added automatically to the study description
| 218 | Patient studies marked with a worksheet function will not be automatically deleted
| 219 | Import image data verification service available (verified and unverified data)
| 220 | Imported image data can be distributed by image distribution without long-term archiving
| 221 | Import non-DICOM image formats such as digital images, endoscopic images or sonographic images
| 222 | Collecting of non-DICOM video data
| 223 | Import non-DICOM live video sources
| 224 | Obtaining data about non-DICOM images, such as pdf files
| 225 | All non-DICOM data will be converted into DICOM format and can be distributed to all clients
| 226 | All studies can be selected by using defined search criteria to be burned on patient CD/DVD including a ready-to-run DICOM viewer
| 227 | Image data and related reports can be burned to a CD / DVD as a whole
| 228 | Image data export in TIFF/JPEG, BMP and AVI formats is available
| 229 | Automatic file import service (pdf data, JPEG data, etc.)
| 230 | FILM/PAPER PRINTING
| 231 | The system has a printer interface, which is operated with the current printer drivers of current standard printers
| 232 | Printing on paper laser printers is possible
| 233 | Printing on hardcopy (film) printers is possible
| 234 | Printing on standard paper printers is possible
| 235 | IMPORT FUNCTIONS
| 236 | Import image and preview data
| 237 | Query on DICOM worksheet
| 238 | Import from other systems (based on DICOM)
| 239 | Import from hard drive
| 240 | Import from CD/DVD
| 241 | Import from other data media
| 242 | Import DICOM data format
| 243 | EXPORT FUNCTIONS
| 244 | Export images/examination data to a local hard drives
| 245 | Export to another server (based on DICOM) via DICOM sending or automatic routing
| 246 | Export to a hard drive
| 247 | Export to CD/DVD
| 248 | Export via DICOM
| 249 | Export via JPEG
| 250 | Export in PDF format
| 251 | Objects can be anonymized before export, including the image header
| 252 | HARDWARE/SOFTWARE REQUIREMENTS
| 253 | Highly available PACS Server with the following performance: CPU: Intel Core i3 10th generation; min. 16 GB RAM DDR4; 4x250GB SSD (RAID 6); 4x5TB (RAID 6); Win 10 pro or appropriate.
| 254 | Computer for diagnostic workstation -2 pcs, with the following performance: CPU: I3 9 / 10gen 4cores or equivalent, MB: 1GB Lan, RAM: 16GB DDR4, HDD: SSD 120GB, DVD: DVDRW, VGA: Discrete 2-4GB GDDR5, possibility of connection with 3 monitors: 1 VGA HD resolution and 2 diagnostic 2MP monitors with DVI-D and port input display, Mouse, keyboard, CASE: MIDI 500W, MON: 1x21", OS: WIN 10 64bit
| 255 | SYSTEM
| 256 | TCP/IP as basic communication standard
| 257 | Ethernet, TP-100/1000 Mbit/s
| 258 | VIRTUALIZATION
| 259 | The PACS system has been released for VMware ESX
| 260 | The PACS system runs smoothly on VMware ESX and works reliably
| 261 | CLIENT
| 262 | Updates, fixes, patches, or hotfixes are loaded and installed automatically when you restart your computer
| 263 | Customer-independent hardware can be used for customers
| 264 | PACS programs can run at the same time as other utilities, such as HIS, RIS, MS Office, etc.
| 265 | 2MP DIAGNOSTIC MEDICAL MONITOR- 4 PCS
| 266 | Color medical monitor, TFT IPS technology diagonal min. 21 ″, min. 2MP.
| 267 | Resolution minimum 1200x1600
| 268 | Luminance minimum 500cd/m²
| 269 | Contrast ratio mon. 1800:1
| 270 | Viewing angle: At least 175° vertical and horizontal
| 271 | Ambient light detection sensor to compensate for optimal calibration
| 272 | Front sensor for light and color stabilization
| 273 | Human Presence built-in Sensor - Automatically enters power save mode when the sensor does not detect a human presence
| 274 | Automatic color image recognition to ensure optimal contrast and brightness
5 |  | Digital X-ray machine (ceiling, 2 flat panels) |  | 1 |  | 0
| 1 | GENERATOR
| 2 | kV range 40-150kV or wider.
| 3 | mA range 10-640mA or more.
| 4 | mAs range 0.5 mAs to 500 mAs or more.
| 5 | Exposure time-range 0.002 sec. up to 2 sec. or wider
| 6 | Automatic exposure (AEC) minimum 3 fields.
| 7 | Power supply - Three-phase.
| 8 | CEILING MOUNTED X-RAY TUBE
| 9 | Motorized or manual longitudinal movement of X-ray tube support minimum 4000mm.
| 10 | Motorized or manual transverse movement of X-ray tube support minimum 1900mm
| 11 | Rotation of X-ray tube along the vertical axis min. ± 180º.
| 12 | Motorized vertical movement of the X-ray tube support at least 1800 mm.
| 13 | Built-in touch screen on the X-ray tube support with a diagonal of min. 6.25 ″ which is used to control the system and the generator
| 14 | Electromagnetic brake
| 15 | X-ray tube
| 16 | Dual focus max. 0.6 / min. 1.2 mm.
| 17 | Anode heat capacity minimum 300 kHU.
| 18 | Maximum X-ray tube voltage min. 150 kV.
| 19 | Motorized collimator with hand meter and laser direction indicator
| 20 | Collimator rotation range min +/- 45 degrees
| 21 | PATIENT TABLE
| 22 | Patient table with floating table top.
| 23 | Patient table dimensions: min. 2400mm x 800mm.
| 24 | Motorized vertical movement of the patient's table in the range of a minimum of 350 mm.
| 25 | Longitudinal movement of the patient's table
| 26 | Transverse movement of patient table
| 27 | Maximum load on the patient table: min. 300kg.
| 28 | Grid against scattered radiation.
| 29 | VERTICAL BUCKY
| 30 | Vertical Bucky, with anti-scattered grid.
| 31 | Motorized vertical movement of the Vertical Buky from min. 1450 mm.
| 32 | P hand holder for lateral imaging and support for chest exam
| 33 | Tilting range -20 +90 degrees
| 34 | Minimum height of the detector center from floor is max 30 cm
| 35 | WIRELESS FLAT PANEL DETECTOR
| 36 | Flat panel detector size 35x43cm.
| 37 | Scintillator material: Csl (cesium iodide)
| 38 | Resolution: min. 2800x2300 pixels.
| 39 | Pixel size: max. 150µm.
| 40 | Minimum two batteries for flat panel detector with accompanying charger.
| 41 | DQE at 0 or 0,5 lp/mm min 65%
| 42 | FLAT PANEL DETECTOR IN VERTICAL BUCKY
| 43 | Flat panel detector size 43x43cm.
| 44 | Scintillator material: Csl (cesium iodide)
| 45 | Resolution: min. 2860x2874 pixels
| 46 | A / D conversion: min. 16 bit.
| 47 | DQE at 0 lp/mm min 65%
| 48 | ACQUISITION STATION FOR RECEIPT AND PROCESSING OF RECORDINGS
| 49 | Min 23” monitor screen
| 50 | Storage capacity minimum 15000 images
| 51 | Image preview time max 3 sec
| 52 | Image viewing and adjustment function (contrast, rotation, magnification, displacement, annotation).
| 53 | Printing with a network DICOM printer with different combinations of images on one film.
| 54 | DICOM functions: Send, Print and MWL
| 55 | Add patient information to the image, add markers as, contrast correction capability and leave on the image.
| 56 | Possibility of binding to the PACS / RIS system.
6 |  | Ultrasound system for the gastro department |  | 1 |  | 0
| 1 | Ultrasound system, not portable, with wheels and front and rear handles
| 2 | Beamformer of advanced architecture
| 3 | Min 3 active probe connectors not counting the pensil probe connector
| 4 | Systemic dynamic range min. 275 dB
| 5 | Max. frame rate 1750 fps, probe and mode dependent
| 6 | Frequency range 2 - 18 MHz or wider
| 7 | LED monitor min. 21" high resolution min. 1920 x 1080, with the possibility of rotation and tilt adjustment
| 8 | Min 13'' touch panel support multiple gesture recognition and user configured layout
| 9 | Possibility to adjust the console - height/rotation
| 10 | Integrated SSD min 500 GB
| 11 | Minimum 10 TGC zone
| 12 | LGC on the touch screen
| 13 | Max. number of focal points min. 8, probe and application dependent
| 14 | Coded harmonic imaging
| 15 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 16 | Method of spatial compouding of the image from several different angles
| 17 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 18 | Full alphanumeric keypad covered with washable protection film
| 19 | Maximum depth of field min. 30 cm, probe dependent
| 20 | Technique which provides a convex field of view on linear and sector transducers
| 21 | HD write zoom and read zoom min. 50X
| 22 | Voice comments recording
| 23 | Transducer type:
| 24 | Convex array
| 25 | Linear array
| 26 | Microconvex array
| 27 | Sector phased array
| 28 | Volume Probe (4D)
| 29 | Biplane convex / convex
| 30 | Biplane convex / linear
| 31 | Operating modes:
| 32 | B-Mode
| 33 | M-Mode
| 34 | Color M-Mode
| 35 | PW Doppler with high PRF
| 36 | Color Flow mode (CFM)
| 37 | Power Doppler Imaging (PDI) and Directional PDI
| 38 | Hardware and software options which can be added later:
| 39 | Lithium-ion battery that allows the device to operate without mains power min. 15 minutes
| 40 | Method for showing hemodynamics in grayscale using B-Mode technology (not the Doppler technology) which is based on reflection of blood cells (direct, real-time visualization of blood flow echoes)
| 41 | Inputs and outputs:HDMI output (1920 x 1080 resolution), VGA output (1920 x 1080 resolution), CVBS output, S-Video output, 4 x USB, Ethernet (RJ45)
| 42 | B/W thermal printer
| 43 | Probes:
| 44 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 45 | Linear probe Frequency range: 4 – 12 MHz or wider Field of view (FOV): 50 mm (+/- 5 mm)
7 |  | Ultrasound system for use in Radiology |  | 2 |  | 0
| 1 | Ultrasound system, not portable, with wheels and front and rear handles
| 2 | Beamformer of advanced architecture
| 3 | Min 3 active probe connectors not counting the pensil probe connector
| 4 | Systemic dynamic range min. 275 dB
| 5 | Max. frame rate 1750 fps, probe and mode dependent
| 6 | Frequency range 2 - 18 MHz or wider
| 7 | LED monitor min. 21" high resolution min. 1920 x 1080, with the possibility of rotation and tilt adjustment
| 8 | Min 13'' touch panel support multiple gesture recognition and user configured layout
| 9 | Possibility to adjust the console - height/rotation
| 10 | Integrated SSD min 500 GB
| 11 | Minimum 10 TGC zone
| 12 | LGC on the touch screen
| 13 | Max. number of focal points min. 8, probe and application dependent
| 14 | Coded harmonic imaging
| 15 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 16 | Method of spatial compouding of the image from several different angles,
| 17 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 18 | Full alphanumeric keypad covered with washable protection film
| 19 | Maximum depth of field min. 30 cm, probe dependent
| 20 | Technique which provides a convex field of view on linear and sector transducers
| 21 | HD write zoom and read zoom min. 50X
| 22 | Voice comments recording
| 23 | Transducer type:
| 24 | Convex array
| 25 | Linear array
| 26 | Microconvex array
| 27 | Sector phased array
| 28 | Volume Probe (4D)
| 29 | Biplane convex / convex
| 30 | Biplane convex / linear
| 31 | Operating modes:
| 32 | B-Mode
| 33 | M-Mode
| 34 | Color M-Mode
| 35 | PW Doppler with high PRF
| 36 | Color Flow mode (CFM)
| 37 | Power Doppler Imaging (PDI) and Directional PDI
| 38 | Hardware and software options which can be added later:
| 39 | Lithium-ion battery that allows the device to operate without mains power min. 15 minutes
| 40 | Method for showing hemodynamics in grayscale using B-Mode technology (not the Doppler technology) which is based on reflection of blood cells (direct, real-time visualization of blood flow echoes)
| 41 | Inputs and outputs:HDMI output (1920 x 1080 resolution), VGA output (1920 x 1080 resolution), CVBS output, S-Video output, 4 x USB, Ethernet (RJ45)
| 42 | B/W thermal printer
| 43 | Probes:
| 44 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 45 | Linear probe Frequency range: 4 – 12 MHz or wider Field of view (FOV): 50 mm (+/- 5 mm)
8 |  | Ultrasound system for use in Ginecolgy |  | 1 |  | 0
| 1 | Ultrasound system, not portable, with wheels and front and rear handles
| 2 | Minimum 3 active probe ports
| 3 | Dynamic range min. 265 dB
| 4 | Max. frame rate 1200 fps, probe and mode dependent
| 5 | Frequency range 2 - 18 MHz or wider
| 6 | 17'' High Resolution LED Monitor. Tilt/Rotate Adjustable. Resolution min 1280 x 1024
| 7 | At least 10 predefined monitor backlight intensity and temperature settings
| 8 | Integrated HDD min. 500 GB
| 9 | Auto TGC
| 10 | Maximum depth of field min. 40 cm, probe dependent
| 11 | Coded Harmonic Imaging with Pulse Inversion
| 12 | Speckle reduction imaging which provides multiple levels (min. 5) of speckle reduction
| 13 | Method of spatial compouding of the image from several different angles
| 14 | Automatic tissue optimization. Automatic Optimization or corresponding
| 15 | Write Zoom up to 8x and read zoom with HD-Zoom functionality up to 22x .
| 16 | Semiautomatic Fetal Biometry for BPD, AC, HC, HL, FL.
| 17 | Integrated DVD+R(W)/CD-R(W) drive
| 18 | At least 3 programmable keys for direct access to commonly used probes, programs and modes
| 19 | Real-time automatic Doppler calculations
| 20 | Transducer type:
| 21 | Convex array
| 22 | Linear array
| 23 | Microconvex array
| 24 | Sector array
| 25 | Volume Probe (4D) Convex array
| 26 | Volume Probe (4D) Microconvex array
| 27 | Operating modes:
| 28 | B-Mode
| 29 | M-Mode
| 30 | Color Flow Doppler (CFM) mod
| 31 | Pulsed wave Doppler (PW) with sample size - sample volume (Doppler gate) in the range from 0.7 to 15 mm
| 32 | Power Doppler (PD) mod
| 33 | Directional Power Doppler high definition
| 34 | Volume mod (Static 3D i 4D Realtime)
| 35 | Data export in formats:Raw files (2D i Volume); BMP, TIFF, JPEG; AVI, MP4; DICOM Files
| 36 | Connections: Ethernet network connection, 3 USB ports, HDMI Out, VGA port, Audio Out
| 37 | DICOM 3.0 connectivity
| 38 | B/W thermal printer
| 39 | Probes:
| 40 | 2D Convex probe Bandwidth: 2,0 – 5,0 MHz (+/- 0,5) MHz Maximum FOV (wide): min. 80°
| 41 | Wideband Micro-Convex Probe Bandwidth: 3,0 – 10,0 MHz (+/- 0,5 MHz) Maximum FOV (wide): min. 180°
9 |  | Portable ultrasound system for use on intensive care |  | 1 |  | 0
| 1 | Laptop style portable ultrasound system
| 2 | System weight maximum 5 kg (with battery)
| 3 | Cart for the system, height adjustable, with min. 3 probe ports
| 4 | Recent beamforming technology
| 5 | Systemic dynamic range min. 265 dB
| 6 | Max. frame rate 1700 fps, probe and mode dependent
| 7 | Frequency range 2 - 18 MHz or wider
| 8 | High-resolution (min.1920 x 1080 pixels) LCD monitor min.15'' with brightness, contrast and color temperature adjustment
| 9 | Integrated SSD min. 256 GB
| 10 | Min. 8 TGC pods
| 11 | Max. number of focal points min. 8, probe and application dependent
| 12 | Coded harmonic imaging
| 13 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 14 | Spatial compounding
| 15 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 16 | Full alphanumeric keypad covered with washable protection film
| 17 | Maximum depth of field min. 30 cm, probe dependent
| 18 | Technique which provides a convex field of view on linear and sector transducers
| 19 | HD write zoom and read zoom min. 50X
| 20 | Real-time automatic Doppler calculations
| 21 | Voice comments recording
| 22 | Transducer type:
| 23 | Convex array
| 24 | Linear array
| 25 | Microconvex array
| 26 | Sector phased array
| 27 | Volume Probe (4D)
| 28 | Operating modes:
| 29 | B-Mode
| 30 | M-Mode
| 31 | Color M-Mode
| 32 | PW Doppler with high PRF
| 33 | Color Flow mode (CFM)
| 34 | Power Doppler Imaging (PDI) and Directional PDI
| 35 | Continuous wave doppler
| 36 | Inputs and outputs: min. 3 x USB, HDMI output (1920 x 1080), Ethernet (RJ45)
| 37 | Storage formats: - DICOM – compressed/uncompressed, single/multi-frame, with/without raw data - Export JPEG, WMV (MPEG 4) and AVI formats
| 38 | DICOM 3.0 connectivity
| 39 | B/W thermal printer
| 40 | Probes:
| 41 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 42 | Sector probe Frequency range: 2 – 4 MHz or wider Field of view (FOV): min. 120°
| 43 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
10 |  | Portable ultrasound system for use in Ginecolgy |  | 1 |  | 0
| 1 | Laptop style portable ultrasound system
| 2 | System weight maximum 5 kg (with battery)
| 3 | Cart for the system, height adjustable, with min. 3 probe ports
| 4 | Recent beamforming technology
| 5 | Systemic dynamic range min. 265 dB
| 6 | Max. frame rate 1700 fps, probe and mode dependent
| 7 | Frequency range 2 - 18 MHz or wider
| 8 | High-resolution (min.1920 x 1080 pixels) LCD monitor min.15'' with brightness, contrast and color temperature adjustment
| 9 | Integrated SSD min. 256 GB
| 10 | Min. 8 TGC pods
| 11 | Max. number of focal points min. 8, probe and application dependent
| 12 | Coded harmonic imaging
| 13 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 14 | Spatial compounding
| 15 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 16 | Full alphanumeric keypad covered with washable protection film
| 17 | Maximum depth of field min. 30 cm, probe dependent
| 18 | Technique which provides a convex field of view on linear and sector transducers
| 19 | HD write zoom and read zoom min. 50X
| 20 | Real-time automatic Doppler calculations
| 21 | Voice comments recording
| 22 | Transducer type:
| 23 | Convex array
| 24 | Linear array
| 25 | Microconvex array
| 26 | Sector phased array
| 27 | Volume Probe (4D)
| 28 | Operating modes:
| 29 | B-Mode
| 30 | M-Mode
| 31 | Color M-Mode
| 32 | PW Doppler with high PRF
| 33 | Color Flow mode (CFM)
| 34 | Power Doppler Imaging (PDI) and Directional PDI
| 35 | Inputs and outputs: min. 3 x USB, HDMI output (1920 x 1080), Ethernet (RJ45)
| 36 | Storage formats: - DICOM – compressed/uncompressed, single/multi-frame, with/without raw data - Export JPEG, WMV (MPEG 4) and AVI formats
| 37 | DICOM 3.0 connectivity
| 38 | B/W thermal printer
| 39 | Probes:
| 40 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 41 | Endo micro convex probe Frequency range: 4 – 10 MHz or wider Field of view (FOV): min. 168°
| 42 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
11 |  | Portable ultrasound system for use in General surgery |  | 1 |  | 0
| 1 | Laptop style portable ultrasound system
| 2 | System weight maximum 5 kg (with battery)
| 3 | Cart for the system, height adjustable, with min. 3 probe ports
| 4 | Recent beamforming technology
| 5 | Systemic dynamic range min. 265 dB
| 6 | Max. frame rate 1700 fps, probe and mode dependent
| 7 | Frequency range 2 - 18 MHz or wider
| 8 | High-resolution (min.1920 x 1080 pixels) LCD monitor min.15'' with brightness, contrast and color temperature adjustment
| 9 | Integrated SSD min. 256 GB
| 10 | Min. 8 TGC pods
| 11 | Max. number of focal points min. 8, probe and application dependent
| 12 | Coded harmonic imaging
| 13 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 14 | Spatial compounding
| 15 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 16 | Full alphanumeric keypad covered with washable protection film
| 17 | Maximum depth of field min. 30 cm, probe dependent
| 18 | Technique which provides a convex field of view on linear and sector transducers
| 19 | HD write zoom and read zoom min. 50X
| 20 | Real-time automatic Doppler calculations
| 21 | Voice comments recording
| 22 | Transducer type:
| 23 | Convex array
| 24 | Linear array
| 25 | Microconvex array
| 26 | Sector phased array
| 27 | Volume Probe (4D)
| 28 | Operating modes:
| 29 | B-Mode
| 30 | M-Mode
| 31 | Color M-Mode
| 32 | PW Doppler with high PRF
| 33 | Color Flow mode (CFM)
| 34 | Power Doppler Imaging (PDI) and Directional PDI
| 35 | Inputs and outputs: min. 3 x USB, HDMI output (1920 x 1080), Ethernet (RJ45)
| 36 | Storage formats: - DICOM – compressed/uncompressed, single/multi-frame, with/without raw data - Export JPEG, WMV (MPEG 4) and AVI formats
| 37 | DICOM 3.0 connectivity
| 38 | B/W thermal printer
| 39 | Probes:
| 40 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 41 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
12 |  | Portable ultrasound system for use in Cardiology |  | 1 |  | 0
| 1 | Laptop style portable ultrasound system
| 2 | System weight maximum 5 kg (with battery)
| 3 | Cart for the system, height adjustable, with min. 3 probe ports
| 4 | Recent beamforming technology
| 5 | Systemic dynamic range min. 265 dB
| 6 | Max. frame rate 1700 fps, probe and mode dependent
| 7 | Frequency range 2 - 18 MHz or wider
| 8 | High-resolution (min.1920 x 1080 pixels) LCD monitor min.15'' with brightness, contrast and color temperature adjustment
| 9 | Integrated SSD min. 256 GB
| 10 | Min. 8 TGC pods
| 11 | Max. number of focal points min. 8, probe and application dependent
| 12 | Coded harmonic imaging
| 13 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 14 | Spatial compounding
| 15 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 16 | Full alphanumeric keypad covered with washable protection film
| 17 | Maximum depth of field min. 30 cm, probe dependent
| 18 | Technique which provides a convex field of view on linear and sector transducers
| 19 | HD write zoom and read zoom min. 50X
| 20 | Real-time automatic Doppler calculations
| 21 | Voice comments recording
| 22 | Transducer type:
| 23 | Convex array
| 24 | Linear array
| 25 | Microconvex array
| 26 | Sector phased array
| 27 | Volume Probe (4D)
| 28 | Operating modes:
| 29 | B-Mode
| 30 | M-Mode
| 31 | Color M-Mode
| 32 | PW Doppler with high PRF
| 33 | Color Flow mode (CFM)
| 34 | Power Doppler Imaging (PDI) and Directional PDI
| 35 | Continuous wave doppler
| 36 | Inputs and outputs: min. 3 x USB, HDMI output (1920 x 1080), Ethernet (RJ45)
| 37 | Storage formats: - DICOM – compressed/uncompressed, single/multi-frame, with/without raw data - Export JPEG, WMV (MPEG 4) and AVI formats
| 38 | DICOM 3.0 connectivity
| 39 | B/W thermal printer
| 40 | Probes:
| 41 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 42 | Sector probe Frequency range: 2 – 4 MHz or wider Field of view (FOV): min. 120°
| 43 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
13 |  | Portable ultrasound system for use in Orthopedics |  | 1 |  | 0
| 1 | Laptop style portable ultrasound system
| 2 | System weight maximum 5 kg (with battery)
| 3 | Cart for the system, height adjustable, with min. 3 probe ports
| 4 | Recent beamforming technology
| 5 | Systemic dynamic range min. 265 dB
| 6 | Max. frame rate 1700 fps, probe and mode dependent
| 7 | Frequency range 2 - 18 MHz or wider
| 8 | High-resolution (min.1920 x 1080 pixels) LCD monitor min.15'' with brightness, contrast and color temperature adjustment
| 9 | Integrated SSD min. 256 GB
| 10 | Min. 8 TGC pods
| 11 | Max. number of focal points min. 8, probe and application dependent
| 12 | Coded harmonic imaging
| 13 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 14 | Spatial compounding
| 15 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 16 | Full alphanumeric keypad covered with washable protection film
| 17 | Maximum depth of field min. 30 cm, probe dependent
| 18 | Technique which provides a convex field of view on linear and sector transducers
| 19 | HD write zoom and read zoom min. 50X
| 20 | Real-time automatic Doppler calculations
| 21 | Voice comments recording
| 22 | Transducer type:
| 23 | Convex array
| 24 | Linear array
| 25 | Microconvex array
| 26 | Sector phased array
| 27 | Volume Probe (4D)
| 28 | Operating modes:
| 29 | B-Mode
| 30 | M-Mode
| 31 | Color M-Mode
| 32 | PW Doppler with high PRF
| 33 | Color Flow mode (CFM)
| 34 | Power Doppler Imaging (PDI) and Directional PDI
| 35 | Inputs and outputs: min. 3 x USB, HDMI output (1920 x 1080), Ethernet (RJ45)
| 36 | Storage formats: - DICOM – compressed/uncompressed, single/multi-frame, with/without raw data - Export JPEG, WMV (MPEG 4) and AVI formats
| 37 | DICOM 3.0 connectivity
| 38 | B/W thermal printer
| 39 | Probes:
| 40 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 41 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)