## Price schedule
Bidder:
LOT 1 |  | Medical equipment |  |  | Date:
Total price:
Line item No. | ID | Equipment name | Technical Specification Offered (model) | QTY | Unit price | Total Price per line item
1 |  | Trolley for instruments |  | 16 |  | 0
2 |  | Observation ceiling lamp |  | 19 |  | 0
3 |  | Urology table |  | 2 |  | 0
4 |  | Daily surgery gynaecology table |  | 2 |  | 0
5 |  | Examination couch |  | 15 |  | 0
6 |  | Portable ECG machine, 3 channels |  | 5 |  | 0
7 |  | Fetal Monitor |  | 3 |  | 0
8 |  | Negatoscope |  | 6 |  | 0
9 |  | General purpose patient monitor with capnography |  | 17 |  | 0
10 |  | Transport Incubator |  | 2 |  | 0
11 |  | Uroflowmetar |  | 1 |  | 0
12 |  | Histoprocessor |  | 1 |  | 0
13 |  | Machines for cleaning the floors |  | 6 |  | 0
14 |  | X-ray protective gear |  | 6 |  | 0
15 |  | Operating theater surgical instruments |  | 1 |  | 0
16 |  | Surgical instruments for small interventions |  | 1 |  | 0
17 |  | Cabinet for drugs |  | 10 |  | 0
18 |  | Stretcher |  | 12 |  | 0
19 |  | Wheelchair |  | 13 |  | 0
20 |  | Pulse oximeter |  | 7 |  | 0
21 |  | Instruments set for Cystoscopy |  | 1 |  | 0
22 |  | Headlight |  | 1 |  | 0
23 |  | 4K Artrhroscopy and Laparoscopy visualisation system |  | 1 |  | 0
24 |  | Battery operated power tools for orthopedic surgery |  | 1 |  | 0
25 |  | Biopsy Instrument for Urology |  | 1 |  | 0
26 |  | Diagnostic Audiometer |  | 1 |  | 0
27 |  | Sphygmomanometer |  | 44 |  | 0
28 |  | Stethoscope |  | 44 |  | 0
29 |  | IV stand |  | 30 |  | 0
30 |  | Suction unit |  | 18 |  | 0
31 |  | Therapy trolley |  | 17 |  | 0
32 |  | Hospital bed for intensive care |  | 19 |  | 0
33 |  | Feeding Table |  | 157 |  | 0
34 |  | Bedside cabinet |  | 157 |  | 0
35 |  | Portable ECG machine, 12 channels |  | 7 |  | 0
36 |  | Bowl stand |  | 1 |  | 0
37 |  | Dirty linen container |  | 11 |  | 0
38 |  | Clothes rails for surgical capes |  | 4 |  | 0
39 |  | Open trolley for the transport of medical material |  | 2 |  | 0
40 |  | Closed trasnsport trolley |  | 1 |  | 0
41 |  | Anesthesia trolley |  | 9 |  | 0
42 |  | Defibrillator |  | 9 |  | 0
43 |  | Central monitoring station for patient vital signs monitors |  | 2 |  | 0
44 |  | Intesive care monitor |  | 4 |  | 0
45 |  | Mechanical ventilator for IC |  | 1 |  | 0
46 |  | Infusion pump |  | 14 |  | 0
47 |  | Surgical suction unit |  | 2 |  | 0
48 |  | Inhalation device |  | 19 |  | 0
49 |  | Syringe pump |  | 19 |  | 0
50 |  | Head for the central oxygen |  | 135 |  | 0
51 |  | Water manometer used for thoracic or tracheal suction with rail mount |  | 15 |  | 0
52 |  | Mobile waste bin for operating blocks |  | 4 |  | 0
53 |  | Anesthesia device |  | 4 |  | 0
54 |  | Electrosurgical unit |  | 1 |  | 0
55 |  | Electrosurgical unit for ginaecology |  | 1 |  | 0
56 |  | Electrosurgical unit for urology |  | 1 |  | 0
57 |  | Basin stand |  | 8 |  | 0
58 |  | Surgeon's chair |  | 8 |  | 0
59 |  | Instrument table |  | 4 |  | 0
60 |  | Surgical table, gynaecology surgery |  | 1 |  | 0
61 |  | Surgical table, general surgery |  | 1 |  | 0
62 |  | Surgical table, urology surgery |  | 1 |  | 0
63 |  | Tympanograph |  | 1 |  | 0
64 |  | Surgical ceiling supply unit |  | 4 |  | 0
65 |  | Anesthesia ceiling supply unit |  | 4 |  | 0
66 |  | Ceiling supply unit for ICU |  | 6 |  | 0
67 |  | Operating lamp |  | 4 |  | 0
68 |  | Mobile light |  | 1 |  | 0
69 |  | Surgical suction unit 2 |  | 6 |  | 0
70 |  | Flexible Intubation Video Endoscope with acc. |  | 1 |  | 0
71 |  | Doctor chair with backrest |  | 4 |  | 0
72 |  | Doctor chair without backrest |  | 10 |  | 0
73 |  | Bedclothes set |  | 157 |  | 0
74 |  | Fiberoptic laryngosope set |  | 4 |  | 0
75 |  | Four-segment mechanical hospital bed |  | 138 |  | 0
76 |  | Trolley for rewinding |  | 5 |  | 0
77 |  | Rollator Walker |  | 8 |  | 0
78 |  | Walker with “Step-By-Step” motion |  | 3 |  | 0
79 |  | Crutches |  | 30 |  | 0
80 |  | Mechanical child bed |  | 6 |  | 0
81 |  | Physical therapy ultrasound unit |  | 1 |  | 0
82 |  | Exemination bed with steper |  | 1 |  | 0
83 |  | Portable ECG machine, 3 channels pediatric |  | 2 |  | 0
84 |  | Vacuum regulator with rail mount |  | 30 |  | 0
85 |  | General purpose patient monitor with cart |  | 4 |  | 0
86 |  | Baby scale with alitmetar |  | 5 |  | 0
87 |  | Buck percussion hammer |  | 3 |  | 0
88 |  | Digital 38-ch electroencephalograph |  | 1 |  | 0
89 |  | Mobile 6-ch system for EMNG and SEP |  | 1 |  | 0
90 |  | Examination lamp |  | 2 |  | 0
91 |  | Delivery table |  | 2 |  | 0
92 |  | Gynaecology Daily Surgery Chair |  | 1 |  | 0
93 |  | Infant Bassinet |  | 10 |  | 0
94 |  | Resuscitation table |  | 4 |  | 0
95 |  | Infant incubator for intensive care unit |  | 4 |  | 0
96 |  | General purpose patient monitor, neonatal |  | 3 |  | 0
97 |  | Intensive care mechanical ventilator for adults, pediatric and neonatal application |  | 1 |  | 0
98 |  | Bedside pulse oximeter |  | 4 |  | 0
99 |  | Steriliser |  | 2 |  | 0
100 |  | Mobile plazma air desifenction and purificaton device |  | 5 |  | 0
101 |  | Phototherapy lamp |  | 2 |  | 0
102 |  | General purpose patient monitor |  | 1 |  | 0
103 |  | Wall supply unit |  | 1 |  | 0
TOTAL Medical equipment | [5 empty] | 0

## Specifications
Bidder:
LOT 1 |  | Medical equipment |  |  | Date:
Total price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Trolley for instruments |  | 16 |  | 0
| 1 | Carts are made of stainless steel 18/10
| 2 | Dimensions 950 x 650 mm, height 800 mm +/-10%
| 3 | Two inox plates
| 4 | 4 plastic double point castors diameter 75mm, anti-static, wear resistant, of which 2 are with brakes
| 5 | 4 double castors diam. 75mm antistatic with 2 locks
| 6 | 4 bumpers
| 7 | Nominal capacity per plate: min. 50 kg
2 |  | Observation ceiling lamp |  | 19 |  | 0
| 1 | LED technology with single color
| 2 | Round shaped light head, with no sharp edges, easy to clean and maintain.
| 3 | Key pad on the lamp housing
| 4 | Light intensity at distance of 1 meter ≥ 115.000 LUX
| 5 | Color rendering index Ra ≥ 95
| 6 | Focusable light field size from 17-25 cm +- 5%
| 7 | Colour temperature 4500 K
| 8 | Electronic light intensity control at the lamp head 5-100%
| 9 | Life span of the LEDs min. 40.000 h
| 10 | Temperature increase in head area ≤ 0,5 °C
| 11 | Light head diameter ≤ 50cm
| 12 | Two handles placed at the rim for easier manipulation.
| 13 | Additional sterilisable handle
| 14 | Arm 360° rotatable
3 |  | Urology table |  | 2 |  | 0
| 1 | Electromotive adjustment of height, trendelenburg position and back section
| 2 | Height adjustment (electromotive) 600mm-900mm
| 3 | Back section adjustment (electromotive) 0-50º
| 4 | Trendeleburg adjustment (electromotive) ±20º
| 5 | Width of upholstery 550-650mm
| 6 | Safe working load 200kg
| 7 | Activation of all functions by means of foot control
| 8 | Ergonomic upholstery
| 9 | The paper roll compartment up to 500 mm wide located behind the back rest
| 10 | 4 levelling screws for leveling when placing the chair on uneven surface
| 11 | Connection cable at least 3m long
| 12 | The chair must have multilayered powder coating making it resistant to impact and scratches and weather conditions. It has to be easy to maintain and disinfect.
| 13 | Goepel leg supports with integrated laminated hand grips and provision for vertical, horizontal and axial adjustment. The supports must have the sponge cover, tightening strap and stainless steel bars 18/10.
| 14 | Cushion at the head section
| 15 | Plastic cover, for the upholstery of the seat, transparent
| 16 | Mounting step fixed to the chair, movable
| 17 | Side rails (pair) stainless steel
| 18 | Arm rest covered with foam pad, antistatic, 450mm long, with arm strap and attachment clamp
| 19 | Rinsing basin, stainless steel with outlet and stopper and hose
| 20 | Bucket stainless steel, with recess in the lid to insert hose
4 |  | Daily surgery gynaecology table |  | 2 |  | 0
| 1 | Electromotive adjustment of height, trendelenburg position and back section
| 2 | Height adjustment (electromotive) 600mm-900mm
| 3 | Back section adjustment (electromotive) 0-50º
| 4 | Trendeleburg adjustment (electromotive) ±20º
| 5 | Width of upholstery 550-650mm
| 6 | Safe working load 200kg
| 7 | Activation of all functions by means of foot control
| 8 | Ergonomic upholstery
| 9 | Plastic bowl for rinsing that can be pulled in and out
| 10 | The paper roll compartment
| 11 | 4 levelling screws for leveling when placing the chair on uneven surface
| 12 | Connection cable at least 3m long
| 13 | The chair must have multilayered powder coating making it resistant to impact and scratches and weather conditions. It has to be easy to maintain and disinfect.
| 14 | Goepel leg supports with integrated laminated hand grips and provision for vertical, horizontal and axial adjustment. The supports must have the sponge cover, tightening strap and stainless steel bars 18/10.
| 15 | Cushion at the head section
| 16 | Plastic cover, for the upholstery of the seat, transparent
5 |  | Examination couch |  | 15 |  | 0
| 1 | The two segment plate is upholstered with min. 50 millimeters sponge
| 2 | Back plateadjustable up to 40 degrees
| 3 | Base of the table made of steel additionally plastified according to RAL9002 standard
| 4 | Height: min.650mm±5%
| 5 | Width: min.650mm±5%
| 6 | Length: min. 2000mm±5%
| 7 | Paper roll holder
6 |  | Portable ECG machine, 3 channels |  | 5 |  | 0
| 1 | LCD display size 5"
| 2 | Numerical hearth rate monitoring, indication of lead dettachment.
| 3 | Integrated thermal printer with 1, 2 and 3 ECG curves available, paper width 60 mm.
| 4 | Recording speed: 25 and 50 mm/s, resolution 200 dpi.
| 5 | Automatic and manual ECG recording
| 6 | Periodic ECG recording and analysis, with adjustable repeatition interval.
| 7 | Automatic analysis of ECG records and interpretation, for neonatal and adult.
| 8 | Hypertrophy parameters included (RV5/SV1 i RV5+SV1).
| 9 | CMRR: > 105 dB
| 10 | Input impedance: >50 MOhm
| 11 | Offset tolerance: ±550 mV
| 12 | EMG filter 25/35Hz, AC filter 50Hz, high-cut filters 75/100/150Hz.
| 13 | Defibrillation-proof ECG inputs.
| 14 | Storing ECG files to internal and external memory.
| 15 | Integrated web server, providing ECG data to be reviewed using PC with internet browser.
| 16 | Data transfed to PC in PDF format.
| 17 | Digital connections: LAN and USB x2.
| 18 | Integrated WiFi module.
| 19 | Integrated guide finction for error resolving.
| 20 | Upgradeable with bar-code and QR-code reader and magnetic card reader.
| 21 | Power supply: 100-240V AC and battery with capacity for 180 min.
| 22 | Delivered with the following accessory: Power cable, 1 pc Battery, 1 pc ECG cable, 1 pc Set of 6+4 ECG electrodes, 1 pc Recording paper, 10 pcs, Cart, 1 pc
7 |  | Fetal Monitor |  | 3 |  | 0
| 1 | Three dimensional touch screen
| 2 | Four types of scoring methods, NST/Fischer/Improved Fischer/Krebs
| 3 | Timing monitoring function
| 4 | Manual/automatic recording of fetal movements
| 5 | FHR signal strength indicator, twins cross channel verification
| 6 | Support connecting to central monitoring system, both wired and wireless
| 7 | 12” LED touch screen, folding up screen
| 8 | Automatic identification probe
| 9 | Integrated probes bracket
| 10 | Built in thermal printer
| 11 | Device can solve paper jam and deviation defect on herself
| 12 | Alarm system visible from wide angle
| 13 | Rechargeable lithium-ion battery
| 14 | Trend graph/Trend table: 120hours
| 15 | Ultrasound pulse doppler, Frequency 1.0Mhz
| 16 | FHR measurement: 30-250bpm
| 17 | Toco Range: 0-100%
8 |  | Negatoscope |  | 6 |  | 0
| 1 | Light Source: straight daylight fluorescent tubes, by 15W up to 58W.
| 2 | Diffusing Panel: acrylic sheet of 3 mm thickness.
| 3 | Film Grip: zinced metal rollers.
| 4 | Power Supply: 230V; 50/60 Hz
| 5 | Viewing area: 40x90 cm +/-10%
| 6 | Colour temperature: 6500 K
9 |  | General purpose patient monitor with capnography |  | 17 |  | 0
| 1 | TFT LCD color display, min. 10", touchscreen
| 2 | Simultaneous display of min. 5 waveforms and 10 numeric parameters
| 3 | Connections for monitoring of ECG, Respiration, SpO2, NIBP, Temperature
| 4 | Connection for monitoring of IBP
| 5 | Connection for intubated and nonintubated capnography
| 6 | Pacemaker signal detection and rejection
| 7 | Automatic ST segment level detection, alarming and storing
| 8 | Automatic arrhythmia detection, alarming and storing
| 9 | Graphic and numerical parameters trend, min. 72h
| 10 | ECG full disclosure min. 72h
| 11 | Visual and sound alarm
| 12 | Continuous noninvasive or minimally invasive monitoring of estimated hemodynamic parameters CO and SV
| 13 | ECG sync sound with pitch variable according to current SpO2 or IBP value
| 14 | Choice for enlarged numerical display mode, readable from long distance
| 15 | Power supply AC 220V/50Hz with battery backup min. 3h
| 16 | Defibrillation proof patient connectors
| 17 | Holder for fixation to patient bed and medical rail
| 18 | Active connection to central monitoring system, interbed monitoring and wireless telemetry
| 19 | Onscreen operation guide with illustrations
| 20 | Holder for fixation to medical rail, 1 pc
| 21 | To be delivered with all accessories needed, including:  ECG cable with 3 leads, 1 pc  Reusable waterproof SpO2 probe, clip type, 1 pc  Reusable NIBP cuff, large, 1 pc  Reusable NIBP cuff, medium, 1 pc  Accessories for continous noninvasive or minimally invasive monitoring of CO and SV, 1 set Capnography accessories (EtCO2) - 50 adapter pieces  Battery, 1 pc
10 |  | Transport incubator |  | 2 |  | 0
| 1 | Main Unit at Hi-Low stand with with four locking casters
| 2 | LED display, minimum size 4”
| 3 | Four access ports
| 4 | Temterature control: Incubator air temperature setting range: 23-38°C Incubator air temperature display range: 20-42°C Skin temperature display range: 30-42°C
| 5 | Heater power adjustable
| 6 | Built in examination light
| 7 | Built in O2 module: Oxygen concentration display range: 15-105% Maximal oxygen concentration: 60% or more O2 sensor calibration at 21%
| 8 | Noise level: maximum 45dB
| 9 | Rechargeable Li-ion battery for continuous operation One fully charged battery: 180 minutes Optional second fully charged battery: additional 180 minutes
| 10 | AC power source: AC100-240V 50/60Hz
| 11 | Alarms: air temperature, skin temperature, fan, O2 concentracion high/low limits
| 12 | Degree of protection: Type BF applied part
| 13 | To be delivered with all accessories needed, including: - Skin temperature probe reusable, 1piece - Preassure regulator for oxygen cylinders, 1 piece - Flow meter, 1 piece - Power cable for external DC 12V power source, 1 piece - Baby safety band, 3 pieces - Electrostatic filter, 1 piece - Dust cover, 1 piece - Operation manual, 1 piece
11 |  | Uroflowmetar |  | 1 |  | 0
| 1 | Wireless urinal flow meter
| 2 | Wireless Bluetooth communication with PC
| 3 | Bluetooth PC interface and software
| 4 | Digital flowsensor
| 5 | Urinary flow software
| 6 | Bluetooth Dongle for PC
| 7 | Funnel
| 8 | Urine container
| 9 | Possibility of regulating the height of the holder of the urine container
12 |  | Histoprocessor |  | 1 |  | 0
| 1 | Two processing steps (ethanol and isopropanol)
| 2 | Protocols for processing overnight
| 3 | Quick protocols for processing
| 4 | Minimum Two retorts
| 5 | Possibility of two processes simultaneously
| 6 | Capacity minimum 300 cassettes
| 7 | Tissue processing minimum 10 mm thick
| 8 | Protocols for working with and without xylol
| 9 | Automatic molding function at the end of the fixation process
| 10 | Magnetic stirrer of reagents in retort
| 11 | The ability to create protocols by users
| 12 | Compatible with using a commercial reagent packaging
| 13 | Double vapor ventilation system
| 14 | Bottle bar code labeling system and bottle position
| 15 | Remote monitoring of the operation of the histoprocessor via Internet and a warning system
13 |  | Machines for cleaning the floors |  | 6 |  | 0
| 1 | Must be suitable to work in highly overcrowded areas
| 2 | Must have adjustable flow (automatic start-stop with solenoid valve) of the detergent solution depending on type of surface
| 3 | The brush must have a quick release
| 4 | Recovery tank must be removable
| 5 | Must be powered by an internal 12V battery which is recharged by the included charger
| 6 | Must have a battery included
| 7 | Squeegee lifting lever
| 8 | The machine must have easy access and simple service of the components.
| 9 | Scrubbing/squeegee with maximum width 360/460 mm
| 10 | Maximum working capacity 1260 m²/h
| 11 | Traction: mechanical
| 12 | Brush motor power minimum 250 W
| 13 | Detergent solution system must be with solenoid valve
| 14 | Depression/vacuum motor 650 minimum mmH2O/180 W
14 |  | X-ray protective gear |  | 6 |  | 0
| 1 | X-Ray Protective vest front 0,50/back 0,25mm; 100 x 60cm
| 2 | X-Ray Protective vest front 0,50/back 0,25mm; 110 x 60cm
| 3 | X-Ray Protective vest front 0,50/back 0,25mm; 120 x 60cm
| 1 | X-Ray Neck cover: 0,50 mm Pb
| 2 | X-Ray Half apron Pb equivalent: 0,50mm
| 3 | X-Ray Protective glassess protection level: front 0,75/0,50 mmpb
15 |  | Operating theater surgical instruments |  | 1 |  | 0
| 1 | SCALPEL HANDLE, NO 3 & 4, DOUBLE, 16 CM |  | 4
| 2 | SCALPEL HANDLE, NO. 4, 13.5 CM |  | 10
| 3 | SCALPEL HANDLE, NO. 3, SHAPE AS NO. 4 |  | 10
| 4 | DISSECTING SCISSORS, CVD., 18 CM |  | 20
| 5 | SCISSORS, MAYO-STILLE, STR., 15 CM |  | 20
| 6 | OPERATING SCISSORS, BL/BL, STR., 12 CM |  | 10
| 7 | UMBILICAL SCISSORS, US TYPE, 10.5 CM |  | 5
| 8 | DRESSING FORCEPS, 16 CM |  | 25
| 9 | TISSUE FORCEPS, 1X2 T., 16 CM |  | 30
| 10 | ATR. ORGAN SEIZING FORCEPS, 25 CM |  | 5
| 11 | FCPS., KOCHER-NIP. 1X2 T., STR., 16.5 CM |  | 5
| 12 | FCPS., KOCHER-NIP. 1X2 T., CVD., 16 CM |  | 5
| 13 | HOOKLET, BLUNT, 2-PR., 16.5 CM |  | 6
| 14 | Retractor, Doyen, 45x88 mm, 24 cm |  | 4
| 15 | RETR., BRUNNER, CVD., 80X20 MM, 25.5 CM |  | 4
| 16 | PELVIS RETRACTOR, 180X60 MM, 23.5 CM |  | 1
| 17 | RETRACTOR, DOUBLE, ROUX, NO. 1, 14,5 CM |  | 6
| 18 | RETRACTOR, DBL., FARABEUF, SET, 15.5 CM |  | 6
| 19 | NEEDLEHOLDER, MAYO-HEGAR, 20.5 CM |  | 5
| 20 | NEEDLEHOLDER, METZENB., FENESTR., 18.5 C |  | 10
| 21 | NEEDLEHOLDER, MAYO-HEGAR, 16 CM |  | 30
| 22 | LOEFFEL, VOLKMANN, OVAL, NO. 0000, 17 CM |  | 2
| 23 | BONE CUR., VOLKM., OVAL NO. 00, 17 CM |  | 2
| 24 | BONE CURETTE, VOLKMANN, OV./OV., 12.5 CM |  | 2
| 25 | BONE CURETTE, VOLKMANN, OV./OV., 14.5 CM |  | 2
| 26 | BONE CURETTE, VOLKMANN, OV./OV., 17 CM |  | 2
| 27 | BONE RONGEUR, STR., 22.5 CM |  | 1
| 28 | BONE RONGEUR, STILLE, CVD., 22.5 CM |  | 2
| 29 | TISSUE FORCEPS, ALLIS, 5X6 T., 20 CM |  | 5
| 30 | ATR. TISSUE FORCEPS, BABCOCK, 21 CM |  | 6
| 31 | TC-TISSUE FORCEPS, DUVAL, 20.5 CM |  | 3
| 32 | DISS. FORCEPS, KELLY, SLIGHT CVD., 19 CM |  | 2
| 33 | FORCEPS, OVERHOLT, FINE, CVD., 22.5 CM |  | 6
| 34 | HAEMOSTATIC FORCEPS, PEAN, CRV., 18.5 CM |  | 20
| 35 | HAEMOSTATIC FORCEPS, PEAN, CVD., 16,5 CM |  | 20
| 36 | FORCEPS, PEAN, DELICATE, CVD., 14 CM |  | 10
| 37 | FORCEPS, PEAN, DELICATE, CVD., 13 CM |  | 10
| 38 | ABD. RETRACTOR, KIRSCHNER, OVAL. CPL. |  | 1
| 39 | VAGINAL SPECULUM, CUSCO, 85X37 MM |  | 10
| 40 | VAGINAL SPECULUM, CUSCO, 100X35 MM |  | 10
| 41 | VAGINAL SPECULUM, KALLMORGEN, SET, NO. 2 |  | 5
| 42 | TENAC. FCPS., SCHROEDER, 1X1 T., 25.5 CM |  | 10
| 43 | TENACULUM FORCEPS, CVD., 2X2 T., 25 CM |  | 1
| 44 | TENACULUM FORCEPS, STR., 2X2 T., 25 CM |  | 1
| 45 | UTERINE POLYPUS FORCEPS, NOTO, 27 CM |  | 5
| 46 | BIOPSY FORCEPS, FAURE, 21 CM |  | 5
| 47 | UTER. CURETTE, 7 MM, BL., MALL., 30.5 CM |  | 2
| 48 | UTER. CURETTE, 9 MM, BL., MALL., 31 CM |  | 2
| 49 | UTER. CURETTE, 12 MM, BL, MALL., 31.5 CM |  | 2
| 50 | UTER. CURETTE, 7 MM, SH., MALL., 30.5 CM |  | 2
| 51 | UTER. CURETTE, 11 MM, SH., MALL, 31.5 CM |  | 2
| 52 | ASPIRATING CURETTE, Ø 14 MM, 29 CM |  | 2
| 53 | ASPIRATING CURETTE, Ø 10 MM, 29 CM |  | 2
| 54 | ASPIRATING CURETTE, Ø 12 MM, 29 CM |  | 2
| 55 | ABDOMIN. RETRACTOR, BALFOUR, WIRE BLADES |  | 1
| 56 | Bone curette, Volkmann, no.2, si., 27cm |  | 1
| 57 | Bone curette, Volkmann, no.4, si., 27cm |  | 1
| 58 | Osteotome, Lexer, si., 30 mm, 22 cm |  | 1
| 59 | Osteotome, Lexer, si., 20 mm, 22 cm |  | 1
| 60 | Osteotome, Lexer, si., 10 mm, 22 cm |  | 1
| 61 | Osteotome, Lexer, si., 7 mm, 22 cm |  | 1
| 62 | Osteotome, Lexer, si., 5 mm, 22 cm |  | 1
| 63 | Gouge, Lexer, si., 30 mm, 22 cm |  | 1
| 64 | Gouge, Lexer, si., 20 mm, 22 cm |  | 1
| 65 | Gouge, Lexer, si., 10 mm, 22 cm |  | 1
| 66 | BONE CURETTE, VOLKMANN, OV./OV., 12.5 CM |  | 3
| 67 | BONE CURETTE, VOLKMANN, OV./OV., 14.5 CM |  | 3
| 68 | BONE CURETTE, VOLKMANN, OV./OV., 17 CM |  | 3
| 69 | BONE LEVER, STRONG CVD., 6 MM, 15 CM |  | 2
| 70 | BONE LEVER, MINI-HOHMANN, 6 MM, 16 CM |  | 2
| 71 | BONE LEVER, 15 MM, 16 CM |  | 2
| 72 | BONE LEVER, HOHMANN, 18 MM, 23.5 CM |  | 2
| 73 | BONE LEVER, HOHMANN, 43 MM, 23.5 CM |  | 2
| 74 | BONE LEVER, HOHMANN, 70 MM, 25 CM |  | 2
| 75 | BONE LEVER, BENETT, 60 MM, 23.5 CM |  | 2
| 76 | BONE RONGEUR, LUER, STR., 18.5 CM |  | 1
| 77 | BONE RONGEUR, LUER, CVD., 18 CM |  | 1
| 78 | BONE HOLD. FORCEPS, SEMB, CVD., 19.5 CM |  | 1
| 79 | BONE HOLDING FORCEPS, FERGUSSON, 21 CM |  | 2
| 80 | BONE HLD. FORCEPS, W. THREAD LOCK, 26 CM |  | 1
| 81 | HOOKLET, FLEXIBLE, SHARP, 2-PR., 15.5 CM |  | 2
| 82 | HOOKLET, FLEXIBLE, BLUNT, 2-PR., 15.5 CM |  | 2
| 83 | SPREADER, ADSON, SHARP, 3X4 T., 26 CM |  | 1
| 84 | ATR. CAPSULE FORCEPS, MILLIN, 22.5 CM |  | 2
| 85 | CAPSULE FORCEPS, MILLIN, 24.5 CM |  | 2
| 86 | GALLSTONE FORCEPS, DESJARDINS, NO. 2 |  | 2
| 87 | CONTAINER MS, 60X30X14 CM, HANDLE GREY |  | 1
| 88 | CONTAINER MS, 60X30X14 CM, HANDLE ORANGE |  | 1
| 89 | CONTAINER MS, 60X30X14 CM, HANDLE RED |  | 1
| 90 | CONTAINER MS, 60X30X14 CM, HANDLE GREEN |  | 1
| 91 | CONTAINER MS, 60X30X14 CM, HANDLE BLUE |  | 1
| 92 | CONTAINER MS, 60X30X16 CM, HANDLE GREY |  | 1
| 93 | CONTAINER MS, 60X30X16 CM, HANDLE ORANGE |  | 1
| 94 | CONTAINER MS, 60X30X16 CM, HANDLE RED |  | 1
| 95 | CONTAINER MS, 60X30X16 CM, HANDLE GREEN |  | 1
| 96 | CONTAINER MS, 60X30X16 CM, HANDLE BLUE |  | 1
| 97 | CONTAINER MS, 60X30X21 CM, HANDLE GREY |  | 5
| 98 | Silicone mesh f. tray DIN, 48x25 CM |  | 5
| 99 | Feet for marTray, 4 pcs. |  | 5
| 100 | PLASTER SHEARS, BRUNS, SERR., 24 CM |  | 3
16 |  | Surgical instruments for small interventions |  | 1 |  | 0
| 1 | SCALPEL HANDLE, NO. 4L, 21 CM |  | 5
| 2 | SENN-GREEN RETRACTOR, 10X6MM, 16CM |  | 2
| 3 | CUSHING VEIN RETRACTOR 9X10MM, 20CM |  | 1
| 4 | RETRACTOR, FRITSCH, HANDLE VERTICAL TO THE BLADE, 45X75 MM, 25CM |  | 2
| 5 | RETRACTOR, FRITSCH, HANDLE VERTICAL TO THE BLADE, 60X85 MM, 26 CM |  | 2
| 6 | RETRACTOR FOR DEEP SURGERY, SIMON, 115X27 MM, 28 CM |  | 4
| 7 | WOUND STITCH RETRACTOR, 150MM |  | 1
| 8 | BONE FILE, 13MM, 24,5 CM |  | 1
| 9 | KOCHER INTESTINAL CLAMP FORCEPS, SOFT ELASTIC, CURVED, 25,5CM |  | 2
| 10 | SIMS RECTAL SPECULA, 75X17MM, 15,5CM |  | 1
| 11 | KRISTELLER VAGINAL SET OF 1 SPECULUM AND 1 RETRACTOR EACH, 80X30MM |  | 5
| 12 | HEYWOOD-SMITH POLYPUS AND OVUM FORCEPS, 25CM |  | 5
| 13 | RECAMIER UTERINE CURETTE SHARP, MALLEABLE, FIG.2, 9,5MM |  | 2
| 14 | KIRSCHNER ABDOMINAL RETRACTOR COMPLETE, CONSISTING OF: 1 FIXED FRAME 250X250MM, 2 VALVES 43X55MM, 2 VALVES 48X65MM |  | 1
| 15 | Mallet WITH HANDLE FROM ARTIFICIAL WOOD, DIAM.40MM, 740G, 24CM |  | 1
| 16 | VOLKMANN BONE CURETTE WITH HANDLE FROM ARTIFICIAL WOOD, FIG 00, 17,5CM |  | 1
| 17 | KEY RASPATORY, 12,5MM, 19CM |  | 1
| 18 | LUER-STILLE BONE RONGEUR, STRAIGHT, 25CM |  | 1
| 19 | LUER-STILLE BONE RONGEUR, CURVED, 25CM |  | 1
| 20 | BONE HOLDING SELF-CENTERING FORCEPS WITH THREAD FIXATION, 11MM, 28CM |  | 1
| 21 | DESMARRES VEIN RETRACTOR 4,5X8MM, 14CM |  | 2
| 22 | DESMARRES VEIN RETRACTOR 12X16MM, 14CM |  | 2
| 23 | BLAKE GALL STONE FORCEPS, CURVED, 21CM |  | 2
| 24 | OVERHOLT-MIXTER DISSECTING AND LIGATURE FORCEPS, 20,5CM |  | 5
17 |  | Cabinet for drugs |  | 10 |  | 0
| 1 | Made of steel profiles, sheet steel and glass
| 2 | All metal surfaces must be protected with powder coating, white color RAL 9016
| 3 | Two segments in one content
| 4 | Upper segment with glass door
| 5 | Bottom segment with steel door
| 6 | Both doors with locks
| 7 | Dimensions: 80x40x180cm +- 5% (WxDxH)
18 |  | Stretcher |  | 12 |  | 0
| 1 | 2 section stretcher with height adjustment
| 2 | Hydraulic height adjustment from 55 to 90 cm ± 5cm
| 3 | Stretcher length 200 cm ± 5cm
| 4 | Stretcher structure to be made of high thickness steel tubes, with rounded corners. The structure should be covered with epoxy powder for protection against corrosion.
| 5 | The stretcher to be equipped with removable grids, made of electro-welded and galvanized steel, properly locked on plastic supports.
| 6 | The external frame to be equipped with bumpers made of soft non-marking plastic material.
| 7 | The lying surface to be divided into 2 sections and it is provided with ergonomic push handles and housing for accessories, such as I.V.
| 8 | The backrest section is to be activated by bilateral levers and helped in its movements by a pneumatic spring
| 10 | The stretcher to be equipped with 4 Ø 200-mm swiveling castors on ball bearings, antistatic, made of non-marking plastic material with chrome steel supports. The castors will be locked by means of lateral pedals.
| 11 | The stretcher should include a pair of both collapsible and removable side rails.
| 12 | The stretcher should have the push handle made of shaped tubolar steel varnished with epoxy powders. The push handle must have a provision for I.V. stand
| 13 | The stetcher will come with the I.V. stand made of stainless steel with two hooks.
| 14 | The stretcher should have the mattress covered with self-extinguishing CLASS 1 material, breathable, waterproof, antibacterial, anti-mite and latex- and formaldehyde-free. It to be resistant to all common disinfectants, washable with common cleaning products and with hydrogen peroxide, fireproof Class 1 UNI9174-8457.
19 |  | Wheelchair |  | 13 |  | 0
| 1 | Foldable armrest and Detachable footrests
| 2 | Reinforced cushion for great comfor
| 3 | Breaks on the handles
| 4 | Solid wheels
| 5 | Explode light back
| 6 | Rear Wheels
| 7 | Front Wheels
| 8 | Max weight 100kg
| 9 | Seat width: 45cm,
| 10 | Seat depth: 40 cm
| 11 | Total width: 65 cm
| 12 | Steel Frame
20 |  | Pulse oximeter |  | 7 |  | 0
| 1 | A small fingertip pulse oximeter
| 2 | Must have display to present the measured value
| 3 | Technology needs to be motion and low perfusion resistant
| 4 | Evolution to noninvasive respiration rate
| 5 | Hemodynamic variability from the Pleth must be available.
21 |  | Instruments set for Cystoscopy |  | 1 |  | 0
| 1 | Sheath for cystoscope 16.5Fr., slanted distal end, capacity for instruments up to 5Fr, compatible with telescopes Ø 4mm, with 2 flushing ports, quick-lock, reusable |  | 1
| 2 | Obturator for cystoscope 16.5Fr. |  | 1
| 3 | Sheath for cystoscope 19.Fr., slanted distal end, capacity for instruments up to 7Fr, compatible with telescopes Ø 4mm, with 2 flushing ports, quick-lock, reusable |  | 1
| 4 | Obturator for cystoscope 19.5Fr. |  | 1
| 5 | Adapter for cystoscope |  | 1
| 6 | Working insert for cystoscope 30/70 with Allbaran lever |  | 1
| 7 | Flexible grasping forceps 7Fr., working length ≈365mm |  | 2
| 8 | HD Telescope 0°, diameter 4mm, length 300 mm, autoclavable |  | 1
| 9 | HD Telescope 30°, diameter 4mm, length 300 mm, autoclavable |  | 1
| 10 | HD Telescope 70°, diameter 4mm, length 300 mm, autoclavable |  | 1
| 11 | Bladder Syringe 150 ML |  | 1
| 12 | Led Light Source 180W |  | 1
| 13 | Working life of lamp approx. 30.000h
| 14 | Brightness control 0-100%
| 15 | Fiber optic light cable |  | 1
22 |  | Headlight |  | 1 |  | 0
| 1 | Professional headlight with unplugged power source and plug-in transformer on the headband
| 2 | Illumination – at least 90.000 Lux
| 3 | Color temperature - at least 5.500 K
| 4 | LED lifetime – at least 50.000 work hours
| 5 | Cable-free head-worn battery integrated in the headband
| 6 | Fully adjustable to the desired position
| 7 | Adjustment of the diameter range of the work area
| 8 | Work capacity – at least 3 hours of continous operating time at full power
| 9 | Charges to full capacity in 2 hours
23 |  | 4K Artrhroscopy and Laparoscopy visualisation system |  | 1 |  | 0
| 1 | Integrated Visualization System comprises of camera control unit integrated in monitor with minimal display resolution 4096x2160 and minimum size of 32"
| 2 | Video outputs including at minimum UHD3840x2160p, display port 1.2, HDMI 3G-SDI
| 3 | 3 CMOS camera head compatible with 4K monitor resolution, with minimum 4 programmable function, autoclavable or equivalent
| 4 | Electronic LED light guide or equivalent |  | 2
| 5 | Image capture system with minimal resolution 1920x1080p and 1TB internal memory. Data output by minimum 4 x USB 3.0. HDMI video input/output. Touch screen data input
| 6 | Equipment cart with minimum 4 shelves, adjustable height, flex monitor holder, and mount possibility for secondary monitor
| 7 | Electrosurgical unit with dual output minimum power of 300W, minimum two monopolar and bipolar cutting modes with hand and foot controle, minimum two coagulation modes. Easy to use, display contole for each mode. Including safety technology
| 8 | Insufflator with minimum CO2 flow of 40l/min and minimum 3 adjustable flow level. Minimum set pressure 15 mmHg. Different modes including at least mode with continuous smoke and aerosol evacuation
| 9 | Suction-irigation pump with flow rating 2000ml/min and set pressure up to 150mmHg
| 10 | Arthroscopic resection system with two ports and displays and shaver handpiece with minimum forward speed 12000 rpm and oscillate speed 2500cpm
| 11 | Arthroscope dia 10mm 30° deg compatible with resolution of camera control unit
| 12 | Laparoscope dia 10mm 0° compatible with resolution of camera control unit
| 13 | Laparoscope dia 10mm 30° compatible with resolution of camera control unit
| 14 | Set of standard arthroscopic instruments diametar 3.4 or equivalent including linear, wide bite, left and right punch, scissors, aligator grasper and sterilization case
| 15 | Consumables include: tri-lumen tube set with integraded ULPA filter, tri-lumen port dia 5-12mm and min lenght 100mm, single lumen insufflation tube set, suction-irigation tube sets, irrigation cannula, shaver blades and burs
24 |  | Battery operated power tools for orthopedic surgery |  | 1 |  | 0
| 1 | Battery operated modular handpiece for orthopedic surgery with speed range in drill mode 0-1500 rpm and 0-300rpm in ream mode. Compatible with automated washer desinfector. Ingress protection rating of IPX6 minimum |  | 2
| 2 | Handpiece drill attachment model Jacobs, AO and Hudson
| 3 | Handpiece reamer attachment model Synthes and Hudson
| 4 | Battery operated oscillating saw for orthopedic surgery with speed range in drill mode 0-1500rpm and 0-250rpm in reaming mode. Compatible with automated washer desinfector. Ingress protection rating of IPX6 minimum |  | 2
| 5 | Battery charger for simultaneous charging of multiple batteries. Include display information for battery status information. Wall mounting option
| 6 | Lithium battery 13.2V compatible with steam and sterrad sterilization
| 7 | Consumables include oscilationg saw blades
25 |  | Biopsy Instrument for Urology |  | 1 |  | 0
| 1 | Automatic Biopsy Instrument |  | 1
| 2 | Biopsy needles 18G x 20cm |  | 10
26 |  | Diagnostic Audiometer |  | 1 |  | 0
| 1 | Tone audiometry
| 2 | Speech audiometry
| 3 | SISI test with built in counter
| 4 | ABLB test - alternating
| 5 | Standard frequencies: 125, 250, 500, 750, 1000, 1500, 2000, 3000, 4000, 6000, 8000 Hz
| 6 | High frequencies: 12.5 kHz
| 7 | Hearing level ranges:
| 8 | Air Conduction: -10 dB / 120 dB HL at mid-frequencies
| 9 | Bone Conduction: -10 dB / 70 dB HL at mid-frequencies
| 10 | Free Field: 10 dB / 100 dB HL, with active loudspeaker
| 11 | Free Field: 10 dB / 85 dB HL, with passive loudspeaker
| 12 | Level settings click free in steps of 1dB up to 5dB
| 13 | Normal and invers tone presentation
| 14 | USB device interface to PC
| 15 | Power supply: 110/230 V +/- 10%, 50/60 Hz
27 |  | Sphygmomanometer |  | 44 |  | 0
| 1 | Apparatus with air - release valve
| 2 | Upper housing made of aluminium alloy with a shock absorbing ring protecting indicator and glass
| 3 | Lower housing made of shockproof plastic with shock absorbing, anti-slip coating
| 4 | Apparatus must be latex-free
| 5 | With pressure-proof manometer and large diameter scale, at least 55 mm diameter
| 6 | With adjustable spoon for both right or left-handed operation
28 |  | Stethoscope |  | 44 |  | 0
| 1 | With double chest piece with two different frequency ranges to choose from: A high-quality membrane on one side and a conical acoustic head on the other side.
| 2 | Single-tube construction
| 3 | With high-quality membrane, at least 42 mm diameter
| 4 | With cone diameter at least 30mm
29 |  | IV stand |  | 30 |  | 0
| 1 | Infusion stand made of steel with chrome-plated coating
| 2 | Five point star bottom
| 3 | Top with four hooks, adjustable hook height from 1300 to 2000mm
| 4 | Double castors 50mm diameter with anti-corrosive coating
| 5 | Nominal load per hook minimum 2 kg
30 |  | Suction unit |  | 18 |  | 0
| 1 | Flow rate: not less than 40 liters/minute
| 2 | Vacuum -0.8 bar
| 3 | Power supply: 220V
| 4 | Appliance weight maximum 4kg
| 5 | Standard accessories: - polycarbonate autoclavable 1liter bottle, - anti-bacterial filter, - set of silicone tubes, - suction catheter and conical connector
31 |  | Therapy trolley |  | 17 |  | 0
| 1 | Trolley with 4 plastic, anti-static double castors Ø125mm, with two brakes and one directional guide.
| 2 | Plastic work surface
| 3 | Dimensions minimum : width 600mm, height 850 mm
| 4 | Including:
| 4.1 | side rail
| 4.2 | Bucket for medical waste, 1.5lit.
| 4.3 | Infusion stand
| 4.4 | Dispenser carrier for disinfection with plate and clamp
| 4.5 | Bucket carrier
32 |  | Hospital bed for intensive care |  | 19 |  | 0
| 1 | 4-segment lying surface height adjustable by two telescopic system with vertical rectangular section
| 2 | Base made of epoxy painted steel tubes with the addition of antibacterial additive
| 3 | Lying surface made of thermoformed ABS.
| 4 | Ribbed structure od the ABS without metal reinforcement to ensure radiotransparency and maximum accessibility to cleaning.
| 5 | The back and femoral sections electrically adjustable
| 6 | The leg section adjusted by a gas spring
| 7 | External dimensions: L 220cm / W 100cm / H 45 - 85cm +/- 10%
| 8 | Lying surface dimensions: L 195/ W 90cm +/-10%
| 9 | trendelenburg 17 ° max
| 10 | Controtrendelenburg max 17° max
| 11 | Safe working load: 250kg +/- 10%
| 12 | Backrest adjustment: 0 - 70°
| 13 | Femoral section adjustment: 0 - 30°
| 14 | Calf secti on adjustment: 0 - 15
| 15 | The split side rails of polyethylene positioned on the backrest and femoral sections
| 16 | Amortized backrest release (CPR), by means of bilateral levers
| 17 | calf section adjustable through gas spring
| 18 | Bumpers at the 4 corners
| 19 | Integrated rechargeable battery
| 20 | Night light
| 21 | Control unit with acoustic alarm malfunctioning or brakes disconnected
| 22 | 4 receptacles for IV stand and lifting pole
| 23 | Integrated bed extension
| 24 | Head/Foot boards made of polyethylene, easy to remove
| 25 | Set of wheels with total brake one directional, Ø 150mm
| 26 | Degree of protection:: IP X6
| 27 | Power supply: 100-240 ± 10V, 50-60 Hz
| 28 | Antidecubitus mattress must be composed of two differentiated plates and a removable cover
| 29 | Cover
| 30 | Cover made of waterproof fabric in coated polyurethane, breathable, anti mite, antistatic and antibacterial, machine wahable
| 31 | Resistant to urine and other biological liquids
| 32 | Antibacterial, bacteriological, fungistatic and anti-mold properties
| 33 | Permanent anti mite barrier
| 34 | To be delivered with: Support for records Lifting pole Urinal support IV stand with 4 hooks, heingt adjustable Support for drainage bags Bedpan support
33 |  | Feeding Table |  | 157 |  | 0
| 1 | Made of polypropylene with rotation movement
| 2 | Structure and column for the raising system made of chromed steel
| 3 | Column height adjustable through a mechanical spring lifting device
| 4 | Serving table made of thermoformed ABS
| 5 | Reclining movement to use it as a reading desk
| 6 | Serving table slides lenghtwise for 25cm
| 7 | Base equipped with 4 swivelling wheels, diameter 50mm
| 8 | Adjustable height from 95-120 cm ±5 cm
34 |  | Bedside cabinet |  | 157 |  | 0
| 1 | Inner structure made of steel sheets, with reinforced parts, varnished with epoxy powders
| 2 | Upper top printed with thermoformed and shockproof ABS and edge on four sides to contain liquids
| 3 | Table height 80cm +/- 5cm
| 4 | Frontal parts and doors made of bilaminate thickness 20mm with ABS shockproof handles
| 5 | Equipped with external bottle holder
| 6 | Four castors, diameter 50mm, two of them with brake with bumpers at the corners and holed bottom for airing
35 |  | Portable ECG machine, 12 channels |  | 7 |  | 0
| 1 | The main unit with TFT LCD tiltable display size minimum of 7" and viewing all 12-lead ECG waveforms at the same time.
| 2 | Integrated carrying handle
| 3 | Integrated alphanumeric keyboard
| 4 | Integrated thermal printer records 12 ECG curves at same time, paper width of a minimum 210mm
| 5 | Print Speed: 5; 10; 12.5; 25 and 50 mm/s with min. 200 dpi resolution
| 6 | Automatic and manual recording
| 7 | Periodic recording of an ECG record and analysis with adjustable repeat interval
| 8 | Automatic analysis of ECG records with 200 reference findings
| 9 | Common mode rejection ratio: minimal 105dB
| 10 | Input impedance: minimum 50MOhm
| 11 | EMG filter 25/35Hz, AC filter 50Hz i High Frequency filter 75/100Hz
| 12 | AC power and battery power (minimum 60 min operation)
| 13 | Integrated battery charger
| 14 | Internal memory
| 15 | Defibrillation-proof feature
| 16 | Two USB ports and one LAN connection
| 17 | Option: data transfer via wired or wireless LAN
| 18 | To be delivered with all accessories needed, including:  ECG cable with 10 leads, 1 pc  Chest electrodes, 6 pcs  Limb electrodes, 4 pcs  Battery, rechargable, 1 pc  Thermal paper, 10 pcs Cart, 1 pc
36 |  | Bowl stand |  | 1 |  | 0
| 1 | Temperature range of heating physiological solutions and other fluids from 35 ° C to a maximum of 60 ° C.
| 2 | Must have safety thermostat
| 3 | Directly adjustable temperature via button, the temperature set is easily readable.
| 4 | Teflon-coated heating plate for easy cleaning.
| 5 | 5 liters container, made of stainless steel with engraved literary scale.
| 6 | Device mounted on castors Ø 75 mm, two with brakes.
| 7 | Power cable, minimum 5 meters long
37 |  | Dirty linen container |  | 11 |  | 0
| 1 | Made in light alloy plate, anodised after trolley ﬁnished
| 2 | Equipped with tip-downdoor and lid for a total closure
| 3 | Horizontal reinforcement ribs
| 4 | Horizontal push handle ﬁxed at one head
| 5 | Anti-mark PVC bumper in light alloy proﬁle, along the whole lower trolley perimeter proﬁle
| 6 | Equipped with nr.4 wheels in anti-mark rubber diam.160 or 200 (2 ﬁx and 2 swivel)
| 7 | Dimensions: H: 1500mm +- 5% W: 1000mm +-5% D: 700mm +-5%
38 |  | Clothes rails for surgical capes |  | 4 |  | 0
| 1 | Clothes rail - 1 bottom and 1 top shelf 1220x450 mm - hang rail length 1200 mm +/- 10%
| 2 | Dimensions: 1200x450x1750 – LxWxH +/- 10%
39 |  | Open trolley for the transport of medical material |  | 2 |  | 0
| 1 | Made of steel profiles final protected with powder coating
| 2 | The supporting surface made of polysterene
| 3 | The bearing surfaces are formed on the edges of both to prevent swelling of fluid
| 4 | Easily removed for regular maintenance and emergency cleaning
| 5 | Container for disposal of medical waste
| 6 | Mobility provides the wheels with brakes
| 7 | Dimensions: length x width x height – 70x50x80cm
| 8 | With 2 levels
| 9 | Capacity: 30kg
40 |  | Closed transport trolley |  | 1 |  | 0
| 1 | Cabinet totally in stainless steel AISI 304
| 2 | With 1 door closing by key
| 3 | 2 internal heigh adjustable shelves
| 4 | Top with 3 sides railing
| 5 | 4 angular rubber bumpers
| 6 | 4 wheels Ø 125 mm - 2 with brake
| 7 | Dimensions: 650x550x1100 – LxWxH +/- 10%
41 |  | Anesthesia trolley |  | 9 |  | 0
| 1 | Cabinet with 2 smaller drawers and 1 drawer with of 650mm and open compartment at the bottom
| 2 | Cabinet total dimensions: 650x600x850mm (WxDxH)+/- 10%
| 3 | 4 swivel-mounted castors diam. 125mm with revolving bumpers and plastic top, electroconductive, ball-bearing, 2 fulllock, type ones and 1 w/directional lock
| 4 | Guiding bar
| 5 | Rack for supporting accessories mounted on the back of the trolley
| 6 | Fixing bracket for dressing and syringe/cannula dispenser
| 7 | Syringe and cannula dispenser grey plastic housing, with 3 large and 5 small crystal-clear boxes, First-in/First-out system, with twin set swab moistener and dispenser
| 8 | Holder for gloves dispenser, adjustable, for all common disposable glove cartons, made of steel powder coated white aluminium , with plastic rail clamp.
| 9 | Medical waste bin set 1.5 litres
| 10 | Disinfectant dispenser for mounting of container, with SST pump, with fastening plate
| 11 | Waste bag holder,plastic casing, with knee operation, steel holder
42 |  | Defibrillator |  | 9 |  | 0
| 1 | Portable biphasic defibrillator for external defibrillation, synchronous cardioversion and AED
| 2 | Color TFT LCD display, min. 6", capable for min. 3 waveforms and 5 numeric parameters
| 3 | Defibrillator needs to have integrated carrying handle
| 4 | External defibrillation paddles with adult, pediatric and neonatal size
| 5 | Charge and shock buttons needs to be integrated in external paddles
| 6 | Visual indicator for skin impedance
| 7 | Power supply: AC 220V/50Hz and backup battery for continuous ECG monitoring min.2.5h
| 8 | Battery capacity sufficient for min. 100 defibrillations with 200J
| 9 | Defibrillation energy selectable: from 2J to 200J or higher, biphasic
| 10 | Energy charging time from 0 to 200J: max 6s, powered by AC or powered by battery
| 11 | Fast ECG base line recovery after defibrillation
| 12 | Automatic detection of R peak during synchronized cardio version
| 13 | Automatic internal energy discharge after predefined time
| 14 | Visual and sound indication of energy charged
| 15 | Adjustable alarm limits for HR, SpO2, EtCO2 and apnea time
| 16 | Automatic detection of pacemaker signal
| 17 | Automatic arrhythmia detection with sound and visual alarms
| 18 | Automatic defibrillation reports, with storing and recording from memory
| 19 | Defibrillation report with ECG waveform
| 20 | Automatic arrhythmia storing, with numeric parameters and printing reports
| 21 | Capable for storing trends min. 24h
| 22 | Integrated 2-ch thermal recorder
| 23 | Operational temperature: 0 to +40°C or better
| 24 | Dust and moisture protection: min. IP34
| 25 | To be delivered with all accessories needed, including:  ECG cable with 3 leads, 1 pcs  Battery, 1 pcs  Thermal paper, 10 pcs Cart, 1 pc
43 |  | Central monitoring station for patient vital signs monitors |  | 2 |  | 0
| 1 | PC based central station for simultaneous monitoring of vital signs from min. 16 patients
| 2 | TFT LCD color display, min 22"
| 3 | Monitoring of waveforms and numerical parameters and storing for all beds
| 4 | Upgradeable for up to min. 48 monitors
| 5 | Capable to display min 12 simultaneous waveforms for selected patient bed
| 6 | Numeric and graphical trend min. 120h, for each patient
| 7 | Memory for storing
| 8 | Displaying of 12-ch ECG for selected patient, if available at bedside monitor
| 9 | Displaying of patient graphic and visual alarms, three priority levels
| 10 | Adjustment of alarm limits
| 11 | Memory for storing min. 1000 alarm episodes per bed
| 12 | Upgradeable for 12 lead ECG interpretation and 3-ch thermal recorder
| 13 | Dedicated insulation units (transformers) for each patient monitor
| 14 | Patient reports printed on laser printer A4
| 15 | Unconditional power supply (UPS)
| 16 | Communication via LAN, network is provided by hospital
| 17 | Must be compatible with patient monitors offered in other lots
44 |  | Intesive care monitor |  | 4 |  | 0
| 1 | TFT LCD color display, min. 12", touchscreen
| 2 | Connection for optional secondary display of any size
| 3 | Simultaneous display of min. 14 waveforms and corresponding numeric parameters
| 4 | Connections for monitoring of ECG, Respiration, SpO2, NIBP, Temperature
| 5 | Upgradeable for 12-ch ECG monitoring and ECG interpretation with min. 100 findings
| 6 | Connection for monitoring of two simultaneous IBP
| 7 | Connection for intubated and non-intubated capnography and End-tidal CVP
| 8 | Additional NIBP measurements triggered by sudden changes in system pressure or pulse wave propagation time or any similar technique
| 9 | Pacemaker signal detection and rejection
| 10 | Automatic ST segment level detection, alarming and storing for min. 72h
| 11 | Automatic detection and alarming for min. 20 arrhythmia types with storing
| 12 | Graphic and numerical parameters trend, min. 72h
| 13 | Short OCRG trend on display
| 14 | ECG full disclosure min. 72h
| 15 | Visual and sound alarm, 3 priority levels, adjustable limits, storage for min. 72h
| 16 | Alarm lamp on top of the unit, with three colors, visible from all directions
| 17 | ECG sync sound with pitch variable according to current SpO2 or IBP value
| 18 | Choice for enlarged numerical display mode, readable from long distance
| 19 | Power supply AC 220V/50Hz with battery backup min. 1h
| 20 | Defibrillation proof patient connectors
| 21 | Active connection to central monitoring system and interbed monitoring
| 22 | Anesthesia depth monitoring module (BIS or equivalent)
| 23 | To be delivered with all accessories needed, including:  ECG cable with 3 leads, 1 pc  Reusable waterproof SpO2 probe, clip type, 1 pc  Reusable NIBP cuff, large, 1 pc  Reusable NIBP cuff, medium, 1 pc  Accessory for BIS monitoring (or equivalent), 25 pcs  Adapter cable for monitoring of IBP, 1 pc  Battery, 1 pc
45 |  | Mechanical ventilator for IC |  | 1 |  | 0
| 1 | Intensive care mechanical ventilator for adult and pediatric patients, bodyweight 3kg to 130kg.
| 2 | Ventilator based on dedicated trolley with castors and brakes, with arm for patient circuit and two quick-lock holders for the arm (left and right side of the trolley).
| 3 | Ventilator must have carrying handle, with easy one-button removal from the trolley.
| 4 | Power supply AC 220V/50Hz, DC 12-24V from wehicle, with integrated Li-Ion batery backup min. 3 h.
| 5 | Integrated air turbine, quiet and efficient, with HEPA filter, flow min. 240l/min.
| 6 | DIN connection for O2 high pressure central gas supply.
| 7 | Low pressure connection for O2 concentrator.
| 8 | TFT LCD color touchscreen, min. 12", with tilt adjustment.
| 9 | Ventilation modes needed: - Volume controlled CMV and pressure controlled CMV - Volume controlled SIMV and pressure controlled SIMV - Pressure support - BiPAP (or equivalent) and APRV - NIV with guaranteed minimal frequency - High-Flow O2 therapy
| 10 | Ventilator must have closed loop ventilation mode with MV guaranteed, adapting ventilation parameters according to current patient respiratory mechanics (R, C and activity). This adaptive ventilation should work automatically: since intubation, through weaning until extubation. Intended for both intubated and nonintubated patients, for obstructive, restrictive or healthy lungs.
| 11 | Ventilation parameters minimal settings range: - Pinsp: 5 - 60 cmH2O over PEEP - Psupport: 0 - 60 cmH2O over PEEP - PEEP: 0 - 35 cmH2O - I:E: 1:9 do 4:1 - f: 1 - 80 bpm - FiO2: 21 - 100% - Vt: 2 - 2000 ml
| 12 | Flow trigger with automatic adjustment to current leak.
| 13 | Automatic tube resistance compensation.
| 14 | Reversible apnea backup function (off when apnea is over).
| 15 | Capability to drive and control pneumatic nebuliser, upgradeable with electronic nebuliser via USB port.
| 16 | Needed functions: preoxygenation, suction, manual breath, inspiration hold.
| 17 | Visual and audible alarms with priorities and adjustments of alarm limits.
| 18 | Alarm lamp on top of ventilator, visible from all directions (360°).
| 19 | Monitoring trends min. 72h.
| 20 | Capability to display min. four simultaneous waveforms, loops and trends.
| 21 | Ventilator must have graphic animation for visualization of lung dynamics (Vt, R, C) and patient activity, as well as graphic animation for visualization of patient - ventilator dependency (oxygenation, CO2 elimination and spontaneous breathing) with adjustable weaning criteria.
| 22 | Monitoring parameters: - Tidal volume, spontaneous MV and total MV - Spontaneous frequency, total frequency, TI, TE - Minute volume of leak - Ppeak, Pmean, AutoPEEP - Airway resistance, lung compliance, RC and RSB - Oxygen concentration
| 23 | Proximal flow sensor.
| 24 | Event log for alarms and ventilation settings.
| 25 | Volumetric capnography, graph FCO2 vs. V, measurement of etCO2, Vdeadspace and VTalveolar.
| 26 | Program tool for automated lung assessment and recruitment maneuver. Recording of complete quasi static compliance curve (inflation and deflation), for evaluation of hysteresis, with adjustable maneuver parameters and pause for recruitment. Cursor for measurements and finding optimal pressures (PEEP, Pplato). Saving history of loops and reference loops.
| 27 | Communication: USB port for screen snapshots and saving log files, upgradeable with RS232 for remote alarms, nurse call and telemetry.
| 28 | Warm humidifier with manual and automatic temperature / humidity controll and watter filling.
| 29 | To be delivered with all accessories needed, including: - Coaxial patient breathing sets for adult/pediatric, 20 pcs - Breathing sets for adult/pediatric, with humidifier chamber, 15 pcs - Flow sensors, 20 pcs - HME filters adult, 20 pcs - NIV masks, three different sizes, 3 pcs - "High Flow" O2 therapy cannulas, pediatric, 10 pcs - Accessory for volumetric capnography, for adult/pediatric, 1 set
46 |  | Infusion pump |  | 14 |  | 0
| 1 | Volumetric infusion pump for Codan standard and cyto IV sets and other validated sets made of PVC, PVC-free, No Dehp.
| 2 | Aplications: Intensive care, anesthesia, standard care, neonatology, IV administration of fluids, citostatics, oncology, parenteral nutrition, blood transfusion
| 3 | Intuitive and symple to use
| 4 | High accuracy ±5% depending on the used IV system,
| 5 | Flow rate 0,1-1200ml/h in steps of 0,1ml/h, modification without infusion interruption. Automatic rate calculation on the basis of volume and time.
| 6 | Adjustable infusion volume 0,1-9999ml (0,1 - 999,9 ml in steps of 0,1ml from 1000 - 9999 in steps of 1 ml),
| 7 | Adjustable infusion time 1 min 99h 59 min. In steps of 1 min.
| 8 | Dosage units: g, mg, ug, ng, IU, mmol, hour, min., kg, day
| 9 | Provision to limit the flow rate according to the medication name
| 10 | Purging of the IV system at the rate up to 1200ml/h
| 11 | Bolus manual and automatic
| 12 | Bolus programming: Bolus flow rate 0,1-1200ml/h, Bolus volume 0,1-999ml
| 13 | Drop detector (device can also operate without drop detector)
| 14 | Air bubbles detection: individual bubbles configurable 50-1000µl, air accumulation 100-2000µl
| 15 | KVO 0.1-10.0ml/h configurable; deactivation depending on medication
| 16 | Adjustable sensitivity of the occlusion alarm in 10 steps
| 17 | Automatic pressure decrease after occlusion
| 18 | Graphic display of pressure in IV line (bar graph)
| 19 | Data display in graphic LCD display. Permanent display of rate, volume and infusion time
| 20 | Displaying additional information: menu for programming specific functions and alarms, display of additional parameters and settings (pressure limit, batery status, bolus dose calculation etc.)
| 21 | Operation in the following modes: Flow rate, Volume + Time
| 22 | Optional medication database medicines with hard and soft dosage limits, provision for dosage calculation. Dose error reduction system – DERS
| 23 | Provision to operate in the IV stand or on its own
| 24 | Connectable to the docking station, connection to the PDMS system, RS-232 interface
| 25 | Transport function
| 26 | Key lock
| 27 | Memory for more than 1000 events
| 28 | Alarms: Visual alarm in red color, discernable from great distance + dispalyed message on the alarm type, accustic alarm with adjustable volume
| 29 | Pump adjustment and service through computer
| 30 | Mains and battery operated. Battery operation minimum 6 hour at the rate of 25ml/h.
| 31 | Mains adapter 220V 50Hz and charger integrated into the pump housing.
| 35 | Compliance with: EN 60601-1, EN 60601-1-2, EN 60601-2-24, ISO 13485, ISO 9001
47 |  | Surgical suction unit |  | 2 |  | 0
| 1 | Flow rate: not less than 90 liters/minute
| 2 | Vacuum -0.9 bar
| 3 | Noise level ≤ 50dB
| 4 | Foot control
| 5 | Mobile frame with antistatic castors
| 6 | Power supply 230V – 50Hz
| 7 | 2 complete aspiration jar 2000ml Conical fitting Silicon set tubes Antibacterial filter
48 |  | Inhalation device |  | 19 |  | 0
| 1 | An inhalation device with additional equipment for treatment of the lower respiratory tract of patients 4 years and older
| 2 | Selectable aerosol spectrum enables targeted use for the central or the peripheral lung areas
| 3 | Effective medication yield with optional intermittent nebulizing
| 4 | For a large number of inhalations during the day
| 5 | Control System is an integral part of the nebuliser (refer to image) and ensures that a flow rate of approx. 25l/min. is not exceeded during inhalation
| 6 | All nebulisers can be: washed in a dishwasher, disinfected by boiling, chemically disinfected, autoclaved
| 7 | Technical data of the Compressor: Alternating current 230 V~ 50 Hz / 0,7 A / 80 W, Pressure: 1.5 bar
| 8 | Aerosol characteristics: Blue nozzle insert Total Output Rate: 590 mg/min, MMD: 2.9 µm, Mass percentage below 5 μm: 75%
| 9 | Red nozzle insert Total Output Rate: 450 mg/min, MMD: 2.2 µm, Mass percentage below 5 μm: 89%
49 |  | Syringe pump |  | 19 |  | 0
| 1 | Syringe infusion pump for standard use, anaesthesia and intensive care
| 2 | High accuracy ±2% depending on the syringe type
| 3 | Flow rate 0,1-1100ml/h in steps of 0,1ml/h, modification without infusion interruption. Automatic rate calculation on the basis of volume and time.
| 4 | Adjustable infusion volume 0,1- 999,9 in steps of 0,1 and 1 ml,
| 5 | Adjustable infusion time 1 min - 99h 59 min. In steps of 1 min.
| 6 | Provision to limit the flow rate according to the medication name using an optional DERS system
| 7 | Purging and filling the IV line at the rate up to 1100ml/h
| 8 | Bolus manual and automatic
| 9 | Bolus programming: Bolus flow rate 0,1-1100ml/h, Bolus volume 0,1-50ml depending on the size of the syringe.
| 10 | Able to use syringes of: 10, 20 and 50/60ml
| 11 | Automatic syringe detection on display
| 12 | KVO 0.1-5,0 ml/h configurable; deactivation depending on medication
| 13 | Adjustable sensitivity of the occlusion alarm from 200 to 1200 mbar. Permanent monitoring of the line pressure and permanent display on the screen
| 14 | Automatic pressure decrease after occlusion
| 15 | Data display in graphic LCD display. Permanent display of rate, volume and infusion time
| 16 | Displaying additional information: menu for programming specific functions and alarms, display of additional parameters and settings (pressure limit, batery status, bolus dose calculation etc.)
| 17 | Operation in the following modes: Flow rate, Volume + Time,
| 18 | Connectable to the docking station, connection to the PDMS system, RS-232 interface
| 19 | Provision to operate in the IV stand, rail system or be placed on an even surface
| 20 | Key lock
| 21 | Memory for more than 1000 events
| 22 | Pre-alarms. Warnings: infusion near end, syringe almost empty, battery running out. Alarms: Visual alarm in red discernable from great distance + alarm type message on the display, accustic alarm with adjustable volume. Alarms for the end of infusion, occlusion of IV line, badly positioned syringe, empty syringe.
| 23 | Protection against electrical shock from a defibrillator
| 24 | Pump adjustment and service through computer
| 25 | Mains and battery operated. Provision to display the remaining time for the battery operation
| 26 | Mains adapter 220V 50Hz and charger integrated into the pump housing.
| 27 | Compliance with: EN 60601-1, EN 60601- 8, EN 60601-2-24, ISO 13485
50 |  | Head for the central oxygen |  | 135 |  | 0
| 1 | Direct connection to the wall outlet
| 2 | Flowmeter with flow regulator, from (0 to 15) l/min ±10%
| 3 | The flowmeter scale protected with unbreakable polyamid or equivalent material
| 4 | Supply to the central oxygen system
| 5 | Compensated regulator relative to the pressure.
| 6 | Standard DIN connector
| 7 | Outlet: thread 1/4 G (ili ekvivalent)
| 8 | Filter at the scale inlet protecting the device against any gas network impurities and thus ensuring the protection of both the patient and the flowmeter. Filter available as a spare part
| 9 | Oxygen humidifier with reusable bottle 250ml and thread 1/4G
| 10 | Outlet switch: universal rib cone for the hose to the patient
| 11 | Humidifier diffusor, 10 pieces. Diffusor available as consumable material
51 |  | Water manometer used for thoracic or tracheal suction with rail mount |  | 15 |  | 0
| 1 | Twin water manometer with the polycarbonate water tube length of minimum 60cm. Water tube is available as a spare part
| 2 | Regulator head made of chromium-plated brass
| 3 | Second polycarbonate water tube for limitation of the minimum suction. Water tube is available as a spare part
| 4 | Complete rail mounting system with hose, adapter and DIN connector
| 5 | Polycarbonate rod graduated every cm, ensuring vacuum level reading. Rod is available as a spare part
| 6 | 2 L polycarbonate collection jar with moulded hanger and pressure-fit lid with silicone ring, polycarbonate rail clamp
| 7 | Silicone suction tubing of 3m length
| 8 | Single-use suction tip for drainage, 10 pcs
| 9 | Reusable polycarbonate catheter holder with a rail adaptor
52 |  | Mobile waste bin for operating blocks |  | 4 |  | 0
| 1 | Anti-static mobile waste bin for operating blocks
| 2 | Minimum 12l bucket, made of stainless steel, stainless steel ring with protective plasticized buffer
| 3 | Double wheels, diameter 50mm, rust-proof, antistatic
53 |  | Anesthesia device |  | 4 |  | 0
| 1 | Dedicated for anstesia for neonates, children and adults
| 2 | Power supply 220VAC and battery (operating capacity for 90 minutes)
| 3 | color, touch-screen display min. 15“
| 4 | Central gas supply for O2, N2O and Air
| 5 | Backup O2 cylinder can be mounted on the back of the device
| 6 | integrated suction unit on (aspirator)
| 7 | anesthetic gas scavenging system
| 8 | all parts that comes in contact with patient gas can be easily disassembled and sterilized
| 9 | anesthesia device accessories: reusable patient breathing circuit, connection hoses, breathing bag hose, breathing bag and all necessary accessories needed for normal functioning of device
| 10 | device should have minimum one drawer for accessories Also device must be portable (mobile) on castors with brake. Device should have writing table. With device should be delivered all belonging components and connections, breathing hoses, breathing bag hoses, connection hoses for central gas supply
| 11 | device has option to display concentration of anestetic agent
| 12 | Breathing system and gas delivery:
| 13 | breathing circuit system
| 14 | device has posibility for low-flow or minimal-flow anesthesia
| 15 | manual and mechanical ventilation
| 16 | possible emergency manual ventilation in case of power supply failure (external and battery supply)
| 17 | O2 flush command
| 18 | automatic N2O and anesthetic gas shut-off in case of O2 failure
| 19 | electronic regulation of O2 concentration in O2 and N2O mixture of 25%
| 20 | possibility for mechanical ventilation with ambiental air even in case of medical gas supply failure (central gas supply and cylinders) - (electrically or pnematically driven and electronically controlled ventilator
| 21 | adjustable valve for pressure control in manual mode of ventilation
| 22 | Simple transition between mechanical and manual ventilation
| 23 | reusable CO2 absorber, minimal volume 1,3lit
| 24 | Sevoflurane vapor
| 25 | posibility for instalation of two vapors, with block for simultaneous work of both vapors
| 26 | posibility for filling vapors during the operation of device
| 27 | Ventilation modes:
| 28 | manual with CPAP
| 29 | volume controlled ventoilation
| 30 | presssure controlled ventoilation
| 31 | Pressure support breathing
| 32 | Apnea ventilation
| 33 | simple transition between ventilation modes
| 34 | Tidal volume: 20ml to minimal 1500ml
| 35 | breathing frequency: from 4 to minimal 100/min
| 36 | PEEP: 4 to minimal 30 cm Н2О
| 37 | Pinsp (inspiratory pressure): 80 mbar
| 38 | Pressure support: 38 mbar
| 39 | I:Е minimal range: 2.1-1.8
| 40 | adjustable flow trigger 0,3 - 10 l/min
| 41 | inspiratory flow 120 l/min
| 42 | Monitoring:
| 43 | Pressure
| 44 | Volume, inspiratory and expiratory
| 45 | Minute Volume, inspiratory and expiratory
| 46 | Integrated gas analyzer with display for inspiratory and expiratory concentrations of: CO2, N2O, anesthetic gas, MAC values
| 47 | device can show simultaneously ventilation and gas values numerically, graphically and trends
| 48 | digital display of medical gas pressures (О2, Air и N2O) for central gas supply and cylinders
| 49 | Anaesthetic gas consumption with possibility to track paitent responses to anaesthetic during surgery
| 50 | simultaneous display of two loops
| 51 | graphical or tabelar view of trends or minitrends simultaneously with ventilation curves
| 52 | Calibration and alarms:
| 53 | device has automated selftest wich includes leakage test, compliance test and all sensors calibration
| 54 | Programable semi automated self-test, that can be started on programed time
| 55 | alarms for: gas supply low pressure, power supply failure, battery charge
| 56 | audio-visual alarms with multi level priority: minute ventilation, leakage,high airway pressure, gas concentration, alarm for sensor failure.
| 57 | Hemodynamic monitor:
| 58 | Monitor for all patient categories – neonatal, pediatric and adult
| 59 | Touchscreen high resolution color display. Minimal size 15“
| 60 | Monitor has rotary knob to navigate around the screen, select and confirm function. Also has a the following functions: - pause sound alarm - start NIBP measurements - enter trends - freezing curves on screen - printing on the integrated printer - enter the main menu
| 61 | Monitor can show on display 12 or more curves
| 62 | Audio-visual alarm system, with centraly positioned alarm lamps
| 63 | Integrated 3-channel recorder or standalone thermal printer with recording types: ‒ Continuous real-time recording ‒ 8/20 seconds real-time recording ‒ Trend graph recording ‒ Trend table recording ‒ NIBP review recording ‒ Arrhythmia review recording ‒ Alarm review recording ‒ Frozen waveform recording ‒ Drug calculation titration recording
| 64 | Internal battery
| 65 | Monitor has modules for measurement of following parameters: - ECG - NIBP - SpO2 – saturation - temperature - 3 IBP - BIS or Entropy or similar
| 66 | With monitor following accessories must be delivered: - 3 lead ECG kit, - SpO2 reusable sensor for adults - 3 NIBP cuffs (different sizes) - Reusable temp sensor - kit for measurement of 1 IBP
54 |  | Electrosurgical unit |  | 1 |  | 0
| 1 | Possibility of standard monopolar and bipolar electro surgery with integrated system for ligation and division of blood vessels and tissue bundles. Complete system is managed by LCD screen.
| 2 | Two independent monopolar receptacle, both with handswitch activation and one with footswitch activation.
| 3 | Bipolar receptacle with available option of auto-bipolar function, and one receptacle for blood vessel ligation and tissue bundles with bipolar resection possibility.
| 4 | Maximum power for monopolar cut 300W and maximum power for coagulation 120W.
| 5 | Monopolar modes function only with return electrodes equipped with unwonted patient burn system prevention.
| 6 | Minimal three modes of monopolar coagulation including soft coagulation mode.
| 7 | Completely automatic and individual patient return electrode alarm adjustment.
| 8 | Active program for automatic instrument recognition and optimal power settings adjustment.
| 9 | Active program for automatic proper tissue fusion recognition with audio notifications.
| 10 | Possibility of ESU upgrade to the latest released software.
| 11 | Compatibility with reusable ES instruments with single use electrodes for sealing blood vessels and tissue bundles up to and including 7mm in diameter, with independent cutting and without loss of performance according to specific needs of demanding surgical interventions.
55 |  | Electrosurgical unit for ginaecology |  | 1 |  | 0
| 1 | Possibility of standard monopolar and auto bipolar electro surgery. Complete system is managed by LCD screen, screen size 15cm min.
| 2 | Two independent monopolar receptacle, both with handswitch activation and footswitch activation.
| 3 | Bipolar receptacle with available option of auto-bipolar function.
| 4 | Maximum power for monopolar cut 300W and maximum power for coagulation 120W.
| 5 | Monopolar modes function only with return electrodes equipped with unwonted patient burn system prevention.
| 6 | Minimal four modes of monopolar coagulation including soft coagulation mode.
| 7 | Completely automatic and individual patient return electrode alarm adjustment.
| 8 | Active program for automatic instrument recognition and optimal power settings adjustment.
| 9 | Active program for automatic detection of handset and automatic configuration of energy platform.
| 10 | Possibility of ESU upgrade to the latest released software.
56 |  | Electrosurgical generator for urology |  | 1 |  | 0
| 1 | Possibility of standard monopolar and auto bipolar electro surgery. Complete system is managed by LCD screen, screen size 15cm min.
| 2 | Two independent monopolar receptacle, both with handswitch activation and footswitch activation.
| 3 | Bipolar receptacle with available option of auto-bipolar function.
| 4 | Maximum power for monopolar cut 300W and maximum power for coagulation 120W.
| 5 | Monopolar modes function only with return electrodes equipped with unwonted patient burn system prevention.
| 6 | Minimal four modes of monopolar coagulation including soft coagulation mode.
| 7 | Completely automatic and individual patient return electrode alarm adjustment.
| 8 | Active program for automatic instrument recognition and optimal power settings adjustment.
| 9 | Active program for automatic detection of handset and automatic configuration of energy platform.
| 10 | Possibility of ESU upgrade to the latest released software.
57 |  | Basin stand |  | 8 |  | 0
| 1 | 18/10 stainless steel frame
| 2 | two 18/10 stainless steel bowl
| 3 | Each bowl capacity minimum 5 l
| 4 | double castors diam. 50 mm, antistatic, rust-proof
| 5 | Basin stand minimum height 800 mm
58 |  | Surgeon's chair |  | 8 |  | 0
| 1 | Mobile chair
| 2 | 5 concealed plastic double castors Ø 50 mm, interval-braked, electrically conductive
| 3 | frame made of stainless steel, electropolished
| 4 | height adjustment via threaded spindle, with locking against unscrewing
| 5 | seat cushion 60 mm thick, with synthetic leather cover, black, electrically conductive, cleanable and disinfectant resistant
| 6 | Safe workload 135 kg
| 7 | Footprint: Ø 600 mm
| 8 | Height adjustment: 500 to 650 mm
| 9 | Seat: Ø 350 mm
59 |  | Instrument table |  | 4 |  | 0
| 1 | Manually height adjustable from at least 800 to 1200 mm
| 2 | Mobile 4-foot construction made of 18/10 stainless steel
| 3 | Stainless steel tabletop
| 4 | Surface reflex-free and disinfectant-proof
| 5 | 4 castors, abrasion-resistant, electrically conductive, Ø 75 mm,
| 6 | Safe workload: 30 kg
60 |  | Surgical table, gynaecology surgery |  | 1 |  | 0
| 1 | Mobile operating table with electrohydraulic adjustment of specific segments with the certified functionality ≥450kg
| 2 | Mobile with 4 double castors with the diameter 125mm, The castors must provide antistatic table protection.
| 3 | Minimum 4 electrohydraulic adjustments and particularly: - upwards, downwards - trendeleburg, counter-trendeleburg - lateral adjustment - backsection adjustment
| 4 | Basic power supply of the table through batteries with the capacity of 100 operations and more. Batteries are charged during the night or out of working hours of the operating theatre from the hospital grid 230V 50Hz through integrated supply module. In case the batteries run out of power, the table must have provision to operate through the network power supply of 230V 50Hz.
| 5 | Provision for adjustment with manual commands through the hydraulic pump |  |  |  | 0
| 6 | The table must have the terminal with suitable connections for potential equalization
| 7 | Centrally positioned column in respect to the table top providing maximum stability and accessibility
| 8 | The table top - lying surface, consists of 6 segments: - seat section, - short back section, - extended back section, - detachable head section, - two-piece foot section, detachable
| 9 | Radiolucent upper plate - lying surface
| 10 | The pads antistatic, detachable, resistant to disinfectants, minimum thickness: 60mm
| 11 | Provision to set the table in reverse position. Provision to set into zero position, flex & reflex position
| 12 | Control over the table through hand controller connected by cable and through manual selection of the position.
| 13 | Hand control must feature color display for active table adjustments, statutory messages, alarm messages, warnings to the danger of collision of certain segments, battery level indication and display of code errors
| 14 | The table construction and side rails in standard dimensions of 25*10mm, made of stainless steel resistant to disinfectants
| 15 | Total length of the table top with all segments ≥2200mm
| 16 | Total width of the upper plate with the side rails 600mm
| 17 | Height adjustment (without mattress, electrohydraulic) from 600mm ± 20mm to 1100mm ±20mm
| 18 | Setting into trendeleburg, couter-trendeleburg position (electrohydraulic) ±30°
| 19 | Lateral adjustment (electrohydraulic) ±20°
| 20 | Backsection adjustment (electrohydraulic) +70°/-40°
| 21 | Head section adjustment (assisted by the gass spring) +25° / -45°
| 22 | Leg section adjustment (assisted by the gass spring) +10° / -90°
| 23 | Sensor system warning the user of the collision of section table sections
| 24 | Visual indication of significant alarm situations on the hand control unit.
| 25 | Accessories:
| 25.1 | Potential equalization cable 5m
| 25.2 | Head rest
| 25.3 | Split type leg section
| 25.4 | Back section extension, fixed, lenght 300 mm±20mm
| 25.5 | Upholstered arm rest with the fixation band. Vertically, horizontaally adjustable and at an angle, including the clamp for the side rail.
| 25.6 | Arm rest in the position against the body
| 25.7 | Infusion pole
| 25.8 | Anaesthetic frame including the clamp
| 25.9 | Anaesthetic frame extension
| 25.10 | Arm fixation for the anesthetic frame as electrostatic band with which the arm is tied to the frame.
| 25.11 | Gopel type leg support with clamps, 2 pcs.
| 25.12 | Band for the fixation of the patient's body
61 |  | Surgical table, general surgery |  | 1 |  | 0
| 1 | Mobile operating table with electrohydraulic adjustment of specific segments with the certified functionality ≥450kg
| 2 | Mobile with 4 double castors with the diameter 125mm, The castors must provide antistatic table protection.
| 3 | Minimum 4 electrohydraulic adjustments and particularly: - upwards, downwards - trendeleburg, counter-trendeleburg - lateral adjustment - backsection adjustment
| 4 | Basic power supply of the table through batteries with the capacity of 100 operations and more. Batteries are charged during the night or out of working hours of the operating theatre from the hospital grid 230V 50Hz through integrated supply module. In case the batteries run out of power, the table must have provision to operate through the network power supply of 230V 50Hz
| 5 | Provision for adjustment with manual commands through the hydraulic pump
| 6 | The table must have the terminal with suitable connections for potential equalization
| 7 | Centrally positioned column in respect to the table top providing maximum stability and accessibility
| 8 | The table top - lying surface, consists of 6 segments: - seat section, - short back section, - extended back section, - detachable head section, - two-piece foot section, detachable.
| 9 | Radiolucent upper plate - lying surface. |  |  |  | 0
| 10 | The pads antistatic, detachable, resistant to disinfectants, minimum thickness: 60mm.
| 11 | Provision to set the table in reverse position. Provision to set into zero position, flex & reflex position
| 12 | Control over the table through hand controller connected by cable and through manual selection of the position.
| 13 | Hand control must feature color display for active table adjustments, statutory messages, alarm messages, warnings to the danger of collision of certain segments, battery level indication and display of code errors.
| 14 | The table construction and side rails in standard dimensions of 25*10mm, made of stainless steel resistant to disinfectants
| 15 | Total length of the table top with all segments ≥2200mm
| 16 | Total width of the upper plate with the side rails 600mm
| 17 | Height adjustment (without mattress, electrohydraulic) from 600mm ± 20mm to 1100mm ±20mm
| 18 | Setting into trendeleburg, couter-trendeleburg position (electrohydraulic) ±30°
| 19 | Lateral adjustment (electrohydraulic) ±20°
| 20 | Backsection adjustment (electrohydraulic) +70°/-40°
| 21 | Head section adjustment (assisted by the gass spring) +25° / -45°
| 22 | Leg section adjustment (assisted by the gass spring) +10° / -90°
| 23 | Sensor system warning the user of the collision of section table sections.
| 24 | Visual indication of significant alarm situations on the hand control unit.
| 25 | Accessories:
| 25.1 | Potential equalization cable 5m
| 25.2 | Head rest
| 25.3 | Split type leg section
| 25.4 | Back section extension, fixed, lenght 300 mm±20mm
| 25.5 | Upholstered arm rest with the fixation band. Vertically, horizontaally adjustable and at an angle, including the clamp for the side rail
| 25.6 | Arm rest in the position against the body
| 25.7 | Infusion pole
| 25.8 | Anaesthetic frame including the clamp
| 25.9 | Anaesthetic frame extension
| 25.10 | Arm fixation for the anesthetic frame as electrostatic band with which the arm is tied to the frame.
| 25.11 | Gopel type leg support with clamps, 2 pcs.
| 25.12 | Band for the fixation of the patient's body
62 |  | Surgical table, urology surgery |  | 1 |  | 0
| 1 | Mobile operating table with electrohydraulic adjustment of specific segments with the certified functionality ≥450kg
| 2 | Mobile with 4 double castors with the diameter 125mm, The castors must provide antistatic table protection.
| 3 | Minimum 5 electrohydraulic adjustments and particularly: - upwards, downwards - trendeleburg, counter-trendeleburg - lateral adjustment - backsection adjustment - leg section adjustment
| 4 | Basic power supply of the table through batteries with the capacity of 100 operations and greater. Batteries are charged during the night or out of working hours of the operating theatre from the hospital grid 230V 50Hz through integrated supply module. In case the batteries run out of power, the table must have provision to operate through the network power supply of 230V 50Hz.
| 5 | Additional emergency control, removable, cable connected, situated at the column.
| 6 | Provision for adjustment with manual commands through the hydraulic pump
| 7 | The table must have the terminal with suitable connections for potential equalization.
| 8 | Centrally positioned column in respect to the table top providing maximum stability and accessibility.
| 9 | The table top - lying surface, consists of 6 segments: - seat section, - short back section, - extended back section, - detachable head section, - two-piece foot section, detachable.
| 10 | Radiolucent table top - lying surface.
| 11 | The pads antistatic, detachable, resistant to disinfectants, minimum thickness: 60mm.
| 12 | Provision to set the table in reverse position. Provision to set into zero position, flex & reflex position
| 13 | Control over the table through hand controller connected by cable, emergency controller mounted at the column, and through manual selection of the position
| 14 | Hand controller and emergency controller must feature color display for active table adjustments, statutory messages, alarm messages, warnings to the danger of collision of certain segments, battery level indication and display of code errors.
| 15 | The table construction and side rails in standard dimensions of 25*10mm, made of stainless steel resistant to disinfectants
| 16 | Total length of the table top with all segments ≥2200mm
| 17 | Total width of the upper plate with the side rails 600mm
| 18 | Height adjustment (without mattress, electrohydraulic) from 600mm ± 20mm to 1100mm ±20mm
| 19 | Setting into trendeleburg, couter-trendeleburg position (electrohydraulic) ±30°
| 20 | Lateral adjustment (electrohydraulic) ±20°
| 21 | Backsection adjustment (electrohydraulic) +70°/-40°
| 22 | Head section adjustment (assisted by the gass spring) +25° / -45
| 23 | Leg section adjustment (electrohydraulic) +70° / -90°
| 24 | Leg section adjustment (assisted by the gass spring) +10° / -90°
| 25 | Sensor system warning the user of the collision of section table sections.
| 26 | Visual indication of significant alarm situations on the hand control unit.
| 27 | Accessories:
| 27.1 | Potential equalization cable 5m
| 27.2 | Head rest
| 27.3 | Split type leg section
| 27.4 | Back section extension, fixed, lenght 300 mm±20mm
| 27.5 | Upholstered arm rest with the fixation band. Vertically, horizontaally adjustable and at an angle, including the clamp for the side rail.
| 27.6 | Arm rest in the position against the body
| 27.7 | Infusion pole
| 27.8 | Anaesthetic frame including the clamp
| 27.9 | Urological adapter. To extend the seat section by 420mm±20mm, radiolucent.
| 27.10 | Gopel type leg support with clamps 2 pcs.
| 27.11 | Band for the fixation of the patient's body
| 27.12 | Swivel mounted rinsing basin with holder, stainless steel, including clamp for attachment to the side rails.
| 27.13 | Sieve insert for basin, steinless steel.
| 27.14 | Bracket for body support. Needed for attachment of the body support. 2 pcs.
| 27.15 | Back/ Buttock Body support, antistatic, 200 x 120 mm, 2 pcs.
| 27.16 | Orthopeadic extension device, 2 extension bar supports with swivelling double joints and traction units, 2 foot plates, 2 traction boots for adults, and countertraction post for dorsal position.
| 27.17 | Positioning plate for lateral positions.
| 27.18 | Countertraction post for lateral position.
| 27.19 | Transfer leg plates (pair)
| 27.20 | Kirschner wire device.
| 27.21 | Transport trolley for extension device
| 27.22 | Traction device for tibial fractures.
| 27.23 | Knee psitioning device with the clamp
| 27.24 | Shoulder surgery plate with individually fold-down and removable padded plates and one lateral body support, with fixture and U shaped head rest.
| 27.25 | Humerus positioning device with the clamp
| 27.26 | Humerus countertraction post with the clamp
| 27.27 | Weinberger hand traction device
| 27.28 | Axilla support
| 27.29 | Elbow support
| 27.30 | Traction unit for traction of the arm in forearm fractures with the clamp.
63 |  | Tympanograph |  | 1 |  | 0
| 1 | Power supply: 110-220 V~, 50/60 Hz, 60 VA
| 2 | Microprozessor controlled
| 3 | LCD display placed on front panel
| 4 | Integrated thermal printer
| 5 | Probetone: 226Hz, 86dB SPL
| 6 | Pressure range: +200 to -400 daPa
| 7 | Pressure delta: 2 seconds for compliance
| 8 | Reflex tones: 500, 1000, 2000 and 4000 Hz at 85, 95 and 105 dB
| 9 | Reflex meassurement: ipsi and contralateral
| 10 | Automatic reflex recognition
| 11 | State indicator: 1. LED 3-colors in probe 2. detailed display - simbols
| 12 | Accessories: 1. Tympprobe 2. 30 ear-plugs in 5 different sizes 3. Headset for contralateral test 4. Probe holder
64 |  | Surgical ceiling supply unit |  | 4 |  | 0
| 1 | Arm system:
| 2 | Arm system with two horizontal swivel arms
| 3 | The arms are manufactured from extruded aluminum profile, powder coated , and the ends being sealed with plastic covers
| 4 | Arm system equipped with friction and pneumatic brake
| 5 | In case pressure drop ceiling supply unit must remain in same position
| 6 | Combined arm length 1700 mm ±3%
| 7 | Swivel range: 330°±3%
| 8 | Stop positions ceiling bearing: every 15° or less
| 9 | Supply column:
| 10 | The supply column is manufactured from extruded aluminum profiles
| 11 | Max load supply column min 100kg and more
| 12 | It is equipped with media installation panels on 3 sides
| 13 | These panels run the full length of the column and can be used for the installation of gas outlets, electrical receptacles, and other media components
| 14 | Two full length, front mounted vertical rails allow simple attachment and infinitely variable height adjustment of shelves
| 15 | For simple hygienic disinfection the supply column features smooth surfaces, rounded corners and installation plates without gaps or screws
| 16 | Push-buttons for operating the pneumatic brakes can be positioned on the back side of the column
| 17 | Swivel range: 330°±3%
| 18 | Column length: 1250-1400mm
| 19 | Outlets:
| 20 | 2 x oxygen
| 21 | 2 x medical-grade air
| 22 | 2 x vacuum
| 23 | 1 x Air motor
| 24 | 4 x 230V UPS
| 25 | 4 x 230V Aggregate
| 26 | 2 x RJ45
| 27 | 8 x potential equalization pin
| 28 | Workplace components:
| 29 | 2 x shelves without drawer, surface area min 550x450mm, load capacity min 35kg, standard rails (25x10 mm) at left and right hand side
| 30 | 1 x shelves with drawer, surface area min 550x450mm, load capacity min 35kg, standard rails (25x10 mm) at left and right hand side
65 |  | Anesthesia ceiling supply unit |  | 4 |  | 0
| 1 | Arm system:
| 2 | Arm system with one swivel arm, electromotoric height adjustment and ceiling bearing
| 3 | The arms are manufactured from extruded aluminum profile, and the ends being sealed with plastic covers
| 4 | Height adjustment is done with electrically-driven lift motor, during compete stoke column must be vertical
| 5 | Arm system equipped with friction and pneumatic brake
| 6 | In case pressure drop ceiling supply unit must remain in same position
| 7 | Combined arm length 1700 mm ±3%
| 8 | Swivel range: 330°±3%
| 9 | Stop positions ceiling bearing: every 15° or less
| 10 | Supply column
| 11 | The supply column is manufactured from extruded aluminum profiles
| 12 | Column must be height adjustable to provide lifting of anesthesia device and best positioning in workplace
| 13 | Max load supply column min 250kg for carrying anesthesia device and other equipment
| 14 | It is equipped with media installation panels on 5 sides These panels run the full length of the column and can be used for the installation of gas outlets, electrical receptacles, and other media components
| 15 | Two full length, front mounted vertical rails allow simple attachment and infinitely variable height adjustment of shelves
| 16 | Supply column features smooth surfaces, rounded corners and installation plates without gaps or screws
| 17 | Push-buttons for operating the pneumatic brakes and height adjustment can be positioned on the back side of the column
| 18 | Swivel range: 330°±3%
| 19 | Height adjustment range min. 600 mm
| 20 | Column length: 1250-1400mm
| 21 | Outlets:
| 22 | 2 x oxygen
| 23 | 2 x medical-grade air
| 24 | 2 x vacuum
| 25 | 1 x nitrous oxide
| 26 | 1 x Anesthesia Gas Scavenging System
| 27 | 4 x 230V UPS
| 28 | 4 x 230V Aggregate
| 29 | 2 x RJ45
| 30 | 8 x potential equalization pin
66 |  | Ceiling supply unit for ICU |  | 6 |  | 0
| 1 | Arm system left side (infusion and suction)
| 2 | Arm system with two horizontal swivel arms
| 3 | The arms are manufactured from extruded aluminum profile, and the ends being sealed with plastic covers
| 4 | Arm system equipped with friction and pneumatic brake
| 5 | In case pressure drop ceiling supply unit must remain in same position
| 6 | Combined arm length 1200 mm ±3%
| 7 | Swivel range: 330°±3%
| 8 | Stop positions ceiling bearing: every 15° or less
| 9 | Arm system right side (respiration and monitoring)
| 10 | Arm system with two horizontal swivel arms
| 11 | The arms are manufactured from extruded aluminum profile, and the ends being sealed with plastic covers
| 12 | Arm system equipped with friction and pneumatic brake
| 13 | In case pressure drop ceiling supply unit must remain in same position
| 14 | Combined arm length 750 mm ±3%
| 15 | Swivel range: 330°±3%
| 16 | Stop positions ceiling bearing: every 15° or less
| 17 | Supply column left side:
| 18 | The supply column is manufactured from extruded aluminum profiles
| 19 | Max load supply column min 100kg
| 20 | It is equipped with media installation panels on 3 sides. These panels run the full length of the column and can be used for the installation of gas outlets, electrical receptacles, and other media components
| 21 | Two full length, front mounted vertical rails allow simple attachment and infinitely variable height adjustment of shelves
| 22 | For simple hygienic disinfection the supply column features smooth surfaces, rounded corners and installation plates without gaps or screws
| 23 | Push-buttons for operating the pneumatic brakes can be positioned on the back side of the column
| 24 | Swivel range: 330°±3%
| 25 | Range of vertical motorized adjustment: 600mm or more
| 26 | Vertical motorized speed: 20 sec or less
| 27 | Column length: 1250-1400mm
| 28 | Outlets:
| 29 | 2 x oxygen
| 30 | 2 x medical-grade air
| 31 | 2 x vacuum
| 32 | 4 x 230V UPS
| 33 | 4 x 230V Aggregate
| 34 | 2 x RJ45
| 35 | 8 x potential equalization pin
| 36 | Supply column right side:
| 37 | The supply column is manufactured from extruded aluminum profiles
| 38 | Max load supply column min 100kg
| 39 | It is equipped with media installation panels on 3 sides
| 40 | Panels run the full length of the column and can be used for the installation of gas outlets, electrical receptacles, and other media components
| 41 | Two full length, front mounted vertical rails allow simple attachment and infinitely variable height adjustment of shelves
| 42 | Supply column features smooth surfaces, rounded corners and installation plates without gaps or screws
| 43 | Push-buttons for operating the pneumatic brakes can be positioned on the back side of the column
| 44 | Colour of media installation panels: Anodized aluminum
| 45 | Swivel range: 330°±3%
| 46 | Range of vertical motorized adjustment: 600mm or more
| 47 | Vertical motorized speed: 20 sec or less
| 48 | Column length: 1250-1400mm
| 49 | Outlets:
| 50 | 2 x oxygen
| 51 | 2 x medical-grade air
| 52 | 2 x vacuum
| 53 | 4 x 230V UPS
| 54 | 4 x 230V Aggregate
| 55 | 2 x RJ45
| 56 | 8 x potential equalization pin
| 57 | Workplace components left side:
| 58 | 1 x shelves with drawer, surface area min 550x450mm, load capacity min 38kg, standard rails (25x10 mm) at left and right hand side
| 59 | 1 x infusion holder for min 4 pumps
| 60 | 1 x holder for infusion bottles
| 61 | 1 x aspirator holder
| 62 | Workplace components right side:
| 63 | 2 x shelves without drawer, surface area min 550x450mm, load capacity min 38kg, standard rails (25x10 mm) at left and right hand side
| 64 | 1 x shelves with drawer, surface area min 550x450mm, load capacity min 38kg, standard rails (25x10 mm) at left and right hand side
67 |  | Operating lamp |  | 4 |  | 0
| 1 | Combination of two operating lights on heavy central axis
| 2 | LED technology with multi color
| 3 | Round shaped light head, with no sharp edges, easy to clean and maintain.
| 4 | User-friendly display and keypad
| 5 | Light must have Endo-mode with green light and depth-light
| 6 | Light intensity at distance of 1 meter ≥ 160.000 LUX
| 7 | Color rendering index Ra ≥ 98
| 8 | Focusable light field size from 20-30 cm +- 10%
| 9 | Colour temperature from 3750 to 4750 K
| 10 | Dimming range 50 - 100%
| 11 | Life span of the LEDs min. 60.000 h
| 12 | Temperature increase in head area ≤ 0,5 °C
| 13 | Total power consumption ≤ 80 W
| 14 | Protection type of light head IP 54 and better
| 15 | Two handles placed at the rim for easier manipulation.
| 16 | Additional sterilisable handle
| 17 | Arm 360° rotatable
| 18 | Description of the satellite:
| 19 | LED technology with multi color
| 20 | Round shaped light head, with no sharp edges, easy to clean and maintain.
| 21 | Display and keypad
| 22 | Light intensity at distance of 1 meter ≥ 140.000 LUX
| 23 | Colour temperature from 3750 to 4750 K
| 24 | Dimming range 50 - 100%
| 25 | Light must have Endo-mode with green light and depth-light
| 26 | Color rendering index Ra ≥ 98
| 27 | Light head diameter ≤ 60cm
| 28 | Protection type of light head IP 54 and better
| 29 | Life span of the LEDs min. 60.000 h
| 30 | Total power consumption ≤ 90 W
| 31 | Focusable light field size from 20-30 cm +- 10%
| 32 | Arm 360° rotatable
68 |  | Mobile light |  | 1 |  | 0
| 1 | Stand model
| 2 | LED technology with single color
| 3 | Round shaped light head, with no sharp edges, easy to clean and maintain.
| 4 | Light intensity at 1m distance 115 000 Lux.
| 5 | Centrally positioned light segment which improves the depth illumination.
| 6 | Illumination depth 1000mm or more.
| 7 | Color rendering index Ra≥ 95
| 8 | Color temperature 4500K
| 9 | Diameter of the light field 20 to 30cm +- 10%
| 10 | Provision to replace individual light segments.
| 11 | Optical system to increase the contrast in the light field.
| 12 | Ambient light (Endo light)
| 13 | Diameter of the lamp head ≤500mm
| 14 | Life span of the LEDs ≥60 000h
| 15 | Light intensity control 10-100% and more
| 16 | Temperature increase in the head area ≤ 0,5 °C
| 17 | Total power consumption ≤30W
| 18 | Additional sterilisable handle
| 19 | Two handles placed at the rim for easier manipulation.
69 |  | Surgical suction unit 2 |  | 6 |  | 0
| 1 | Flow rate: not less than 90 liters/minute
| 2 | Vacuum -0.9 bar
| 3 | Noise level ≤ 50dB
| 4 | Foot control
| 5 | Mobile frame with antistatic castors
| 6 | Power supply 230V – 50Hz
| 7 | Standard accessories: - two polycarbonate, autoclavable bowls of 5 liter equipped with the anti-overflow valve, - set of silicone hoses, conical appendage
70 |  | Flexible Intubation Video Endoscope with acc. |  | 1 |  | 0
| 1 | Flexible Intubation Video Endoscope Set 4.0 x 65 with 7” monitor.
| 2 | The flexible intubation video endoscope for universal use
| 3 | For the placement of Endotracheal Tubes (ETT) as of size 6 mm
| 4 | High image resolution and video imaging in 4:3 format – no tunnel vision
| 5 | Practical tube fixation via special adaptor
| 6 | Suitable and validated for the following low temperature reprocessing methods up to max 65 °C
| 7 | Manual/machine cleaning and disinfection
| 8 | Sterilization and EtO gas
| 9 | High-Level Disinfection
| 10 | Direction of view: 0°
| 11 | Angle of view: 100°
| 12 | Working length: 65 cm
| 13 | Total length: 95 cm
| 14 | Working channel inner diameter: 2.1 mm
| 15 | Distal tip outer diameter : 5.5 mm
| 16 | Deflection up/down: 140°/140
| 17 | Monitor for CMOS Endoscopes Set:
| 18 | Screen size 7" with 1280 x 800 pixel resolution
| 19 | Two camera inputs
| 20 | USB and a HDMI port
| 21 | Optimized user interface
| 22 | Video and image capture in real time on SD card
| 23 | Playback of recorded video clips and still images
| 24 | Data transfer from SD card to USB flash drive possible
| 25 | HDMI output for connection to anexternal monitor
| 26 | Run while charging and with rechargeable Li-ion batteries
| 27 | Video Laryngoscope D-BLADE Set:
| 28 | For adults
| 29 | CMOS technology
| 30 | For difficult intubation
| 31 | Documentation of images and video sequences via BlueButton
| 32 | With catheter introduction sizes 16 – 18 Fr.
| 33 | Including guide adapted to blade shape
| 34 | For direct and indirect endotracheal intubation
| 35 | Special curved blade shape for difficult intubation
| 36 | Video Laryngoscope MAC #3:
| 37 | CMOS technology
| 38 | With MACINTOSH laryngoscope blade
| 39 | Size 3
| 40 | With system interface
| 41 | Documentation of images and video sequences via button
| 42 | With catheter introduction sizes 16 – 18 Fr
| 43 | For direct and indirect endotracheal intubation
| 44 | Mobile Stand:
| 44.1 | Rides on 4 antistatic dual wheels
| 44.2 | 2 equipped with locking brakes
| 44.3 | Including: Subrack, for mobile stand Top Cover, with guide sleeve Stainless Steel Tube, length 135 cm Equipment Rail Cross Tube Adaptor Stainless Steel Round Pipe, length 25 cm. Shelf, incl. installation accessories, max. capacity 10 kg, Holder for Flexible Endoscopes, Wire Basket, for accessories.
71 |  | Doctor chair with backrest |  | 4 |  | 0
| 1 | Doctor chair with antistatic surface, thickness minimum 80 mm
| 2 | Height adjustment by foot pedal from 550 to 800 mm±20mm
| 3 | Pedestal with 5 double antistatic castors 50mm in diameter
| 4 | The backrest is antistatic, skin identical as a saddle part
| 5 | Safe load 130 kg
72 |  | Doctor chair without backrest |  | 10 |  | 0
| 1 | Doctor chair with antistatic surface thickness minimum 80 mm
| 2 | Height adjustment by foot pedal from 550 to 800 mm±20mm
| 3 | Pedestal with 5 double antistatic castors 50mm in diameter
| 4 | Safe load 130 kg
73 |  | Bedclothes set |  | 157 |  | 0
| 1 | Cushion: cover: white linen cotton 100% filler: siliconized beads PES 100% dimensions: 60x80cm +/-10%
| 2 | Quilt:cover: cover: white linen cotton 100% filler: konflin PES 100%, weight 200g / m2 Dimensions: 140x200cm +/-10%
| 3 | Set of upholstery for: - mattress (cotton 100%, with a holding) - pillow (cotton 100%, on a raffle) - comforter (cotton 100%, on raffle)
74 |  | Fiberoptic laryngosope |  | 4 |  | 0
| 1 | Laringoscopes set includes: 1.1 Macintosh Fiber Optic Blade 2 114x14 1.2 Macintosh Fiber Optic Blade 3 134x14 1.3 Macintosh Fiber Optic Blade 4 155x15 1.4 Fiber Optic Handle 1.5 case
| 2 | Macintosh Fiber Optic Blades: 2.1 Achieves the best possible illumination of the pharynx even after 4,000 sterilisation cycles 2.2 For simple and safe intubation 2.3 Design from chrome-plated, stainless, high-grade steel 2.4 Blade is more resistant to the oxidation and / or corrosion damage 2.5 Blades have no screwed joints, no external fiber bundles and no openings that could contaminate 2.6 The clean shape without any edges or corners is easy to clean and sterilise 2.7 Compatible with all handles to the ISO 7376 (green standard) 2.8 4000 standard autoclave cycles of 134 °C / 4 min. with more than 1000 lux
| 3 | Laringoscope Handle 3.1 Compatible with all laryngoscope blades to the ISO 7376 (green standard) 3.2 Can be converted from battery to rechargeable handles by simply exchanging the bottom insert 3.3 brightness reduces slowly even with low residual capacity 3.4 Colour Temperature: 5,000 K
| 4 | Table Charger 4.1 Dimensions: Width 142 mm, Height 61 mm, Depth 66 mm 4.2 Two independent charging wells 4.3 Complete functionality for one or two handles at a time 4.4 Automatic charging 4.5 Cannot be overcharged 4.6 Charger switches off automatically when the battery is fully charged 4.7 Charge status indicator
75 |  | Four-segment mechanical hospital bed |  | 138 |  | 0
| 1 | Four segment fixed height bed
| 2 | Dimensions: L: 200cm / W: 97cm/ H: 55cm +/- 5%
| 3 | Structure made of steel pipes coated with epoxy powder
| 4 | Safe workload: 250kg +/- 10%
| 5 | Backrest and the leg are powered by gas pumps with two handles both sides of the bed
| 6 | Foot section by an automatick rack with six adjustments
| 7 | Basin section is fixed to prevent the diaphragm crushing in the highest inclination of the sectors
| 8 | Head/Foot panels made of polyethylene
| 9 | Set of wheels with total brake one directional, Ø 125mm
| 10 | Equipped with side rails couple with compass movement
| 11 | Side rails composed of three horizontal bars with a diameter of 25mm
| 12 | To be delivered with: Support for records Lifting pole Urinal support IV stand with 4 hooks, heingt adjustable Support for drainage bags Bedpan support
| 13 | Fireproof mattress with removable cover, composed of central core in polyurethane foam, waterproof barrier covering, all machine washable
76 |  | Trolley for rewinding |  | 5 |  | 0
| 1 | Trolley with 4 closed, plastic, anti-static double castors Ø125mm, with two brakes and one directional guide.
| 2 | Plastic work surface
| 3 | Minimal dimensions width 600mm, height 850 mm
77 |  | Rollator Walker |  | 8 |  | 0
| 1 | Rollator with breaks on handles
| 2 | With seat, basket and backrestt
| 3 | Handles adjustable by height
| 4 | Max patient weight: 100kg
| 5 | Seat width: 45cm
78 |  | Walker with “Step-By-Step” motion |  | 3 |  | 0
| 1 | Foldable frame
| 2 | Double function: fix & swinging
| 3 | Adjustable by height
| 4 | Optional walking movement at the arms
| 5 | Max patient weight: 100 kg
79 |  | Crutches |  | 30 |  | 0
| 1 | Steel frame
| 2 | Adjustable height
| 3 | Max load: 220kg
80 |  | Mechanical child bed |  | 6 |  | 0
| 1 | With 2 lowerable sideguards
| 2 | With bumpers
| 3 | Levina castors diam. 100 mm with brake
| 4 | Dimensions: 150 x 70 cm
| 5 | Material: Steel, epoxy-coating
| 6 | Frame made of rounde steel tubes.
| 7 | Mattress platform made of perforaded steel
81 |  | Physical therapy ultrasound unit |  | 1 |  | 0
| 1 | Total number of channels : 5 ( 4 completely independent bipolar electrotherapy channels ) and 1 channel for ultrasound therapy frequencies 1MHz and 3MHz with application in continuous and pulsed mode 10%, 20%, 50% and duty cycle 16Hz, 48Hz and 100Hz with possibility of activation of the ultrasound head contact surface, connectors for 2 probes | `
| 2 | Combined ultrasound and 4-pole interferential currents, ultrasound and 2-pole interferential currents, ultrasound and high voltage pulsed current, ultrasound and asymmetrical biphasic TENS current and ultrasound and symmetrical biphasic TENS current
| 3 | Full-color rotational touch/screen LCD diagonal 10' with screen angle adjustment
| 4 | Pre-loaded clinical protocols with therapy treatment proposals according to body part, clinical indication and treatment goal and in the case of pain treatment according to acute and chronic conditions with photographs of electrodes positions and information of the therapy current type
| 5 | Memory positions for memorizing most frequently used protocols
| 6 | Anatomical library that illustrates most common array of pathologies
| 7 | Possibility of creation therapy sequence consisting of 3 electro-therapy currents
| 8 | Possibility of memorizing each patient session : electrode position, pain area, pain scale on USB memory stick
| 9 | Possibility for creation of icons ( shortcuts ) on the display for frequently used therapy protocols
| 10 | 'Print screen'' option
| 11 | High-speed USB port for software update and patient treatment data transfer
| 12 | WAVEFORMS
| 13 | Four-pole Interferential currents with vector function, base frequency in the range of 2-10kHz, modulation frequency in the range of 1-200Hz, possibility of base and modulation frequency setting and vector function angle setting
| 14 | Two-pole ( pre-modulated ) interferential with possibility of setting of base and modulation frequency
| 15 | Galvanic, continuous and medium frequency modulated current, with possibility of manual reversal of electrode polarity
| 16 | TENS - asymmetrical and symmetrical biphasic wave form - Monophasic rectangular and alternating rectangular wave form
| 17 | MONOPHASIC ELECTRICAL STIMULATION with rectangular and triangular waveform, with possibility of adjustment phase and pause duration
| 18 | KOTS – Russian current
| 19 | Symmetrical biphasic electrical stimulation for muscle strengthening with high intensity and short pulse duration ( phase duration adjustable from 20-1000μs, intensity adjustable from 0-200mA )
| 20 | Symmetrical biphasic electrical stimulation ( Burst ) for muscle strengthening with high intensity and short pulse duration ( phase duration adjustable from 20-700μs, intensity adjustable from 0-200mA in burst form
| 21 | Symmetrical biphasic electrical stimulation ( Burst ) for muscle strengthening where one channels stimulates agonist and second antagonists with pulse duration adjustable 20-400μs, intensity adjustable from 0-150mA
| 22 | HAN stimulation Diadynamic waveform with possibility of automatic polarity reversal on the half of treatment time in following forms : - DF - MF - CP – LP - CP-iso - CP-id - MF+CP - DF+LP - DF+CP - MF+CP-id
| 23 | High voltage pulsed current
| 24 | Microcurrent
| 25 | Trabert / Ultra Reiz current
| 26 | Faradic current with rectangular and triangular form with adjustable frequency phase duration and surges in minutes
| 27 | Iontophoresis Jwith calculation of the dose mA/min and therapy time according to entered amplitude (mA)
| 28 | I/T diagnostic with I/T curve display
| 29 | Measuring of electrical potential of muscle activity via 2-channel EMG feedback ( sensitivity below 3μV ) with possibility of one-channel EMG triggered electrical situation. Display of muscle electrical activity via bar-graph or EMG curve with electrode positioning guidelines for different body parts
| 30 | Power supply : 220 V, 50Hz
| 31 | Accessories : - Velcro straps - 4 pieces of electrode cables for channels 1-4 - 2 pieces of electrode cables for surface EMG - Rubber-graphite electrodes 6x8 cm (8 pieces ) with sponges - adhesive EMG surface electrodes 4 pcs - User manual on Serbian language - ultrasound transducer ( probe ), ERA 2cm2
| 32 | Cart with 3 drawers
| 33 | 2-channel vacuum
82 |  | Exemination bed with steper |  | 1 |  | 0
| 1 | The two segment plate is upholstered with min. 70 millimeters sponge
| 2 | Back plate adjustable up to 40 degrees
| 3 | Base of the table made of steel additionally plastified according to RAL9002 standard
| 4 | Height: min.800mm
| 5 | Width: min.650mm
| 6 | Length: min. 1900mm
| 7 | Paper roll holder
| 8 | Step stool with one step, 18/10 stainless steel frame, ine-ribbed anti-slip tread in 18/10 stainless steel, floor pads antistatic
| 9 | Step stool Safe working load 180 kg±5kg
83 |  | Portable ECG machine, 3 channels pediatric |  | 2 |  | 0
| 1 | LCD display size 5"
| 2 | Numerical hearth rate monitoring, indication of lead dettachment.
| 3 | Integrated thermal printer with 1, 2 and 3 ECG curves available, paper width 60 mm.
| 4 | Recording speed: 25 and 50 mm/s, resolution 200 dpi.
| 5 | Automatic and manual ECG recording, printing copy of the last recording.
| 6 | Last 3 minutes of ECG waveforms can be storred internally before recording and analysing the selected segment.
| 7 | Periodic ECG recording and analysis, with adjustable repeatition interval.
| 8 | Automatic 1-ch rhytm recording, paper savings by multi-row printing.
| 9 | Automatic analysis of ECG records and interpretation with 200 findings, for neonatal and adult.
| 10 | Hypertrophy parameters included (RV5/SV1 i RV5+SV1).
| 11 | CMRR: > 105 dB
| 12 | Input impedance: >50 MOhm
| 13 | Offset tolerance: ±550 mV
| 14 | EMG filter 25/35Hz, AC filter 50Hz, high-cut filters 75/100/150Hz.
| 15 | Defibrillation-proof ECG inputs.
| 16 | Storing files to internal and external memory.
| 17 | Integrated web server, providing ECG data to be reviewed using PC with internet browser.
| 18 | Data transfed to PC in PDF format.
| 19 | Digital connections: LAN and USB x2.
| 20 | Integrated WiFi module.
| 21 | Integrated guide finction for error resolving.
| 22 | Delivered with the following accessory: Power cable, 1 pc Battery, 1 pc ECG cable, 1 pc Set of 6+4 ECG electrodes, 1 pc Recording paper, 10 pcs, Set of 6+4 pediatric ECG electrodes, 1pc Cart, 1 pc
84 |  | Vacuum regulator with rail mount |  | 30 |  | 0
| 1 | Vacuum regulator: 0-1000 mbar
| 2 | Medical vacuum pipeline systems drive
| 3 | Complete rail mounting system with hose, adapter and DIN connector
| 4 | ON/OFF switch-button providing a quick restoration of the pre-adjusted vacuum level
| 5 | Manual adjustment of the vacuum gauge from -45° to +45°for a better visibility
| 6 | 100 ml safety jar equipped with a mechanical anti-overflow safety valve and a single-use antibacterial filter
| 7 | Tube of 10 antibacterial plastic filters
| 8 | 2 L polycarbonate collection jar with moulded hanger and pressure-fit lid with silicone ring, polycarbonate rail clamp
| 9 | Silicone suction tubing of 3m length
| 10 | Single-use suction tip for drainage, 10 pcs
| 11 | Reusable polycarbonate catheter holder with a rail adaptor
85 |  | General purpose patient monitor with cart |  | 4 |  | 0
| 1 | TFT LCD color display, min. 10", touchscreen
| 2 | Simultaneous display of min. 5 waveforms and 10 numeric parameters
| 3 | Connections for monitoring of ECG, Respiration, SpO2, NIBP, Temperature
| 4 | Connection for IBP monitoring
| 5 | NIBP measurements triggered by changes in pulse wave propagation time
| 6 | Pacemaker signal detection
| 7 | Automatic ST segment level detection, alarming and storing
| 8 | Automatic arrhythmia detection, alarming and storing: min. 10s per episode, total min. 120h
| 9 | Graphic and numerical parameters trend, min. 120h
| 10 | ECG full disclosure min. 120h
| 11 | Visual and sound alarm, 3 priority levels, adjustable limits
| 12 | Continuous noninvasive monitoring of estimated hemodynamic parameters CO and SV
| 13 | ECG sync sound with pitch variable according to current SpO2 and IBP value
| 14 | Choice for enlarged numerical display mode, readable from long distance
| 15 | Power supply AC 220V/50Hz with battery backup min. 3h
| 16 | Defibrillation proof patient connectors
| 17 | Onscreen operation guide with illustrations
| 18 | Each monitor to be delivered with the following accessory:  ECG cable with 3 leads, 1 pc  Reusable waterproof SpO2 probe, clip type, 1 pc  Reusable NIBP cuff, large, 1 pc  Reusable NIBP cuff, medium, 1 pc  Accessory for continous noninvasive monitoring of CO and SV, 1 set - Holder for monitor, fixation to medical rail or wall, 1 pc - Battery, 1 pc - Cart
86 |  | Baby scale with altimetar |  | 5 |  | 0
| 1 | Network-capable with wireless technology
| 2 | Adjustable damping.
| 3 | High-quality scale technology with fine 5 g graduation.
| 4 | TARE, HOLD, auto-HOLD and BMIF for extra convenience
| 5 | Extra large and ergonomic tray for safe weighing.
| 6 | Optional measuring rod expands scale into a measuring station.
| 7 | Capacity: 20 kg
| 8 | Graduation: 5 g < 7.5 kg > 10 g
| 9 | Power supply: Batteries / mains adapter (optional)
| 10 | Data transmission: wireless technology
| 11 | Functions: TARE, BMIF, HOLD, auto-HOLD, weighing range switch-over, adjustable damping, automatic switch-off
| 12 | RS232 compatible
| 13 | Calibration class: III
| 14 | System compatible with: wireless printers, software analytics, wireless USB adapter
| 15 | Measuring range in cm: 35 - 80 cm
| 16 | Graduation [Measuring Rod]: 1 mm / 1/8 inch
87 |  | Buck percussion hammer |  | 3 |  | 0
| 1 | Highest quality finish
| 2 | Chrome-plated
| 3 | With needle and brush
88 |  | Digital 38-ch electroencephalograph |  | 1 |  | 0
| 1 | Based on brand name Windows PC computer with 24" display and with A4 laser printer.
| 2 | Dedicated metal trolley (medical grade), with isolation transformer.
| 3 | Photo stimulator with flash, mounted on portable stand with castors
| 4 | Digital EEG amplifier with 38 channels (29 EEG/PSG inputs, 3 bipolar respiration inputs, 4 DC inputs, integrated SpO2 and CO2 inputs).
| 5 | Input impedance min. 100 MOhm, noise max. 1.5uV, CMMR min. 105 dB.
| 6 | Sampling frequency 1000 Hz (optional 10,000 Hz).
| 7 | EEG amplifier must have integrated impedance check button with LED indicators for all leads, mounted on portable stand with castors
| 8 | AC filter, EMG filters, separate ECG artefact elimination filter.
| 9 | Event marking and patient marking button
| 10 | Software for EEG recording and review, patterns, 2D frequency and amplitude mapping
| 11 | FFT spectral analyses software and software for digital spectrum vs. time, of whole EEG recording
| 12 | Simultaneous display of 8 temporal digital spectrums of whole EEG recording
| 13 | Power and voltage 3D mapping software
| 14 | Patterns with 1-250 channels or more
| 15 | EEG review function simultaneous with EEG recording function
| 16 | Automatic recognition and quantification of waves (amplitude, interval, frequency).
| 17 | Automatic report generation
| 18 | Patient and EEG recordings database, SQL type, with HL7 compatibility
| 19 | To be delivered with all accessories needed, including: - EEG caps, different sizes, 3 pcs - EEG cap adapter for direct (one connector only) connection to EEG amplifier,1pc - ECG clips and cables, 1 pair
89 |  | Mobile 6-ch system for EMNG and SEP |  | 1 |  | 0
| 1 | Main unit:
| 2 | Main unit based on PC of latest generation, with display 22", laser printer.
| 3 | Mobile cart for medical purpose with integrated insulation transformer
| 4 | Control panel with function keys
| 5 | Double foot switch with programmable functions
| 6 | Input amplifier:
| 7 | Digital amplifier with min. 6 input channels, with adjustable arm
| 8 | Input impedance >1000MΩ (common mode), CMRR >112dB iso, noise <0,6μVrms
| 9 | Impedance check button and indicators available on input box: 2-50 kΩ
| 10 | A/D conversion min. 18 bit, sampling rate min 20 kHz for 6 ch simultaneously
| 11 | Sensitivity: 1µV/div -10mV/div
| 12 | Low frequency filter: 0.01 Hz to 3 kHz or better
| 13 | High frequency filter: 10 Hz to 20 kHz or better
| 14 | Notch AC filter: 50/60 Hz
| 15 | Waveform rectification: half and full.
| 16 | Stimulation:
| 17 | Two electrical stimulators 0-200mA, mono and biphasic pulse. Upgradeable to 4 stimulators
| 18 | Active electrode for active removal of stimulation artefacts
| 19 | Software:
| 20 | Licensed operating system Windows 10
| 21 | Software for min 12 simultaneous recordings
| 22 | Software for EMG, MUP auto detection, trigger for MUP, automatic turn/amp histogram, QEMG with FFT analyses, upgradeable with Single Fiber EMG and Macro EMG
| 23 | Software for somatosensory evoked potentials: SEP, SSEP. ECG-SEP, ESCP
| 24 | Software for neurography: NCS, MCS, SCS, F-wave, Repetitive stimulation with auto sequencing, Blink, H-Reflex, Collisions
| 25 | Cortical potentials back averaging based on EMG trigger
| 26 | Automatic report generation in HTML
| 27 | SQL patient database with neuro recordings and reports
| 28 | Accessories:
| 29 | Reusable disk electrodes, 10 pcs
| 30 | Grounding electrode, 1 pcs.
| 31 | Stimulation electrode with filtz, 1 pcs.
| 32 | Finger electrode, 1 pcs.
| 33 | Conductive paste and abrasive paste, 1 pcs.
| 34 | EMG needles, three sizes, total 75 pcs
| 35 | Stimulating electrode with integrated regulation of the intensity of stimulation, 1 pcs.
| 36 | Disc shaped electrode for hand grounding, 1 pcs.
| 37 | Stimulators
| 38 | Two electric stimulators 0-200mA, mono i biphasic pulse. Provision to extend to 4 stimulators.
90 |  | Examination lamp |  | 2 |  | 0
| 1 | Light intensity in Lux at 1 meter distance 130.000 Lux
| 2 | Colour temperature 4300K
| 3 | Colour rendering index Ra: 95
| 4 | Focussable light field size: 17-24cm
| 5 | Diameter of light head 40cm +/- 5%
| 6 | Working distance: 70-140cm +/- 5%
91 |  | Delivery table |  | 2 |  | 0
| 1 | The delivery table consists of the main section and separable foot section. The main section featuring seat and back sections. Both the main and the separable seat sections are equipped with castors, with central locking device.
| 2 | The separable foot section can be separated from the main section enabling easy access for midwife and doctor. The separable foot section may also be used as a seat or storage. It can be pushed under the main section and can be locked in four positions.
| 3 | Seat section features a recess of 150 mm radius.
| 4 | Two-part pad for main and foot sections made of highly elastic coated foam, moisture repellent and secured against slipping.
| 5 | 1 two-part sheath for pads, may be washed up to 95° C.
| 6 | Casing and frame made of powder-coated sheet steel
| 7 | Back rest may be attached at the head or foot end, with artificial leather covering
| 8 | Two Göpel-type leg supports, stainless steel 18/10, integral foam pad in imitation leather silver-grey, height adjustable.
| 9 | Two rotating joints to attach the leg supports.
| 10 | Two handles, stainless steel 18/10, integrated foam in imitation leather silver grey, may be attached at the sides.
| 11 | Infusion rod with 4 hooks, stainless steel 18/10, may be attached at the head side on both sides, height adjustable.
| 12 | Back support near the seat recess, semicircular, integral foam in imitation leather silver-grey, for birth in squatting position.
| 13 | Two hand controls with cable connection, left and right of the main section, for infinite adjustment driven by an electric motor.
| 14 | Soft start function of all bed movements.
| 15 | Height of main section adjustable from 650 to 900 mm
| 16 | Height of foot section adjustable from 450 to 900 mm
| 17 | External dimensions of delivery table with back rest 2300(L) x 1000 (W)mm±50mm
| 18 | Matress tickness 80mm
| 19 | Back section adjustable from 0 to 78°
| 20 | Seat section adjustable from 0 to 27°
| 21 | Inclination of main section (shock position) from 0 to 12°
| 22 | Safe workload 220kg.
| 23 | Stainless steel bowl
92 |  | Gynaecology Daily Surgery Chair |  | 1 |  | 0
| 1 | Examination and treatment chair with memory functions . Electromotor adjustment of pelvic elevation, and back section.
| 2 | Ergonomic upholstery at the back, seat and head section
| 3 | SST bowl for rinsing that can be pulled in and out
| 4 | The paper roll compartment
| 5 | 4 columns for leveling when placing the chair on uneven surface
| 6 | On the base construction there are the place for connecting the colposcope support on both sides.
| 7 | Connection cable at least 3m long.
| 8 | The chair must have high quqlity powder coating making it resistant to impact and scratches and weather conditions.
| 9 | The padded sections must be of high quality foam material covered in artificial leatherette upholstery
| 10 | Hand and the foot control with memory
| 11 | Goepel leg supports with integrated handles and provision for vertical, horizontal and axial adjustment. The supports must have the sponge cover, tightening strap and stainless steel bars 18/10.
| 12 | The height adjustment of the pelvic section from 550mm to 1000mm
| 13 | The inclination of the back section 6 - 52º
| 14 | Nominal capacity 250kg
| 15 | Power supply 100-24V, 50/60Hz, 0,94kW
| 16 | Cushion at the head section attachable by magnet
| 17 | Protective cover of the seat section.
| 18 | Doctor`s foot support left-hand mounting, patient`s POV
| 19 | Leg plate integrated into the seat frame, extractable, padded, safe working load 50kg, 480 x 373 x 30mm
| 20 | Mobility on 4 swivel antistatic castors with the central brake
93 |  | Infant Bassinet |  | 10 |  | 0
| 1 | Telescoping pedestal raises bassinet from bedside to staffside
| 2 | Large clear tilting bassinet for full view of infant
| 3 | Easy to clean surface
| 4 | 4-wheel locking casters
| 5 | Load capacity: 10kg (on mattress)
| 6 | Mattress and mattress sheet
| 7 | Shade
| 8 | I.V. Pole
| 9 | Drawer
| 10 | Basket
| 11 | Cylinder Rack
94 |  | Resuscitation table |  | 4 |  | 0
| 1 | Main Unit at the stand with four locking casters
| 2 | Mattress high position: 900mm +/-10%
| 3 | Automatic temterature control: Skin temperature setting range: 34-38°C Skin temperature display range: 30-42°C
| 4 | APGAR and CPR timer are available
| 5 | Heater power adjustable in minimum 10 levels
| 6 | Warming head: rotation minimum +/-85° or more
| 7 | Mattress tilting
| 8 | Built in LED lighting lamp
| 9 | X-ray cassette tray
| 10 | Alarms: power failure, skin temperature probe and system error
| 11 | Degree of protection: Type BF applied part
| 12 | Warming Unit have to be delivered with all accessories needed, including: - Skin temperature probe reusable, 1pc - Mattress sheet, 1 pc - Dust cover, 1 pc - Operation manual, 1 pc
| 13 | Resuscitation external module: Pressure gauge indication: -20 to 80cmH2O Pressure limit: 20 to 60cmH2O Oxygen concentration control: 21-100% Flowmeter adjusment: 0-15l/min Accessories: Patient circuit with T-flow safety valve (40 pcs), Test bag, AIR and O2 connecting hoses, Infant face masks size 1 (20 pcs) and size 2 (20 pcs)
| 14 | External suction unit, venturi (ejector) type: Continuous vacuum regulator (analog gauge) Vacuum level: 0-250 mbar Horizontal rotation of the vacuum gauge from -45° to +45° for a better visibility Complete mounting system: rail clamp, AIR hose and DIN connection Accessories: 150 ml safety jar, 1 L polysulfone collection jar with rail mount, antibacterial filters (100 pieces), Suction tip for drainage (10 pieces)
95 |  | Infant incubator for intensive care unit |  | 4 |  | 0
| 1 | Continuous mode of operation
| 2 | Main Unit at high adjustable stand with four locking casters
| 3 | Mttress high position is adjusted: 800-1200mm or larger
| 4 | Mattress tilting
| 5 | Foot switches provided on both sides of the Hi-Low stand
| 6 | TFT LCD “touch screen“ display size: min 8 inches
| 7 | Automatic temterature control: Incubator air temperature setting range: 25-37C Incubator air temperature display range: 20-42°C Skin temperature setting range: 34-37.5°C Skin temperature display range: 30-42°C
| 8 | Heater power adjustable
| 9 | Warm-up time: 60min or less
| 10 | Automatic oxygen control: Oxygen concentration setting range: 22-65% Oxygen concentration display range: 19-105%
| 11 | Automatic humidity control: Relative humidity setting range: 40 – 95% Relative humidity display range: 15 – 99% Maximal relative humidity: 90% or more
| 12 | Pulse oximeter is integrated: SpO2 display range: 10-100% Pulse rate display range: 30-240bpm
| 13 | Noise level: maximum 45dB
| 14 | Double lock system of each admittance panel
| 15 | An X-ray cassette can be inserted and removed without opening an admittance panel so that the incubator air temperature will not drop
| 16 | Alarms: high temperature, power failure, set humidity, skin temperature probe, oxygen concentracion, SpO2 limit, pulse rate limit and system failure
| 17 | Dual wheel caster for more stability of the unit with four locking casters
| 18 | Degree of protection: Type BF applied part
| 19 | Incubator have to be delivered with all accessories needed, including: - Skin temperature probe reusable, 1pc - Infant SpO2 probe reusable, 1 pc - Dust cover, 1 pc - Operation manual, 1 pc
| 20 | External Phototherapy Unit at flexibile arm with attachment for incubator stand Examination LED light built-in to monitor the infant's condition Radiation level: 30 - 40 uW/cm2, at distance of 30cm Light illumination can be adjusted: 500 - 2000Lux, at distance of 30cm AC power source: AC220-240V 50/60Hz Power consumption: max 30VA
96 |  | General purpose patient monitor , neonatal |  | 3 |  | 0
| 1 | TFT LCD color display, min. 10", touchscreen
| 2 | Simultaneous display of min. 5 waveforms and 10 numeric parameters
| 3 | Connections for monitoring of ECG, Respiration, SpO2, NIBP, Temperature
| 4 | Connection for monitoring of IBP
| 5 | Connection for intubated and nonintubated capnography
| 6 | Pacemaker signal detection and rejection
| 7 | Automatic ST segment level detection, alarming and storing
| 8 | Automatic arrhythmia detection, alarming and storing: min. 10s per episode, total min. 72h
| 9 | Graphic and numerical parameters trend, min. 72h
| 10 | ECG full disclosure min. 72h
| 11 | Visual and sound alarm, 3 priority levels, adjustable limits, storage for min. 72h
| 12 | Continuous noninvasive or minimally invasive monitoring of estimated hemodynamic parameters CO and SV
| 13 | ECG sync sound with pitch variable according to current SpO2 or IBP value
| 14 | Choice for enlarged numerical display mode, readable from long distance
| 15 | Defibrillation proof patient connectors
| 16 | Holder for fixation to patient bed and medical rail
| 17 | Active connection to central monitoring system, interbed monitoring and wireless telemetry
| 18 | Onscreen operation guide with illustrations
| 19 | Holder for fixation to medical rail, 1 pc
| 20 | To be delivered with all accessories needed, including:  ECG cable with 3 leads, 1 pc  Reusable waterproof SpO2 probe, clip type, 1 pc  Reusable waterproof SpO2 probe, neonatal, 1 pc  Reusable NIBP cuff, medium, 1 pc  Reusable NIBP cuff, small, 1 pc  Reusable NIBP cuff, extra small, 1 pc  NIBP hose, neonatal, 1 pc  Disposable NIBP cuff, extra small, 10 pcs  Disposable ECG electrodes, neonatal, 120 pcs  Accessories for continous noninvasive or minimally invasive monitoring of CO and SV, 1 set  Battery, 1 pc
97 |  | Intensive care mechanical ventilator for adults, pediatric and neonatal application |  | 1 |  | 0
| 1 | Intensive care mechanical ventilator for adult, pediatric and neonatal patients, bodyweight > 0.2kg.
| 2 | Ventilator based on dedicated trolley with castors and brakes, with arm for patient circuit.
| 3 | Power supply AC 220V/50Hz, with integrated battery backup for minimum 60 min.
| 4 | Standard connection for Air and O2 high pressure central gas supply.
| 5 | TFT LCD color touchscreen, min. 15", with tilt and pan adjustment.
| 6 | Ventilation modes needed: - Volume controlled CMV and pressure controlled CMV - Volume controlled SIMV and pressure controlled SIMV - Pressure support - BiPAP (or equivalent) and APRV - NIV with guaranteed minimal frequency - High-Flow O2 therapy - Neonatal nCPAP therapy
| 7 | Ventilation parameters minimal settings range: Pinsp: 5 - 90 mbar Psupport: 0 - 90 mbar PEEP: 1 - 50 mbar I:E: 1:6 do 4:1 f: 1 - 150 bpm FiO2: 21 - 100% Vt: 2 - 2000 ml
| 8 | Reversible apnea backup function (off when apnea is over).
| 9 | Flow trigger and pressure trigger, adjustable.
| 10 | Tube Resistance Compensation and Leak Compensation.
| 11 | Inspiration hold and Expiration hold functions.
| 12 | Preoxygenation and Manual breathe functions.
| 13 | Capability to drive and control pneumatic nebulizer.
| 14 | Upgradeable with electronic nebulizer, controlled by ventilator.
| 15 | Upgradeable with integrated ET cuff control system: cuff pressure setting, monitoring and alarming via ventilator display.
| 16 | Visual and audible alarms with priorities and adjustments of alarm limits.
| 17 | Alarm lamp on top of ventilator, visible from all directions (360°).
| 18 | Capability to display min. 8 simultaneous waveforms, loops and trends.
| 19 | Monitoring trends min. 96h.
| 20 | Event log for alarms and ventilation settings.
| 21 | Monitoring parameters: - Tidal volume, spontaneous MV and total MV - Spontaneous frequency, total frequency, TI, TE - Minute volume of leak - Ppeak, Pmean, Pplato, AutoPEEP, P01 - Insp. and exp. airway resistance, lung compliance, PTP, RC and RSB - Oxygen concentration
| 22 | Ventilator must have closed loop ventilation mode with MV guaranteed, with automatic adaptation of ventilation parameters according to current patient respiratory mechanics (R, C and activity). With automatic protective limits, preventing barotraumas, autopeep and deadspace ventilation. This adaptive ventilation should work automatically: starting with intubation, through weaning, until extubation. Intended for both intubated and nonintubated patients, for obstructive, restrictive or healthy lungs.
| 23 | Program tool for automated lung assessment and recruitment maneuver. Recording of complete quasi static compliance curves (inflation and deflation), for evaluation of hysteresis, with adjustable maneuver parameters and pause for recruitment. Cursor for measurements and finding optimal pressures (PEEP, Pplato). Saving history of loops and reference loops.
| 24 | Ventilator must have graphic animation for visualization of lung dynamics (Vt, R, C) and patient activity, as well as graphic animation for visualization of patient - ventilator dependency (oxygenation, CO2 elimination and spontaneous breathing) with adjustable weaning criteria.
| 25 | Proximal flow sensor.
| 26 | Auxiliary pressure input for esophageal (transpulmonary) pressure measurement.
| 27 | Output connection for external video screen.
| 28 | Volumetric capnography: graph FCO2 and PCO2 vs. V, measurement of etCO2, Vdeadspace, VTalveolar and display on ventilator's screen.
| 29 | Function of screenshots saving to a memory card.
| 30 | Warm humidifier with manual and automatic temperature / humidity controll and watter filling.
| 31 | To be delivered with all accessories needed, including: - Patient breathing sets for adult/pediatric, with humidifier chamber, 15 pcs - Patient breathing sets for neonatals, with humidifier chamber, 15 pcs - Flow sensor adult/pediatric, 15 pcs - Flow sensor neonatal, 15 pcs - NIV masks, three different sizes, 3 pcs - "High Flow" O2 therapy cannulas, pediatric/neonatal, 30 pcs - Accessory for nCPAP therapy, 20 pcs - Accessory for volumetric capnography, for neonatal and adult/pediatric, 1 set
98 |  | Bedside pulse oximeter |  | 4 |  | 0
| 1 | Touch Screen bedside Pulse oximeter
| 2 | Pulse oximeter with handle, allowing the continuous measurement of SpO2, PR, Perfusion index.
| 3 | With built-in Wifi capacity, Li-Ion rechargeable battery and built-in integrated charger.
| 4 | The Technology needs to be motion and low perfusion resistant, with possibilitty to upgrade to noninvasive total hemoglobin and oxygen reserve index capacity
| 5 | Accessories that must me supplied: Mains power cable (210-240V) 1pc Extension cable for SpO2 sensor 1pc Single use SpO2 sensor Neo/Adult 20pcs Single use SpO2 sensor NeoPt 20pcs Single use SpO2 sensor Pediatric 20pcs
99 |  | Steriliser |  | 2 |  | 0
| 1 | Temperature settings in °C: 5°C above ambient up to +250°C
| 2 | Multifunctional digital PID-microprocessor controller
| 3 | HD TFT colour display
| 4 | Digital counter with target time setting
| 5 | adjustable temperature in Celsius or Fahrenheit, adjustable air flap position, adjustable summertime/wintertime
| 6 | Natural convection
| 7 | Adjustment of pre-heated fresh air mixture by air flap control ( 10% steps)
| 8 | Vent connection with restrictor flap
| 9 | Programme stored in case of power failure
| 10 | Safety of temperature control - protection class 1 according to DIN12880
| 11 | Autodiagnostic system for fault analysis
| 12 | Chamber volume at least 70 liters
| 13 | Maximal loading of chamber 120 kg
| 14 | Interior made of stainless steel, easy to clean
100 |  | Mobile plazma air desifenction and purificaton device |  | 5 |  | 0
| 1 | Decontaminating unit should be intended for rapid remediation in large spaces and situations with high risk of infection, using plasma technology combined with a triple-stage filter system to provide a combined solution for air disinfection and particle removal. Must be on 2-wheeled base + handle, speed fan setting
| 2 | Range of area coverage:
| 3 | minimum: 20 – 90 m2 ( maximum: 4 air changes per hour)
| 4 | minimum: 40 – 180 m2 (maximum: 2 air changes per hour)
| 5 | Volume treatment:
| 6 | minimum:45 – 220 m3 ( maximum: 4 air changes per hour)
| 7 | minimum: 90 – 450m3 (maximum: 2 air changes per hour) or similar
| 8 | Electrical rating
| 9 | Single phase, 230 VAC, 50 Hz
| 10 | Air flow
| 11 | minimum range:180 – 900 m3/hr (107 – 533 CFM)
| 12 | Emloyed filters:
| 13 | Stage 1: prefilter
| 14 | Stage 2: HEPA filter
| 15 | Stage 3: carbon plated filter
101 |  | Phototherapy lamp |  | 2 |  | 0
| 1 | Phototherapy Unit at floor stand with casters
| 2 | Examination LED light built-in to monitor the infant's condition
| 3 | Radiation level: 30 - 40 uW/cm2, at distance of 30cm
| 4 | Light illumination can be adjusted: 500 - 2000Lux, at distance of 30cm
| 5 | AC power source: AC220-240V 50/60Hz
102 |  | General purpose patient monitor |  | 1 |  | 0
| 1 | TFT LCD color display, min. 10", touchscreen
| 2 | Simultaneous display of min. 5 waveforms and 10 numeric parameters
| 3 | Connections for monitoring of ECG, Respiration, SpO2, NIBP, Temperature
| 4 | Connection for IBP monitoring
| 5 | NIBP measurements triggered by changes in pulse wave propagation time
| 6 | Pacemaker signal detection
| 7 | Automatic ST segment level detection, alarming and storing
| 8 | Automatic arrhythmia detection, alarming and storing: min. 10s per episode, total min. 120h
| 9 | Graphic and numerical parameters trend, min. 120h
| 10 | ECG full disclosure min. 120h
| 11 | Visual and sound alarm, 3 priority levels, adjustable limits, storage
| 12 | Continuous noninvasive monitoring of estimated hemodynamic parameters CO and SV
| 13 | ECG sync sound with pitch variable according to current SpO2 and IBP value
| 14 | Choice for enlarged numerical display mode, readable from long distance
| 15 | Power supply AC 220V/50Hz with battery backup min. 3h
| 16 | Defibrillation proof patient connectors
| 17 | Onscreen operation guide with illustrations
| 18 | Each monitor to be delivered with the following accessory:  ECG cable with 3 leads, 1 pc  Reusable waterproof SpO2 probe, clip type, 1 pc  Reusable NIBP cuff, large, 1 pc  Reusable NIBP cuff, medium, 1 pc  Accessory for continous noninvasive monitoring of CO and SV, 1 set - Holder for monitor, fixation to medical rail or wall, 1 pc - Battery, 1 pc
103 |  | Wall supply unit |  | 1 |  | 0
| 1 | Wall unit-1piece
| 2 | For patient preparation-Wall installation channel for one bed, total length of 1.7m with rails for carrying the device.It is used to supply the device with electricity and medical gases. It is made according to the principle of connecting the basic sockets into one whole with the central place of supply of all necessary installations.
| 3 | Consists of :
| 4 | Double alluminum profile L=1.700 mm, totoal height of 275mm, with closed ends
| 5 | Electrical power: -Schuko outlet 230V, generator (red)-4pieces, Schuko outelet 230V, UPS (green)2 pieces, Equipotenital bonding (connectors)-4 pieces
| 6 | Medical gases:O2(1piece), KV5(2pieces), N20(1piece), AGSS with protective cover
| 7 | Equpiement carriers-box rail 25x10mm, upper rail L=1,7 m, lower rail L=1,7 m
| 8 | Lower wall "dirty" box rail made of stainless steel, total length L=1,7 m for carrying accessories and equipment, complete with end fasteners and wall brackets, it is placed at a height of 0.5 m from the level of the finished floor
| 9 | Wall set for anesthesia(MRI and Angio room)-2pieces
| 10 | Wall installation channel for one bed, total length of 1 m. It is used to supply the device with electricity and medical gases (without power sockets). It is made according to the principle of connecting the basic sockets into one whole with the central place of supply of all necessary installations
| 11 | Consists of:
| 12 | Double alluminum profile L=1.000 mm, totoal height of 352mm, with closed ends
| 13 | Medical gases:O2(2pieces), KV5(2pieces), N20(1piece), AGSS with protective cover(1piece)
| 14 | Testing of the installation for strength and impermeability according to the valid standards SRPS EN ISO 7396-1
| 15 | Marking of pipelines for each gas separately with the name of the type of gas that is placed directly on the pipe and marking the route of the pipeline marked or on the wall just below the suspended ceiling "MEDICAL GASES" that are placed on the ceiling