## PriceSchedule
Bidder: |  |  |  | Offered price:
10_Sterilization
Item No. | Equipment name | Qty | Model, Manufacturer and Country of Origine | Unit price | Total Price per line item
1 | Ultrasound counter | 1 |  |  | 0
2 | Pre Wash Couter | 1 |  |  | 0
3 | Stainless steel table | 3 |  |  | 0
4 | Stainless steel shelves | 20 |  |  | 0
5 | Automatic instrumets washer – disinfector machine | 2 |  |  | 0
6 | Pass through box | 1 |  |  | 0
7 | Bedpan washer-disinfctor | 1 |  |  | 0
8 | Packing station | 7 |  |  | 0
9 | Rotary heat sealer | 3 |  |  | 0
10 | Pouchmaking machine | 1 |  |  | 0
11 | Chair with backrest | 11 |  |  | 0
12 | Baskets with wall brackets | 10 |  |  | 0
13 | Paper trolley | 2 |  |  | 0
14 | Linen folding table with light | 1 |  |  | 0
15 | Cloth cutting machine | 1 |  |  | 0
16 | Automatic steam sterilizer 9 STU with authomatic unloading | 2 |  |  | 0
17 | Low temperature plasma sterilizer | 1 |  |  | 0
18 | Closed transport trolley for sterile material | 6 |  |  | 0
19 | Washing-disinfection machine for trolleys | 1 |  |  | 0
20 | Steam sterilizer (autoclave) | 2 |  |  | 0
21 | Worktop Ultrasonic Cleaner | 1 |  |  | 0
22 | Laboratory autoclave, capacity 110l | 1 |  |  | 0
23 | Steam sterilizer for steam disinfection of infectious waste | 1 |  |  | 0
24 | Water preparation system set |  |  |  | 0
Total price per Lot |  |  |  |  | 0

## TechSpecification
Bidder: | [5 empty] | Total DAP price: | 0
10_Sterilization
Item No. | ID | Technical Specification Requested | Technical Specification Offered | Qty | Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Ultrasound counter |  | 1 |  | 0
| 1.1 | Manufactured from 304 grade stainless steel
| 1.2 | Two instrument washing sinks and one ultrasonic cleaner
| 1.3 | Three spiral faucets with adjustable shower spout
| 1.4 | Two water drainage system
| 1.5 | Two pieces air-water gun rinser with different tips
| 1.6 | All worktop edges are raised to prevent water drips
| 1.7 | Height adjustable stainless legs
| 1.8 | Dimension (WxLxH): approx. 2000x700x850 mm.
2 |  | Pre Wash Couter |  | 1 |  | 0
| 2.1 | Manufactured from 304 grade stainless steel
| 2.2 | One sink with triple manifold
| 2.3 | Two spiral faucets with adjustable shower spout
| 2.4 | Two water drainage system
| 2.5 | Extra drawer or cabinet options
| 2.6 | All worktop edges are raised to prevent water drips
| 2.7 | Height adjustable stainless legs
| 2.8 | 2 set air water spray gun with 8 nozzles,
| 2.9 | Dimension (WxLxH): approx. 1900x700x850 mm.
3 |  | Stainless steel table |  | 3 |  | 0
| 3.1 | Manufactured from 304 grade stainless steel
| 3.2 | Stainless steel adjustable feet
| 3.3 | Dimensions (WxLxH): approx. 600/700x1300x850 mm.
4 |  | Stainless steel shelves |  | 20 |  | 0
| 4.1 | Made from 304 grade Stainless Steel
| 4.2 | 4 castors with two brakes
| 4.3 | Solid shelve
| 4.4 | Modular construction, adjustable shelve intervals
| 4.5 | Dismountable shelves
| 4.6 | Dimensions (WxLxH): approx. 450X110x1800 mm.
5 |  | Automatic instrumets washer – disinfector machine |  | 2 |  | 0
| 5.1 | The capacity of the device min 10 DIN trays
| 5.2 | The device for washing and thermal disinfection and drying of instruments
| 5.3 | The device should be pass-through, with two doors.
| 5.4 | The chamber is made of high quality stainless steel AISI 316L or better
| 5.5 | Flow control for dosing units
| 5.6 | Wash pump, capacity min. 600 l/min
| 5.7 | Foam control
| 5.8 | Pump pressure monitoring
| 5.9 | High-capacity drain valve, min 5 l/s
| 5.10 | Electric heating of the chamber (water) power max 20 kW
| 5.11 | Ventilator for drying with HEPA filter
| 5.12 | Control with touch panel with LCD screen or LCD screen with dirty buttons, and LCD display on the clean side
| 5.13 | Disinfection by A0 - value
| 5.14 | Temperature monitoring during complete cycle
| 5.15 | The device must have possibility for PC and bar code reader connection by different interfaces
| 5.16 | Acoustic signal at program end
| 5.17 | Maintenance reminder
| 5.18 | Front covers made of stainless steel AISI 304 or better, with glass doors
| 5.19 | Dimensions (HxWxD): max. 1850x700x700 mm±50mm
| 5.20 | Equipment that has to be delivered with washer-disinfector:
| 5.21 | o Transport cart, 2 pc
| 5.22 | o Instrument rack for loading 10 standard instrument trays, 1 pc
| 5.23 | o Rack for AN material, 1 pc
| 5.24 | o Rack for containers for sterilization, 1 pc
| 5.25 | o Rack for OR shoes, 1 pc
| 5.26 | o MIS Rack
| 5.27 | EN ISO 15883
6 |  | Pass through box |  | 1 |  | 0
| 6.1 | Manufactured from 304 grade stainless steel
| 6.2 | Glass doors made of aluminum profiles
| 6.3 | Sealed door system preventing air passage
| 6.4 | Rounded inner chamber to prevent bacterial growth
| 6.5 | Magnetic lock system prevents to open two doors at the same time
| 6.6 | Visual and audio warning system in case of undesired operation
| 6.7 | Adjustable to keep the locks open during power failure
| 6.8 | Dimensions(WxLxH): 500x680x770 ±3% mm
7 |  | Bedpan washer-disinfctor |  | 1 |  | 0
| 7.1 | The device is used for emptying, cleaning and disinfection urine bottles, bed-pans and their lids and kidney dishes
| 7.2 | Device cladding, chamber for emptying and cleaning and chamber door, are made of stainless steel
| 7.3 | Chamber door are double walled, acoustically and thermally insulated, steam-tight, easy to replace
| 7.4 | Chamber door locking by program-controlled impulse transmitted to a maintenance-free magnet for closing and opening thus eliminating any danger to the operator in the event of a power failure
| 7.5 | Automatic monitoring of locking and unlocking the door
| 7.6 | Cleaning chamber including funnel and DN-100 drain socket deep drawn from a single steel sheet, with large radii and sloping chamber top
| 7.7 | Siphon monitoring, program stop in the event of siphon blockages in the cleaning chamber
| 7.8 | Insulation from the mains water according to DIN EN 1717 with air seal type AA, complying with European standards
| 7.9 | The device has RJ 11 infrared interface for data transfer to external software
| 7.10 | Energy saving mode, stand-by operation
| 7.11 | Plain text display with background illumination
| 7.12 | The device has optical and acoustic signal
| 7.13 | The device has separated commands for the program start for: short, normal, intensive washing and intensive chamber rim flushing
| 7.14 | Selectable programs
| 7.15 | Main rotary jet for internal cleaning of the care utensils, must be positioned outside the emptying zone, to avoid dirt/blockage
| 7.16 | The device has 12 rotating nozzles for external and internal washing
| 7.17 | Dosing technology with 1 dosing unit for a combined water softening and rinse products and one suction pipe for containers with level control
| 7.18 | Control by disinfection process control according to the pre-determined A0 value (according to EN ISO 15883-3: A0 = 60).
| 7.19 | Disinfection process control with two temperature sensors monitoring each other. The control software monitors and assesses the control process.
| 7.20 | Maximum noise level 49 LWA db
| 7.21 | Dimensions: width x depth x height: 500 x 450 x 1750 ±3% mm
8 |  | Packing station |  | 7 |  | 0
| 8.1 | Complete large set: 2 synchronous motor, 1 control box, 1 hand controller, 1 controller holder apparatus
| 8.2 | In all clinical and polyclinic environments that need to be sterile,
| 8.3 | 304 quality stainless steel,
| 8.4 | Height of the table can be adjusted in compliance with sedentary or on foot operation through synchronized height adjusted motors,
| 8.5 | User friendly design with flexible storage shelves or baskets,
| 8.6 | Height adjustable Blester shelf,
| 8.7 | Height and size adjustable Blester holding apparatuses,
| 8.8 | Stand design that does not affect sitting position of the user,
| 8.9 | Plug-socket and electric group, internal illumination lamp and flexible lens table lamb,
| 8.10 | Computer screen and cabinet holder apparatus,
| 8.11 | Two block drawers,
| 8.12 | Dimensions (WxLxH): 700x1800x900/1850 ±3%mm.
9 |  | Rotary heat sealer |  | 3 |  | 0
| 9.1 | Min 4 inch touch touch display for control
| 9.2 | can be linked to batch documentation and tracking systems
| 9.3 | USB, RS 232 and Ethernet interface
| 9.4 | All actions recorded and can be viewed if required
| 9.5 | Fully validatable according ISO 11607-2
| 9.6 | Suitable for sealing thermoplastic pouches, sealable paper pouches, sealable ponuches and reels made of film and paper according to EN 11607-1/EN868-4
| 9.7 | Microprocessor control
| 9.8 | Touch screen coulour display, min. 4“
| 9.9 | Sealing temperature max. 220ºC
10 |  | Pouchmaking machine |  | 1 |  | 0
| 10.1 | Width of seal seam and cutting lenghts, roll holder - 400 mm±20mm
| 10.2 | Suitable for use in medical practice, hospital and medical industry
| 10.3 | Adjustment pouch lengths 100 - 1,000 mm
| 10.4 | Adjustment quantity of pouches 1 - 500 pouches
| 10.5 | Sealing temperature max. 200 °C
| 10.6 | Contact pressure fix
| 10.7 | Have overheating protection
| 10.8 | LCD-Display
| 10.9 | Microprocessor controlled
| 10.10 | Interface USB A/B, RS 232, Ethernet connection
| 10.11 | Barcode reader
| 10.12 | Stainless steel AISI 304 casing cover
| 10.13 | Seal seam 12 mm
| 10.14 | Reel holder
| 10.15 | Sorting unit
| 10.16 | Sealing process automatic/reproducible
| 10.17 | Switch-off tolerance +/- 5°C (ISO/TS 16775)
| 10.18 | Alarm and machine stop in case the monitored parameters exceed predetermined limits
11 |  | Chair with backrest |  | 11 |  | 0
| 11.1 | Grade 304 Stainless Steel base
| 11.2 | Anti-static and anti bacterial seat
| 11.3 | Circular, square and saddle seat types
| 11.4 | Aluminium injection cast foot system, 5 castors.
| 11.5 | With backrest
12 |  | Baskets with wall brackets |  | 10 |  | 0
| 12.1 | Baskets intended for storage and transportation of non-sterile and sterile material
| 12.2 | Baskets are made of wire mesh;raster prevents standard elements from falling out.
| 12.3 | Square shape
| 12.4 | Made of stainless steel AISI 304
| 12.5 | Type: 1 STU,
| 12.6 | Dimensions: approximately 600x300x300 mm (LxWxH)
| 12.7 | Bracket for hanging varioous modular elements as well as sterilization baskets (1 bracket for 2 baskets)
| 12.8 | Made of stainless steel AISI 304
| 12.9 | With integrated hooks, min 2 hooks/bracket
| 12.10 | Tubing must be finished with plastic caps on both sides
| 12.11 | Carriers of 700 mm ± 5% length with interlocking hooks 350 mm ± 5%"
13 |  | Paper trolley |  | 2 |  | 0
| 13.1 | Made from 304 grade Stainless Steel
| 13.2 | 4 wheels with 2 breaks
| 13.3 | 4 hanger bars
| 13.4 | Dimensions (WxLxH): approximately 600x1200x900 mm.
14 |  | Linen folding table with light |  | 1 |  | 0
| 14.1 | Made from 304 grade Stainless Steel
| 14.2 | Illuminated transparent top surface
| 14.3 | Grounded socket with key
| 14.4 | Dimensions (WxLxH): 900x1400x 900 ±3% mm.
15 |  | Cloth cutting machine |  | 1 |  | 0
| 15.1 | Cutting knife diameter max 125 mm
| 15.2 | Power supply 220-230 V, 50Hz or 380-400 V, 50Hz
| 15.3 | Power of the device max 400W
| 15.4 | Capacity min 25 mm / 1 "(seconds)
16 |  | Automatic steam sterilizer 9 STU with authomatic unloading |  | 2 |  | 0
| 16.1 | Automatic steam sterilizer with horizontal electromechanical door
| 16.2 | The sterilizer should be pass-through, with two doors
| 16.3 | The sterilizer must be intended for water-steam sterilization (autoclaving) on the temperatures 121°C and 134°C
| 16.4 | The chamber is pass-through, horizontal, square cross-section. The chamber is adapted for easy loading, loading height max. 500 mm
| 16.5 | The capacity of approx. 650 l for 9 STU (standard sterilizing unit 600x300x300mm)
| 16.6 | Chamber and jacket material:
| 16.7 | Chamber made of stainless steel AISI 316L (CrNiMo) or better
| 16.8 | Jacket made of stainless steel AISI 316L or AISI 316Ti
| 16.9 | Construction and paneling made of stainless steel
| 16.10 | The chamber is insulated with high quality insulating material covered with galvanized metal sheet or better
| 16.11 | Automatic horizontal sliding door open with an electric motor
| 16.12 | Sterilizer is equipped with door security systems:
| 16.13 | o System against door opening during the cycle
| 16.14 | o System against opening both sterilizer doors at the same time
| 16.15 | Sterilizer has a double automatic system for protection of the user form door crushing
| 16.16 | Steam piping and valves are made of stainless steel, the pipes are heat insulated and colour-coded
| 16.17 | Sterilizer is equipped with two-step vacuum pump with water ring. Recirculation water saving system with Pt100 probe or better
| 16.18 | On the non-sterile side of the touch screen diagonal min 10", and on the sterile side of the appliance, a panel for opening and closing the door and light signaling for the status of the sterilizer is sufficient
| 16.19 | Sterilizer is guided by microprocessor that is programmable and has a possibility of PC and web connection (Ethernet).
| 16.20 | Sterilization process control by double Pt100 probe and double pressure transducer.
| 16.21 | Integrated printer that is showing the cycle parameters and graphs of the sterilization process.
| 16.22 | Cycles:
| 16.23 | o Instruments (134°C),
| 16.24 | o Textile (134°C),
| 16.25 | o Rubber (121°C),
| 16.26 | o Bowie & Dick Test (134°C),
| 16.27 | o Vacuum Test,
| 16.28 | o Preheating,
| 16.29 | o Programs on customer’s demand (defined by user and programmed upon delivery and installation).
| 16.30 | o Loading cart return program
| 16.31 | The sterilizer has integrated electronic Bowie&Dick test, in accordance with DIN EN ISO 11140 standard.
| 16.32 | The device must have its own steam generator total power maximum 80кW
| 16.33 | Steam generator and vacuum pump are external, installed in the designated location
| 16.34 | Power: 380V/50Hz, three phase
| 16.35 | Sterilizer conforms to the following standards
| 16.36 | DIN EN ISO 9001
| 16.37 | EN ISO 13485
| 16.38 | CE sign
| 16.39 | EN 285
| 16.40 | Directive MDD 93/42/EEC
| 16.41 | Directive PED 97/23/EC
| 16.42 | Dimensions of the sterilizer should be maximum HeightxWidthxDepth: 1950x1600x1300mm±50mm
| 16.43 | Sterilizer must be equipped with 1 loading cart and 2 transport trolley
| 16.44 | Sterillizer must be equipped with automatic unloading ramp for the automatic unloading of steam sterilizers with push-out device for loading carts.
17 |  | Low temperature plasma sterilizer |  | 1 |  | 0
| 17.1 | Sterilizer is intended for the sterilization of thermolabile materials at temperature of 50° C to 55° C
| 17.2 | The chamber is pass-through, square section, of the total volume min. 140 liters
| 17.3 | A system with two sliding doors
| 17.4 | Sterilant agent H₂O₂ (hydrogen peroxide) max. Concentration 50%
| 17.5 | No. of cycles per 1 sterilant agent bottle/cartridge: min. 10 standard cycles
| 17.6 | The device is mobile with 4 wheels with brakes which ensure stability during operation
| 17.7 | Built in printer
| 17.8 | Four sterilization programs:
| 17.9 | - Surface - cycle time max. 35 min.
| 17.10 | - Standard - cycle time max. 65 min.
| 17.11 | - Advanced - cycle time max. 70 min.
| 17.12 | - ENDO - cycle time max. 50 min.
| 17.13 | Self test-function
| 17.14 | Ethernet connection
| 17.15 | 2 external USB ports
| 17.16 | Power supply 220V / 50Hz
| 17.17 | The device does not require any additional supplys or installations except electricity.
| 17.18 | Managing and control via touch-screen monitor diagonal min. 10 "
| 17.19 | Sterilization of lumens:
| 17.20 | o Flexible: - Ø1x1.000 - min. 10 pcs
| 17.21 | o -Ø1x4.000 – 1 pc
| 17.22 | o Rigid: - Ø1x400 – min. 10 pcs
| 17.23 | Norms and standards:
| 17.24 | o 93/42 / EEC
| 17.25 | o ISO 13485
| 17.26 | o ISO 14971
| 17.27 | o ISO 14937
| 17.28 | o ISO 10993-5
| 17.29 | o ISO 15223-1, -2
| 17.30 | o IEC 60601-1-2
| 17.31 | o IEC 62366
| 17.32 | o IEC 62304
| 17.33 | Sterilizer and bottle with sterilant agent/ cartridge must be registered in ALIMS
| 17.34 | Dimensions: max. 1100X1700X1100 mm (WxHxD)
18 |  | Closed transport trolley for sterile material |  | 6 |  | 0
| 18.1 | Doors swing open up to 270 Double-walled, sealed air-tight locking doors
| 18.2 | Shelf system with 3 rows, height adjustable, removable wire shelves
| 18.3 | Double sided pushing handles Four castors with two brakes
| 18.4 | Made from 304 grade Stainless Steel
| 18.5 | Dimensions (WxLxH): 700 x 1250 x 1450 ±3% mm.
19 |  | Washing-disinfection machine for trolleys |  | 1 |  | 0
| 19.1 | Cleaning and disinfection system for handling uncritical and semi-critical medical devices as well as various other equipment used in hospitals
| 19.2 | Pass through chamber with 2 doors
| 19.3 | Chamber material stainless steel (1.4301, AISI 304) with smooth surface
| 19.4 | Chamber wall insulation using low-pollutant mineral wool clad in aluminium
| 19.5 | Large angle inclination in chamber ceiling and floor for improved drainage
| 19.6 | Manual removal of retention filters in the chamber for quick cleaning of the circulation system
| 19.7 | Self-cleaning program
| 19.8 | Emergency open switch within the chamber
| 19.9 | Chamber draining due to drain pump
| 19.10 | Horizontally sliding doors with double-layered safety glass, vacuum insulated
| 19.11 | Safety device to prevent opening during the program cycle
| 19.12 | Door closing edge secured due to integrated motor force limiter.
| 19.13 | Doors are locked to prevent opening at the same time
| 19.14 | All pipes are self-draining due to slanting, with minimum dead space
| 19.15 | All pipes are made of stainless steel 1.4301 or higher
| 19.16 | All valves are made of stainless steel
| 19.17 | Nozzles are screwed into vertical pipes fitted on each side
| 19.18 | The nozzle pipes pivot vertically on their own axes and reliably ensure the equipment and washing system is heated up gradually.
| 19.19 | Nozzle pipes are attached on both sides of the wall, with a drive for each side.
| 19.20 | Closely spaced stainless steel flat jet nozzles
| 19.21 | The nozzle system is fully drained automatically after every step of the program
| 19.22 | The circulation pump made of stainless steel 1.4571 (AISI 316 Ti)
| 19.23 | Stainless steel medium pressure fan ensures high level of air circulation in the chamber and excellent drying
| 19.24 | Air temperature controlled by analogue Pt100 and can be adapted according to the temperature sensitivity of the items to be treated
| 19.25 | Two large peristaltic dosing pumps
| 19.26 | Precise flowmeters for each agent
| 19.27 | Front cover for the unit with polished stainless steel panels with 1.4301
| 19.28 | The following equipment can be treated: - Trolleys for various purposes - Instrument containers with lids - Surgical shoes - Kidney dishes - Washing bowls, salt dishes - Plastic boxes with lids - Various non-invasive items
| 19.29 | Water and air are heated by electric power. Max power 130 kW.
| 19.30 | Two different types of water connections: soft cold water and DI water
| 19.31 | Dosing from canister 2x. Dispensing of the process chemical and its machine processing is done by canisters
| 19.32 | Exhaust air fan made of stainless steel. Draught power 900 m3/h at 300 Pa counter pressure
| 19.33 | Waste water cooling
| 19.34 | Built-in plotter
| 19.35 | Loading and unloading ramps made of stainless steel 1.4301.
| 19.36 | Machine is delivered with the following trolleys: 1. Loading trolley for sterile material containers made of metal for processing together with their lids, capacity 12 containers 1/1, 1 pc 2. Loading trolley for OR shoes, capacity min 25 pairs of shoes, 1 pc
| 19.37 | Outer dimensions: max. 2500 x 3000 x 2050 mm (H x W x D)
20 |  | Steam sterilizer (autoclave) |  | 2 |  | 0
| 20.1 | Two sterilization temperatures:121°C and 134°C
| 20.2 | Cycles - Three universal sterilization cycles “ B” type - Three test cycles: Vacuum test, Bowie & Dick test/ Helix test ,Spore test - “Prion” cycle - Delayed start cycle
| 20.3 | Automatic door
| 20.4 | Display
| 20.5 | Additional drying at the end of the cycle
| 20.6 | Operator code
| 20.7 | Automatic water filling by pump
| 20.8 | SD Card
| 20.9 | Possibility of automatic water filling by depurator connection
| 20.10 | Printer
| 20.11 | Labeller connection
| 20.12 | Water softener delivered with the sterilizer
| 20.13 | Chamber volume: 18 l
| 20.14 | Chamber dimensions: min. Ø236 mm
| 20.15 | Outer dimensions: 470 x 400x 630 mm ± 5%
21 |  | Worktop Ultrasonic Cleaner |  | 1 |  | 0
| 21.1 | Manufactured from 304 grade stainless steel
| 21.2 | Water drainage tap
| 21.3 | Inluding water tank, wire basket and lid
| 21.4 | 220-240 V/50 Hz, Ultrasonic frequency 28 Khz
| 21.5 | Dimension (WxLxH): 320x520x470 mm ± 5%
22 |  | Laboratory autoclave |  | 1 |  | 0
| 22.1 | Intended for sterilization of media, solutions and decontamination of waste material
| 22.2 | Vertical sterilization chamber made of stainless steel
| 22.3 | Working area dimensions: Ø400 mm, depth 850 mm
| 22.4 | Working capacity min 100 l
| 22.5 | External dimensions_ WxHxD 700mm x 1100mm x 550 mm ± 5%
| 22.6 | Device should have castors, so it can be moved
| 22.7 | External clading made of stainless tin
| 22.8 | 3 stainless steel baskets Ø380mm, depth 250 mm
| 22.9 | 6 programs, 2 of which are programmable
| 22.10 | Temperature and pressure are shown on display. Pressure is shown on manometer, as well.
| 22.11 | Display is showing process phases
| 22.12 | Possibility of temperature setting between 105 to 135°C, sterilization time setting up to 60 minutes.
| 22.13 | Possibility of choosing the way of distilled water drainage
| 22.14 | Alarms: high temperature, low temperature, probe, time
| 22.15 | Independently installed thermostat and attested valve for protection of the device and user
| 22.16 | Technical documentation in accordance with EN 13455 to be delivered with the device
23 |  | Steam sterilizer for steam disinfection of infectious waste |  | 1 |  | 0
| 23.1 | Steam sterilizer for Disinfection of infectious waste
| 23.2 | Standards and norms: Pressure Equipment Directive (PED) 2014/68/EU Machinery Directive 2006/42/EC The CE mark is in accordance with the Pressure Equipment Directive 2014/68/EU DIN 58951-2 DIN EN 61326-1 DIN EN 60204 DIN EN 61010-2-040 DIN EN 62304
| 23.3 | Outer frame and front panel material CrNi steel no. 1.4301 (AISI 304) or adequate
| 23.4 | Double-walled chamber, internal walls of CrNiMo steel 1.4404 (AISI 316L) and steam jacket of CrNiMoTi steel, material no. 1.4571 / 1.4404 (AISI 316 Ti / AISI 316L); Pressure vessel must be insulated with high quality heat insulation
| 23.5 | Chamber volume min 450 l
| 23.6 | Single door sterilizer (non-pass through), horizontal opening
| 23.7 | Automatic, electrically driven, sliding door, door plate of CrNiMo steel, material no. 1.4404 or adequate Must be equipped with safety devices to prevent opening during the program run and to prevent entrapment.
| 23.8 | Piping design as per DIN EN 285, steam pipes and valves of CrNi steel; pipes colour-coded and insulated.
| 23.9 | Two-stage water ring vacuum pump with upstream plate heat exchanger and water circulation tank. Optimized water saving by PT-100 control.
| 23.10 | Increased safety by holding back and sterilizing of the condensate generated during sterilization of contaminated goods that are processed for discard.
| 23.11 | Circulation pump for vacuum device. Fresh water flow by means of "free outlet" according to class 5 of DIN EN 1717.
| 23.12 | According to DIN EN 285, the measuring, control and regulation of process relevant parameters, such as temperature and pressure, are effected by two resistance temperature probes PT100 and absolute pressure transducers.
| 23.13 | Programs: Waste program 134°C acc. RKI Solid goods 134°C Solid goods 121°C Porous goods 134°C Porous goods 121°C Filter 121°C Vacuum test
| 23.14 | It should be possible to modify all program parameters in the service menu to be accessed through access code within the limits of minimum and maximum values and within the range of job and product safety.
| 23.15 | Chamber jacket cooling by demineralized fresh water. Inlet through jacket nozzles, discharge thermostatically controlled.
| 23.16 | Integrated operating panel attached to the front of the sterilizer, min diagonal 10“
| 23.17 | Thermal printer for continuous recording of operating temperature and pressure and print-out of all relevant cycle parameters
| 23.18 | Validated filter for the outgoing air designed for an absolute filter transmittance of 0.003 micrometres (gaseous condition); filtration of the germs in the air removed from the chamber during the program cycle. Automatic in-line sterilization of the filter during every cycle by treatment with steam at process temperature. The filter cartridge must be made of hydrophobic and steam resistant PLTFE or adequate. Pressure tight housing made of ni-chrome molybdenum steel, material No.1.4404 (AISI 316 L). Removable upper part for the easy replacement of the filter.
| 23.19 | Integrated electric steam generator, power max 55 kW
| 23.20 | External dimensions: max. 1900x1600x1300 mm±50mm
| 23.21 | Power supply 380 V, 50 Hz.
| 23.22 | Accessories: - Loading trolley, 1 pc - Transport cart, 1 pc - Water preparation system of adequate capacity (softener + RO).
| 23.23 | Accessories: - Loading trolley, 1 pc - Transport cart, 1 pc - Water preparation system of adequate capacity (softener + RO).
24 |  | Water preparation system set |  | 1 |  | 0
| 24.1 | The head and body of the filter are made of PP. The filter is supplied with a washable filter cartridge with a fineness of 80 μm. The inlet and outlet connections are internally threaded and made of brass. The filter has 1" connections. Fineness of filtration is 80 μm.
| 24.2 | Duplex automatic softener consisting of two softeners in alternating operation. Their operation is managed by automatic valves with a microprocessor controller.
| 24.3 | Technical details of duplex softener:
| 24.4 | column volume min. 170L
| 24.5 | Quantity of ion exchange mass: min 2x125 l
| 24.6 | Softener capacity: min 440 m3 x 1º
| 24.7 | Working flow rate: 5000 l/h
| 24.8 | Number of NaCl solution tanks: 2
| 24.9 | Volume of NaCl solution tank: min 190 l each
| 24.1 | NaCl consumption: max. 20 kg/rageneration
| 24.11 | Inlet water connection: 1 1/2“
| 24.12 | Inlet water pressure: 2,5 bar
| 24.13 | Soft water connection: 1“
| 24.14 | Drainage connection: 3/4“
| 24.15 | Working pressure 2-8 bar
| 24.16 | Working temperature range min. 1-35º
| 24.17 | After the softener, a 10" long filter housing with a 20 μm filter cartridge should be installed
| 24.18 | Active carbon filter for dechlorination and water deodorization, i.e. removal of substances that can affect its organoleptic properties such as taste, smell and color. The filter with 1" connections. Filter cartridge 20"
| 24.19 | RO device containing a fine inlet filtration of 5 μm, a low pressure feed water switch, a high pressure switch for permeate, rotameters for measuring the flows of permeate and concentrate, a conductometer for measuring the conductivity of the permeate with the possibility of adjustment alarm at a certain conductivity value. The PLC controller of the device flushes the RO membranes automatically after each stop and start-up. During system rest, periodic flushing should be performed in order to prevent possible microbiological problems.
| 24.2 | Technical details of RO device:
| 24.21 | Working flow of permeate: min 500 l/h
| 24.22 | Working flow of concentrate: 300-500 l/h
| 24.23 | Working pressure 10 bar
| 24.24 | Device must have multistage centrifugal pump
| 24.25 | Number of RO membranes 2
| 24.26. | Purified water tank with distribution pump: The tank made of polypropylene, volume min 1000 l. Dimensions (height x length x width) 1310 x 1400 x 670 mm. The device must have booster pump; 0.55kW; 5-50 l/min; frequency regulated. The pump must be equipped with protection against dry operation.
| 24.27 | After the purified water tank there must be 20“ long filter, filtration 5μm