## Price Schedule
Bidder: |  |  |  | Offered price (currency): | 0
LOT 1 | Surgery
ID | Equipment name (Technical Specification Requested) | QTY | Technical Specification Offered (Model) | Estimated value per unit | Estimated value Total
1.1 | Surgical instruments
1.1.1 | Instrument set for Caesarean section | 4
1.1.2 | Instrument set for Laparatomy 1 | 4
1.1.3 | Abdominal retractor holding set | 1
1.1.4 | Chest set | 2
1.1.5 | Instrument set for Laparatomy 2 | 5
1.1.6 | Laparoscopic instruments set | 3
1.1.7 | Amputation instruments set | 2
1.1.8 | Set for breast surgery | 2
1.1.9 | Set for goiter surgery | 2
1.1.10 | Urology set | 1
1.1.11 | Sterilization containers and baskets, SET | 1
1.1.12 | Pouch making machine | 1
1.2 | Telescopes for ENT (set) | 1
1.3 | Ultrasound surgical system | 1
1.4 | EUS | 1
1.5 | Uretroton | 1
1.6 | Bipolar resection | 1
1.7 | Lithotripsy device | 1
1.8 | 4K laparoscopic tower | 1
1.9 | Operating table- gynecology/urology | 2
1.10 | Surgical lights
1.10.1 | Surgical lights | 1
1.10.2 | Surgical lights with camera | 1
1.11 | Anasthesia system monitor | 1
1.12 | Anaesthesia monitor | 3
|  |  |  | Total Price per Lot: | 0

## LOT 1 - Surgery
Bidder:
LOT 1 |  | Surgery |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1.1 |  | Surgical instruments |  |  |  | 0
| 1.1.1 | Instrument set for Caesarean section |  | 4 |  | 0
| 1 | Round bowl 60x30mm, 0,07l |  | 2
| 2 | Round bowl 150x65mm, 0,70l |  | 1
| 3 | Fritsch retractor, handle vertical to the blade, 60x85mm, 26cm |  | 1
| 4 | Roux retractor double-ended 31x30mm/36x42mm 17cm |  | 2
| 5 | Backhaus towel clamp, 13,5cm |  | 12
| 6 | Dissecting forceps standard 13cm |  | 2
| 7 | Dissecting forceps standard 18cm |  | 2
| 8 | Tissue forceps , standard model 1x2 teeth 13cm |  | 2
| 9 | Tissue forceps , standard model 1x2 teeth 18cm |  | 2
| 10 | Dissecting forceps with tungsten carbide inserts, 0,5mm, 14,5cm, standard, gold handles |  | 1
| 11 | Hegar uterine dilator with sloped handle, diam.14mm, 18,5cm |  | 1
| 12 | Hegar uterine dilator with sloped handle, diam.15mm, 18,5cm |  | 1
| 13 | Hegar uterine dilator with sloped handle, diam.16mm, 18,5cm |  | 1
| 14 | Hegar uterine dilator with sloped handle, diam.17mm, 18,5cm |  | 1
| 15 | Hegar uterine dilator with sloped handle, diam.18mm, 18,5cm |  | 1
| 16 | Hegar uterine dilator with sloped handle, diam.19mm, 18,5cm |  | 1
| 17 | Hegar uterine dilator with sloped handle, diam.20mm, 18,5cm |  | 1
| 18 | Hegar uterine dilator with sloped handle, diam.21mm, 18,5cm |  | 1
| 19 | Hegar uterine dilator with sloped handle, diam.22mm, 18,5cm |  | 1
| 20 | Hegar uterine dilator with sloped handle, diam.23mm, 18,5cm |  | 1
| 21 | Hegar uterine dilator with sloped handle, diam.24mm, 18,5cm |  | 1
| 22 | Dissecting scissors with welded tungsten carbide blades, curved, 18cm, with stable 3-point positions, ergonomically designed scissors rings with V-shaped handle |  | 1
| 23 | Lexer dissecting scissors with tungsten carbide edges, curved, 16cm, gold handles |  | 1
| 24 | Mayo dissecting scissors with tungsten carbide edges, straight, 14,5cm, with gold handles |  | 1
| 25 | Bush umbilical cord scissors, 16cm |  | 1
| 26 | Wertheim uterine scissors, curved, 20cm |  | 1
| 27 | Surgical scissors, slender pattern, straight, 12,5cm |  | 1
| 28 | Pean haemostatic forceps, delicate pattern, curved, 13cm |  | 6
| 29 | Rochester-Pean haemostatic forceps, curved, 18,5cm |  | 6
| 30 | Mayo-Hegar needle holder with tungsten carbide inserts, heavy model, 0.5mm, 18.5cm, gold handles |  | 2
| 31 | Needle sterilizing exterior and interior box, 65x15mm |  | 1
| 32 | Scalpel handle no. 4 length 130 mm |  | 1
| 33 | Scalpel blade remover to remove the blades from handle without the risk of injury, 14cm |  | 1
| 34 | Gross sponge and dressing forceps, straight, 20cm |  | 2
| 35 | Gross sponge and dressing forceps, curved, 20cm |  | 2
| 36 | Foerster-Ballenger dressing and sponge holding forceps serrated jaws 24,5cm straight |  | 2
| 37 | Foerster-Ballenger dressing and sponge holding forceps serrated jaws 24,5cm curved |  | 2
| 38 | Kocher-Ochsner haemostatic forceps straight 16cm 1x2 teeth |  | 2
| 39 | Mikulicz peritoneum forceps, 1x2 teeth, 20cm |  | 2
| 40 | Foerster-Ballenger dressing and sponge holding forceps serrated jaws 18cm straight |  | 6
| 1.1.2 | Instrument set for Laparatomy 1 |  | 4 |  | 0
| 1 | Round bowl 115x50mm, 0,35l |  | 2
| 2 | Round bowl 150x65mm, 0,70l |  | 1
| 3 | Gosset abdominal retractor 150mm, lateral valves 65x30mm and 65x35mm, with central valve 60x55mm |  | 1
| 4 | Doyen abdominal retractor, 115x60mm, 23cm |  | 1
| 5 | Fritsch retractor, handle vertical to the blade, 60x85mm, 26cm |  | 1
| 6 | Roux retractor double-ended 30x30mm/35x40mm, 17cm |  | 2
| 7 | Surgical scissors, standard, straight, 13cm, sharp/blunt |  | 1
| 8 | Nelson-Metzenbaum dissecting scissors with tungsten carbide edges, delicate model, curved, 26cm, gold handles |  | 1
| 9 | Dissecting scissors with welded tungsten carbide blades, curved, 18cm, with stable 3-point positions, ergonomically designed scissors rings with V-shaped handle |  | 1
| 10 | Sims uterine scissors, curved, blunt/blunt, 23cm |  | 1
| 11 | Mayo-Harrington dissecting scissors, straight, 23cm |  | 1
| 12 | Toennis-Adson dissecting scisasors with tungsten carbide edges, curved 17,5cm, gold handles |  | 2
| 13 | De Bakey vascular forceps with non-traumatic serration, straight, 2,7mm, 20cm |  | 1
| 14 | Dissecting forceps, narrow pattern, 14,5cm |  | 2
| 15 | Dissecting forceps narrow pattern 18cm straight |  | 2
| 16 | Dissecting forceps standard, straight, 30,5cm |  | 2
| 17 | Tissue forceps standard model, 1x2teeth, straight, 14,5cm |  | 2
| 18 | Tissue forceps , standard model 1x2 teeth 18cm |  | 2
| 19 | Dissecting forceps with tungsten carbide inserts, 0,5mm, 14,5cm, standard, gold handles |  | 1
| 20 | Potts-Smith dissecting forceps with tungsten carbide inserts, with dissector end, 0,4mm, 24,5cm, gold handle |  | 1
| 21 | Needle sterilizing exterior and interior box, 65x15mm |  | 1
| 22 | Scalpel handle no. 4 length 130 mm |  | 1
| 23 | Scalpel handle no. 3 length 120 mm |  | 1
| 24 | Scalpel blade remover to remove the blades from handle without the risk of injury, 14cm |  | 1
| 25 | Babcock intestinal and tissue grasping forceps 17cm |  | 2
| 26 | Micro-Crile delicate haemostatic forceps, curved, 14cm |  | 6
| 27 | Rochester-Pean haemostatic forceps, curved, 18,5cm |  | 6
| 28 | Gemini-Mixter dissecting and ligature forceps, 20,5cm |  | 1
| 29 | Rochester-Pean haemostatic forceps, curved, 24cm |  | 2
| 30 | Rochester-Pean haemostatic forceps, straight, 24cm |  | 2
| 31 | Overholt dissecting and ligature forceps, delicate pattern, curved, 21,5cm |  | 2
| 32 | Overholt dissecting and ligature forceps, delicate pattern, curved, 22cm |  | 2
| 33 | Gross sponge and dressing forceps, straight, 20cm |  | 2
| 34 | Gross sponge and dressing forceps, curved, 20cm |  | 2
| 35 | Foerster-Ballenger dressing and sponge holding forceps serrated jaws 24,5cm straight |  | 2
| 36 | Foerster-Ballenger dressing and sponge holding forceps serrated jaws 24,5cm curved |  | 2
| 37 | Lahey uterine tenaculum forceps, 15,5cm |  | 1
| 38 | Blohmke tonsile grasping forceps, 20cm |  | 1
| 39 | Schroder uterine tenaculum forceps, 1x1 teeth, 25cm |  | 2
| 40 | Backhaus towel clamp, 13,5cm |  | 14
| 41 | Kocher-Ochsner haemostatic forceps straight 16cm 1x2 teeth |  | 2
| 1.1.3 | Abdominal retractor holding set |  | 1 |  | 0
| 1 | Main holder with eccentric closure and integrated isolation |  | 1
| 2 | Crossbar with ball |  | 1
| 3 | Retainer |  | 1
| 4 | Universal Clamp |  | 6
| 5 | Holder arm maxi with ball |  | 2
| 6 | Blade Holder with handle nut |  | 3
| 7 | Blade holder with ball-and-socket joint, eccentric closure and handle nut, Allows a flexible and focused site setting. |  | 1
| 8 | Kirschner - Abdominal retractor, swivel 80mm/65mm, |  | 2
| 9 | Kirschner - Abdominal retractor, swivel 60mm/85mm, |  | 2
| 10 | Retractor blade, adjustable, 90mm/140mm, |  | 1
| 11 | Breisky specula, blade, 25 mm/150mm, |  | 1
| 12 | Retractor blade, malleable, 50mm/160mm, |  | 1
| 13 | St. Marks - Pelvis retractor, blade, 60mm/175mm |  | 1
| 1.1.4 | Chest set |  | 2 |  | 0
| 1 | Roux retractor, double-ended, 30x30mm/35x40mm, 17cm |  | 2
| 2 | Farabeuf retractor set, 25x15mm/30x15mm, 20x15mm/25x56mm, 15,5cm |  | 2
| 3 | Russ modell grasping vascular and tumor forceps, 15cm |  | 2
| 4 | Tuttle (Singley) lung dissecting forceps, 23cm |  | 2
| 5 | Tissue forceps, standard model, 1x2 teeth, 14,5cm |  | 3
| 6 | Dissecting forceps, standard model, 14,5cm |  | 2
| 7 | Dissecting forceps, standard model, 18cm |  | 2
| 8 | Dissecting forceps with tungsten carbide inserts, 0,4mm serration, 14cm, golden handles |  | 1
| 9 | Potts-Smith dissecting forceps with tungsten carbide inserts, 0,4mm serration, 18,5cm, with dissector end, golden handles |  | 1
| 10 | De Bakey vascular forceps with non-traumatic serration, 2mm, 15,5cm |  | 2
| 11 | De Bakey vascular forceps with non-traumatic serration, 2mm, 20cm |  | 2
| 12 | Surgical scissors, standard model, sharp/blunt, 13cm, straight |  | 1
| 13 | Metzenbaum dissecting scissors, standard model, curved, 18cm |  | 2
| 14 | Metzenbaum dissecting scissors, standard model, curved, 20,5cm |  | 1
| 15 | Scalpel blade remover to remove the blades from handles without risk of injury, 14cm |  | 1
| 16 | Scalpel handle 13cm, no. 4 |  | 1
| 17 | Scalpel handle straight, no. 3L, 21cm |  | 1
| 18 | Scalpel handle angled, no. 3L, 20,5cm |  | 1
| 19 | Carroll needle holder 16cm |  | 2
| 20 | Crile-Wood needle holder with tungsten carbide inserts, serration 0,4mm, 18cm, golden handles |  | 2
| 21 | De Bakey needle holder with tungsten carbide inserts, serration 0,4mm, 18cm, golden handles |  | 2
| 22 | De Bakey needle holder with tungsten carbide inserts, serration 0,4mm, 20,5cm, golden handles |  | 2
| 23 | Mixter dissecting and ligature forceps, 23cm, curved |  | 1
| 24 | Leriche delicate haemostatic forceps, curved, 15cm |  | 6
| 25 | Rochester-Pean haemostatic forceps, curved, 18,5cm |  | 6
| 26 | Overholt-Geissendoerfer dissecting and ligature forceps, curved, 19,5cm |  | 2
| 27 | Overholt-Geissendoerfer dissecting and ligature forceps, curved, 19cm |  | 2
| 28 | Overholt-Geissendoerfer dissecting and ligature forceps, curved, 21cm |  | 2
| 29 | Overholt-Geissendoerfer dissecting and ligature forceps, curved, 22cm |  | 2
| 30 | Mikulicz peritoneum forceps, 1x2 teeth, 18,5cm |  | 2
| 31 | Babcock intestinal and tissue grasping forceps, 20,5cm |  | 2
| 32 | Maier sponge and dressing forceps, straight, with ratchet, 26cm |  | 2
| 33 | Maier sponge and dressing forceps, curved, with ratchet, 26cm |  | 2
| 34 | Backhaus towel clamp, 11cm |  | 14
| 35 | Foerster-Ballenger dressing and sponge holding forceps, serrated jaws, 24,5cm, straight |  | 4
| 36 | Bowl 130x55mm, 0,45l |  | 2
| 37 | Bowl 165x75mm, 1,1l |  | 2
| 38 | Stille-Giertz rib sheares, 27cm |  | 1
| 39 | Doyen rib raspatory, for children, right, 18cm |  | 1
| 40 | Doyen rib raspatory, for children, left, 18cm |  | 1
| 41 | Semb rib raspatory 14mm, 24cm |  | 1
| 42 | Semb rib raspatory 12mm, 24cm |  | 1
| 43 | Stitch helper (retractor for wound closing), 200mm |  | 1
| 44 | Burford-Finochietto rib spreader with 2 pairs of blades, for adults, 48x64mm/77x64mm, 210mm |  | 1
| 45 | Bailey rib contractor, for children, 20cm |  | 1
| 46 | Duval intestinal and tissue grasping forceps, wider blades, 23cm |  | 2
| 47 | Price-Thomas bronchus forceps, 22cm |  | 2
| 48 | Lees bronchus clamp with toothing De Bakey 22cm |  | 2
| 49 | Gigli hook handle for wire saw, 1 pair, 7cm |  | 1
| 50 | Olivecrona (Gigli) wire saw, 40cm |  | 1
| 51 | Luer bone rongeur, straight, 18cm |  | 1
| 52 | Liston bone cutting forceps, 27cm |  | 1
| 53 | Gallaher bone file 17cm |  | 1
| 54 | Robb intima spatula (3mm) and dilator(1,5mm), 20,5cm |  | 1
| 1.1.5 | Instrument set for Laparatomy 2 |  | 5 |  | 0
| 1 | ROUX Wound Retractor, 165 mm (6 1/2"), double ended, 30 x 30 mm, 36 x 44 mm, non-sterile, reusable
| 2 | FARABEUF Wound Retractor, 150 mm (6"), double ended, large, 20 x 16 mm, 23 x 16 mm, 24 x 16 mm, 28 x 16 mm, non-sterile, reusable, Set of 2
| 3 | KOCHER Wound Retractor (With Prong), 220 mm (8 3/4"), 4 prongs, semi-sharp, depth: 16 mm, width: 20 mm, non-sterile, reusable
| 4 | HARRINGTON-PEMBERTON Lung Retractor, 295 mm (11 5/8"), flexible, depth: 175 mm, width: 64 mm, non-sterile, reusable
| 5 | SIMON Vaginal Specula, 275 mm (10 3/4"), depth: 115 mm, width: 27 mm, non-sterile, reusable
| 6 | KOERTE Abdominal/Intestinal Spatula, 300 mm (12"), width: 46 mm, non-sterile, reusable
| 7 | FRITSCH Abdominal Wall Retractor, 235 mm (9 1/4"), Fig. 4, depth: 46 mm, width: 75 mm, non-sterile, reusable
| 8 | REVERDIN Abdominal/Intestinal Spatula, 285 mm (11 1/4"), width: 45 mm, 63 mm, non-sterile, reusable
| 9 | PATTERN RUSSIA Tissue Forceps, straight, 150 mm (6"), toothed, ring tip, non-sterile, reusable
| 10 | TUTTLE Lung Grasping Forceps (Tweezer, straight, 230 mm (9"), non-sterile, reusable
| 11 | Tissue Forceps, standard, straight, 145 mm (5 3/4"), standard, toothed (1x2), non-sterile, reusable
| 12 | Dressing Forceps (Tweezers), straight, 145 mm (5 3/4"), medium, serrated, non-sterile, reusable
| 13 | Dressing Forceps (Tweezers), straight, 180 mm (7"), medium, serrated, non-sterile, reusable
| 14 | Dressing Forceps (Tweezers), straight, 250 mm (10"), medium, serrated, non-sterile, reusable
| 15 | DUROGRIP TC Atraumatic Forceps, straight, 145 mm (5 3/4"), delicate, non-sterile, reusable
| 16 | DUROGRIP TC Atraumatic Forceps, straight, 180 mm (7"), delicate, non-sterile, reusable
| 17 | DUROGRIP TC Atraumatic Forceps, straight, 250 mm (9 3/4"), delicate, non-sterile, reusable
| 18 | DE BAKEY ATRAUMATA Atraumatic Forceps, straight, 150 mm (6"), toothing DE BAKEY, width: 2 mm, non-sterile, reusable
| 19 | DE BAKEY ATRAUMATA Atraumatic Forceps, straight, 200 mm (7 7/8"), toothing DE BAKEY, width: 2 mm, non-sterile, reusable
| 20 | DE BAKEY ATRAUMATA Atraumatic Forceps, straight, 240 mm (9 1/2"), toothing DE BAKEY, width: 2 mm, non-sterile, reusable
| 21 | DE BAKEY ATRAUMATA Atraumatic Forceps, straight, 300 mm (11 3/4"), toothing DE BAKEY, width: 2 mm, non-sterile, reusable
| 22 | Surgical Scissors, straight, 130 mm (5 1/8"), standard, sharp/blunt, non-sterile, reusable
| 23 | METZENBAUM Dissecting Scissors, curved, 180 mm (7"), blunt/blunt, non-sterile, reusable
| 24 | METZENBAUM Dissecting Scissors, curved, 200 mm (7 7/8"), blunt/blunt, non-sterile, reusable
| 25 | NELSON-METZENBAUM Dissecting Scissors, curved, 230 mm (9"), blunt/blunt, non-sterile, reusable
| 26 | NELSON-METZENBAUM Dissecting Scissors, curved, 285 mm (11 1/4"), blunt/blunt, non-sterile, reusable
| 27 | Scalpel Handle, 135 mm (5 1/4"), No. 4, non-sterile, reusable
| 28 | CRILE-WOOD Needleholder, straight, 160 mm (6 1/4"), heavy pattern, jaw with longitudinal groove, non-sterile, reusable
| 29 | DE BAKEY DUROGRIP TC Needleholder, straight, 150 mm (6"), with 0.4 mm pitch of serration, suture 4/0 - 6/0, non-sterile, reusable
| 30 | DE BAKEY DUROGRIP TC Needleholder, straight, 180 mm (7"), with 0.4 mm pitch of serration, suture 4/0 - 6/0, non-sterile, reusable
| 31 | DUROGRIP TC Needleholder, straight, 200 mm (7 7/8"), very delicate, with 0.2 mm pitch of serration, suture 6/0 - 10/0, non-sterile, reusable
| 32 | DE BAKEY DUROGRIP TC Needleholder, straight, 260 mm (10 1/4"), with 0.4 mm pitch of serration, suture 4/0 - 6/0, non-sterile, reusable
| 33 | WANGENSTEEN Needleholder, straight, 275 mm (10 3/4"), delicate, non-sterile, reusable
| 34 | MIXTER Dissecting a. Ligature Forceps, angled right, 230 mm (9"), non-sterile, reusable
| 35 | MICRO-HALSTED Hemostatic Forceps, curved, 125 mm (5"), delicate, non-sterile, reusable
| 36 | LERICHE Hemostatic Forceps, curved, 150 mm (6"), delicate, blunt, non-sterile, reusable
| 37 | ROCHESTER-PEAN Hemostatic Forceps, curved, 185 mm (7 1/4"), blunt, non-sterile, reusable
| 38 | ROCHESTER-PEAN Hemostatic Forceps, curved, 240 mm (9 1/2"), blunt, non-sterile, reusable
| 39 | ROCHESTER-PEAN Hemostatic Forceps, straight, 240 mm (9 1/2"), blunt, non-sterile, reusable
| 40 | OVERHOLT-GEISSENDOERFER Dissecting a. Ligature Forceps, curved, 195 mm (7 3/4"), Fig. 0, non-sterile, reusable
| 41 | OVERHOLT-GEISSENDOERFER Dissecting a. Ligature Forceps, curved, 205 mm (8"), Fig. 1, non-sterile, reusable
| 42 | OVERHOLT-GEISSENDOERFER Dissecting a. Ligature Forceps, curved, 220 mm (8 3/4"), Fig. 2, non-sterile, reusable
| 43 | OVERHOLT-GEISSENDOERFER Dissecting a. Ligature Forceps, curved, 225 mm (8 7/8"), Fig. 3, non-sterile, reusable
| 44 | OVERHOLT-GEISSENDOERFER Dissecting a. Ligature Forceps, curved, 260 mm (10 1/4"), Fig. 1, non-sterile, reusable
| 45 | MIKULICZ Hemostatic Forceps, curved, 185 mm (7 1/4"), toothed (1x2), non-sterile, reusable
| 46 | BABCOCK Intestinal Grasping Forceps, straight, 215 mm (8 1/2"), width: 10 mm, non-sterile, reusable
| 47 | Intestinal Grasping Forceps, straight, 175 mm (6 7/8"), width: 8 mm, non-sterile, reusable
| 48 | BABCOCK ATRAUMATA Organ And Tisue Grasping Force, straight, 240 mm (9 1/2"), toothing DE BAKEY, non-sterile, reusable
| 49 | MAIER Dressing Forceps, straight, 260 mm (10 1/4"), serrated, box lock, with ratchet, non-sterile, reusable
| 50 | MAIER Dressing Forceps, curved, 260 mm (10 1/4"), serrated, box lock, with ratchet, non-sterile, reusable
| 51 | BACKHAUS Towel Clamp, curved, 110 mm (4 1/4"), sharp, non-sterile, reusable
| 52 | FOERSTER-BALLENGER Sponge Forceps, straight, 245 mm (9 5/8"), oval, serrated, fenestrated, box lock, with ratchet, non-sterile, reusable
| 53 | HEYWOOD-SMITH Haemorrhoidal Forceps, straight, 200 mm (7 7/8"), oval, serrated, fenestrated, screw lock, with ratchet, non-sterile, reusable
| 54 | HEYWOOD-SMITH Organ Grasping Forceps, straight, 250 mm (9 3/4"), oval, screw lock, non-sterile, reusable
| 55 | Round Bowl, 111 mm, height: 56 mm, width: 72 mm, 300 ml
| 56 | Round Bowl, 151 mm, height: 73 mm, width: 100 mm, 1000 ml
| 57 | Round Bowl, 211 mm, height: 96 mm, width: 145 mm, 2500 ml
| 1.1.6 | Laparoscopic instruments set |  | 3 |  | 0
| 1 | Tissue forceps standard model, 1x2teeth, 14,5cm |  | 2
| 2 | Russ. Model vascular and tumor grasping forceps 15cm |  | 2
| 3 | Surgical scissors standard model, sharp/blunt, straight, 13cm |  | 1
| 4 | Metzenbaum dissecting scissors standard model, curved, 18cm |  | 2
| 5 | Scalpel handle No. 3, 12cm |  | 1
| 6 | Carroll needle holder 16cm |  | 1
| 7 | Leriche delicate haemostatic forceps, curved, 15cm |  | 2
| 8 | Rochester-Pean haemostatic forceps, curved, 18,5cm |  | 6
| 9 | Rochester-Pean haemostatic forceps, curved, 24cm |  | 2
| 10 | Micro forceps curved, tip 0,8mm, 21cm, titanium with ceramic coating |  | 1
| 11 | Mikulicz peritoneum forceps, 1x2 teeth, curved, 18,5cm |  | 2
| 12 | Backhaus towel clamp 11cm |  | 14
| 13 | Foerster-Ballenger dressing and sponge holding forceps, serrated jaws, straight, 24,5cm |  | 4
| 14 | Foerster-Ballenger dressing and sponge holding forceps, serrated jaws, straight, 18cm |  | 1
| 15 | Trachea retractor, blunt, 3 teeth, 17cm |  | 1
| 16 | Director probe 16cm |  | 1
| 17 | Bowl 116x50mm, 0,35l |  | 3
| 18 | Mixter dissecting forceps, angled 90°, diam.10mm, jaw length 20,5mm, shaft length 330mm, handle without ratchet with connector, complet instrument |  | 1
| 19 | Instrument insert Claw forceps, 2x3 teeth,diam.10mm, jaw length 30mm, shaft length 330mm |  | 1
| 20 | Tube shaft length 330mm, for instrument insert diam. 10mm |  | 1
| 21 | Handle with connector and ratchet |  | 1
| 22 | Instrument insert grasping forceps, corrugated, atraumatic, diam.5mm, jaw length 21,6mm, shaft length 330mm |  | 1
| 23 | Tube shaft length 330mm, for instrument insert diam. 5mm |  | 1
| 24 | Handle with connector and ratchet |  | 1
| 25 | Instrument insert grasping forceps, corrugated, atraumatic, diam.5mm, jaw length 26,6mm, shaft length 330mm |  | 1
| 26 | Tube shaft length 330mm, for instrument insert diam. 5mm |  | 1
| 27 | Handle with connector and ratchet |  | 1
| 28 | Instrument insert Dorsey (Johann) grasping forceps fenestrated, diam.5mm, jaw length 20mm, shaft length 330mm |  | 1
| 29 | Tube shaft length 330mm, for instrument insert diam. 5mm |  | 1
| 30 | Handle with connector and ratchet |  | 1
| 31 | Instrument insert Clinch grasping forceps, fenestrated, atraumatic, diam.5mm, jaw length 20mm, shaft length 330mm |  | 1
| 32 | Tube shaft length 330mm, for instrument insert diam. 5mm |  | 1
| 33 | Handle with connector and ratchet |  | 1
| 34 | Trocar with three-edged tip, diam. 5,5mm |  | 1
| 35 | Trocar with conical tip, diam. 5,5mm |  | 1
| 36 | Trocar sleeve with flapping valve, stopcock and interchangeable metal tube, smooth, diam. 5,5mm x 95mm |  | 2
| 37 | Trocar sleeve with flapping valve, stopcock and interchangeable metal tube, spiral-shaped, diam. 5,5mm x 95mm |  | 1
| 38 | Veress insufflation cannulae with stopcock and Luer-Lock connection, chrome-plated, diam.2mm x 100mm |  | 1
| 39 | Handle with 2-way stopcock, sliding valve and tube connectors |  | 1
| 40 | Roundhole suction tube with 16 holes,diam.5mm x 330mm |  | 1
| 1.1.7 | Amputation instruments set |  | 2 |  | 0
| 1 | PATTERN RUSSIA Tissue Forceps, straight, 150 mm (6"), toothed, ring tip, non-sterile, reusable |  | 2
| 2 | Tissue Forceps, standard, straight, 145 mm (5 3/4"), standard, toothed (1x2), non-sterile, reusable |  | 2
| 3 | Dressing Forceps (Tweezers), straight, 145 mm (5 3/4"), medium, serrated, non-sterile, reusable |  | 2
| 4 | DUROGRIP TC Atraumatic Forceps, straight, 145 mm (5 3/4"), delicate, non-sterile, reusable |  | 1
| 5 | Surgical Scissors, straight, 130 mm (5 1/8"), standard, sharp/blunt, non-sterile, reusable |  | 1
| 6 | METZENBAUM Dissecting Scissors, curved, 160 mm (6 1/4"), blunt/blunt, non-sterile, reusable |  | 1
| 7 | METZENBAUM Dissecting Scissors, curved, 180 mm (7"), blunt/blunt, non-sterile, reusable |  | 1
| 8 | Scalpel Handle, 135 mm (5 1/4"), No. 4, non-sterile, reusable |  | 1
| 9 | CRILE-WOOD Needleholder, straight, 160 mm (6 1/4"), heavy pattern, jaw with longitudinal groove, non-sterile, reusable |  | 2
| 10 | LERICHE Hemostatic Forceps, curved, 150 mm (6"), delicate, blunt, non-sterile, reusable |  | 6
| 11 | ROCHESTER-PEAN Hemostatic Forceps, curved, 185 mm (7 1/4"), blunt, non-sterile, reusable |  | 4
| 12 | OVERHOLT-GEISSENDOERFER Dissecting a. Ligature Forceps, curved, 195 mm (7 3/4"), Fig. 0, non-sterile, reusable |  | 2
| 13 | OVERHOLT-GEISSENDOERFER Dissecting a. Ligature Forceps, curved, 205 mm (8"), Fig. 1, non-sterile, reusable |  | 2
| 14 | BACKHAUS Towel Clamp, curved, 110 mm (4 1/4"), sharp, non-sterile, reusable |  | 12
| 15 | FOERSTER-BALLENGER Sponge Forceps, straight, 245 mm (9 5/8"), oval, serrated, fenestrated, box lock, with ratchet, non-sterile, reusable |  | 4
| 16 | GROSS Dressing Forceps, straight, 180 mm (7"), serrated, screw lock, with ratchet, non-sterile, reusable |  | 2
| 17 | Round Bowl, 111 mm, height: 56 mm, width: 72 mm, 300 ml |  | 2
| 18 | Round Bowl, 151 mm, height: 73 mm, width: 100 mm, 1000 ml |  | 1
| 19 | PERCY Amputation Retractor, removable handles, non-sterile, reusable |  | 1
| 20 | SATTERLEE Spare blade, 200 mm, 8", to be used with FH325R |  | 1
| 21 | GIGLI Hook handle for saw |  | 1
| 22 | GIGLI Wire saw, fine, 1,4 mm diam., 400 mm, 16", 6 wires twisted, package of 5 pieces |  | 1
| 23 | LISTON Amputation Knife, jaw length: 220 mm(8 3/4") |  | 1
| 24 | Bone File, 295 mm (11 3/4"), serrated (cross), width: 20 mm |  | 1
| 25 | SCHNEIDER-SAUERBRUCH Raspatory, curved, 190 mm (7 1/2"), sharp, width: 14 mm, non-sterile, reusable |  | 1
| 26 | Bone Curette, 200 mm (7 7/8"), double ended, non-sterile, reusable |  | 1
| 27 | STILLE Osteotome, straight, 12 mm, 205 mm, 8" |  | 1
| 28 | STILLE Gouge, straight, 12 mm, 205 mm, 8" |  | 1
| 29 | HEATH Mallet, 240 mm (9 1/2"), diam. 40 mm, peek handle, stainless steel, 850 g, 500 g |  | 1
| 30 | LUER Bone Rongeur, straight, 180 mm (7"), oval |  | 1
| 31 | LISTON Bone Cutting Forceps, straight, 170 mm (6 3/4"), non-sterile, reusable |  | 1
| 32 | KOCHER-OCHSNER Hemostatic Forceps, curved, 225 mm (8 7/8"), toothed (1x2), non-sterile, reusable |  | 1
| 33 | SEMB Bone Holding Forceps, curved to side, 190 mm (7 1/2"), width: 8 mm, non-sterile, reusable |  | 1
| 34 | Vascular Spatula, curved, 215 mm (8 1/2"), blunt, width: 3 mm, non-sterile, reusable |  | 1
| 1.1.8 | Set for breast surgery |  | 2 |  | 0
| 1 | Baby-Roux retractor 18x17mm/24x21mm, 12,5cm |  | 2
| 2 | Farabeuf retractor, double-ended set, 25x10mm/32x12mm, 22x10mm/29x12mm, 12,5cm |  | 2
| 3 | Weitlaner self-retaining retractor sharp 2x3 prongs 10,5cm |  | 1
| 4 | Weitlaner self-retaining retractor sharp 3x4 prongs 13,5cm |  | 1
| 5 | Simon retractor for deep surgery 115x27mm, 28cm |  | 1
| 6 | Langenbeck retractor 21cm ,42x13mm |  | 2
| 7 | RUSS. Model vascular and tumor forceps 15cm |  | 2
| 8 | Tissue forceps standard, 1x2 teeth, 14,5cm |  | 3
| 9 | Dissecting forceps narrow model 14,5cm |  | 2
| 10 | Dissecting forceps with tungsten carbide inserts, 0,5mm, 14,5cm, standard model, gold handle |  | 1
| 11 | Vascular forceps De-Bakey with non-traumatic serration, straight, 2 mm, length 155 mm |  | 1
| 12 | Surgical scissors standard model, sharp/blunt, straight 13cm |  | 1
| 13 | Metzenbaum dissecting scissors curved 15,5cm |  | 1
| 14 | Metzenbaum dissecting scissors with tungsten carbide edges, delicate model, curved, 18cm, gold handles |  | 1
| 15 | Dissecting scissors with welded tungsten carbide blades, curved, 18cm, with stable 3-point positions, ergonomically designed scissors rings with V-shaped handle |  | 1
| 16 | Scalpel handle no. 4 length 130 mm |  | 1
| 17 | Scalpel handle no. 3 length 120 mm |  | 1
| 18 | Scalpel blade remover to remove the blades from handle without the risk of injury, 14cm |  | 1
| 19 | Carroll needle holder 16cm |  | 1
| 20 | De Bakey needle holder with tungsten carbide inserts, 0,5mm, 15cm, gold handles |  | 1
| 21 | Mixter-O'Shaugnessy dissecting and ligature forceps 17,5cm |  | 1
| 22 | Leriche delicate haemostatic forceps curved 15cm |  | 4
| 23 | Overholt-Geissendoerfer dissecting and ligature forceps 19,5cm curved |  | 2
| 24 | Overholt-Geissendoerfer dissecting and ligature forceps 19cm, curved |  | 2
| 25 | Babcock intestinal and tissue grasping forceps 16cm |  | 2
| 26 | Backhaus towel clamp, 11cm |  | 12
| 27 | Foerster-Ballenger dressing and sponge holding forceps serrated jaws 18cm straight |  | 4
| 28 | Gross sponge and dressing forceps, straight, 20cm |  | 2
| 29 | Round bowl 115x50mm, 0,35l |  | 1
| 30 | Round bowl 145x65mm, 0,70l |  | 1
| 31 | Perforated stainless steel tray 550x250x70mm |  | 1
| 1.1.9 | Set for goiter surgery |  | 2 |  | 0
| 1 | Baby-Roux retractor 18x17mm/24x21mm, 12,5cm |  | 2
| 2 | Farabeuf retractor, double-ended set, 25x10mm/32x12mm, 22x10mm/29x12mm, 12,5cm |  | 2
| 3 | Weitlaner self-retaining retractor sharp 2x3 prongs 10,5cm |  | 1
| 4 | Weitlaner self-retaining retractor sharp 3x4 prongs 13,5cm |  | 1
| 5 | Weitlaner self-retaining retractor sharp 3x4 prongs 13,5cm |  | 1
| 6 | Tissue forceps standard, 1x2 teeth, 14,5cm |  | 2
| 7 | Dissecting forceps standard model 13cm |  | 2
| 8 | Dissecting forceps with tungsten carbide inserts, 0,5mm, 14,5cm, standard model, gold handle |  | 1
| 9 | Vascular forceps De-Bakey with non-traumatic serration, straight, 2 mm, length 155 mm |  | 2
| 10 | Surgical scissors standard model sharp/blunt straight 13cm |  | 1
| 11 | Metzenbaum dissecting scissors curved 15,5cm |  | 1
| 12 | Metzenbaum dissecting scissors with tungsten carbide edges, delicate model, curved, 18cm, gold handles |  | 1
| 13 | Dissecting scissors with welded tungsten carbide blades, curved, 18cm, with stable 3-point positions, ergonomically designed scissors rings with V-shaped handle |  | 1
| 14 | Scalpel handle no. 4 length 130 mm |  | 1
| 15 | Scalpel handle no. 3 length 120 mm |  | 1
| 16 | Scalpel blade remover to remove the blades from handle without the risk of injury, 14cm |  | 1
| 17 | Carroll needle holder 16cm |  | 1
| 18 | De Bakey needle holder with tungsten carbide inserts, 0,4mm, 15cm, gold handles |  | 1
| 19 | Converse needle holder with tungsten carbide inserts 0,4mm, 10,5cm with big sized rings |  | 1
| 20 | Mixter-O'Shaugnessy dissecting and ligature forceps 17,5cm |  | 1
| 21 | Baby-Mosquito (Hartmann) delicate haemostatic forceps curved 10cm |  | 4
| 22 | Micro-Halsted delicate haemostatic forceps curved 12cm |  | 2
| 23 | Overholt dissecting and ligature forceps curved 18,5cm |  | 4
| 24 | Babcock intestinal and tissue grasping forceps 16cm |  | 2
| 25 | Backhaus towel clamp, 11cm |  | 12
| 26 | Foerster-Ballenger dressing and sponge holding forceps serrated jaws 18cm straight |  | 4
| 27 | Gross sponge and dressing forceps, straight, 20cm |  | 1
| 28 | Round bowl , 0,35l |  | 3
| 29 | Perforated stainless steel tray 250x250x70mm |  | 1
| 1.1.10 | Urology set |  | 1 |  | 0
| 1 | Roux retractor set, double-ended, 17cm, 21x24/29x28mm, 27x26/32x37mm, 31x30/36x42mm |  | 1
| 2 | Collin retractor double-ended, 29x14/35x17mm, 25x14/31x17mm, 14,5cm |  | 1
| 3 | Koerte retractor, handle vertical to the blade, sharp, 26x39mm, 24cm |  | 2
| 4 | Balfour abdominal retractor, complet, consisting of: 1 frame, 1pair lateral valves 64x45mm, 1 pair lateral valve 89x45mm and 1 central valve 65x84mm |  | 1
| 5 | Denis-Browne abdominal retractor complet, consisting of: 1 frame 180x150mm, 2 valves 40x30mm and 2 valves 40x40mm |  | 2
| 6 | Legueu bladder spatula, 105x35mm, 27cm |  | 2
| 7 | Schoenborn retractor, 6x13mm, 21cm |  | 2
| 8 | Cushing vein retractor, 11x14mm, 24cm |  | 2
| 9 | Allis intestinal and tissue grasping forceps, 5x6 teeth, 22,5cm |  | 4
| 10 | Dieffenbach buldog clamp 48mm |  | 2
| 11 | Johns-Hopkins buldog clamp 65mm |  | 2
| 12 | Bowl 128x55mm, 0,45l |  | 3
| 13 | Bowl 167x75mm, 1,1l |  | 2
| 14 | Rampley dressing and sponge holding forceps, straight, 24,5cm |  | 6
| 15 | Backhaus towel clamp, 9cm |  | 10
| 16 | Backhaus towel clamp, 13,5cm |  | 10
| 17 | Scalpel blade remover to remove the blades from handles without risk of injury, 14cm |  | 1
| 18 | Scalpel handle 13cm, no. 4 |  | 2
| 19 | Scalpel handle 12cm, no. 3 |  | 2
| 20 | Scalpel handle straight, no. 3L, 21cm |  | 1
| 21 | Adson delicate tissue forceps, straight, 1x2 teeth, 12cm |  | 2
| 22 | Tissue forceps, 2x3 teeth, 20cm |  | 3
| 23 | Waugh delicate tissue forceps, 1x2 teeth, 25cm |  | 2
| 24 | Dissecting forceps with tungsten carbide inserts, 0,5mm serration, standard model, 25cm, golden handles |  | 2
| 25 | De Bakey vascular forceps with non-traumatic serration, straight, 2,7mm, 30,5cm |  | 2
| 26 | Potts-Smith dissecting forceps with tungsten carbide inserts, 0,4mm serration, 21cm, with dissector end, golden handles |  | 2
| 27 | De Bakey vascular forceps with non-traumatic serration, straight, 3,5mm, 24cm |  | 2
| 28 | Liston-Key-Horsley bone cutting forceps, S-shaped, for posterior rib stumps, 24,5cm |  | 1
| 29 | Lambotte rib raspatory 10mm, 21cm |  | 1
| 30 | Doyen rib raspatory, for adults, right, 17cm |  | 2
| 31 | Doyen rib raspatory, for adults, left, 17cm |  | 2
| 32 | Bone file, 13mm, 24,5cm |  | 1
| 33 | Cooley-Reynolds anastomosis clamp with toothing Cooley, 16,5cm |  | 1
| 34 | Cooley-Reynolds anastomosis clamp with toothing Cooley, 17,5cm |  | 1
| 35 | Cooley-Derra anastomosis clamp with toothing Cooley, small, 16,5cm |  | 1
| 36 | Cooley-Derra anastomosis clamp with toothing Cooley, medium, 16,5cm |  | 1
| 37 | De Bakey-Satinsky tangential occlusion clamp, 24,5cm |  | 2
| 38 | De Bakey-Satinsky tangential occlusion clamp, 24cm |  | 1
| 39 | Randall kidney stone forceps, 23cm |  | 1
| 40 | Randall kidney stone forceps, 20,5cm |  | 1
| 41 | Randall kidney stone forceps, 19cm |  | 1
| 42 | Randall kidney stone forceps, 18cm |  | 1
| 43 | Mixter dissecting and ligature forceps, 23cm, curved |  | 1
| 44 | Gemini-Mixter dissecting and ligature forceps, 23cm |  | 1
| 45 | Rochester-Pean haemostatic forceps, curved, 22cm |  | 3
| 46 | Rampley dressing and sponge holding forceps, straight, 18cm |  | 3
| 47 | Rochester-Pean haemostatic forceps, curved, 18,5cm |  | 4
| 48 | Baby-Mosquito (Hartmann) delicate haemostatic forceps, curved, 10cm |  | 10
| 49 | Wertheim-De Bakey (Rogers) parametrium forceps, curved, 24,5cm |  | 4
| 50 | Lawrence dissecting and ligature forceps, 28cm |  | 4
| 51 | Schnidt haemostatic and tonsil forceps, curved, 18,5cm |  | 6
| 52 | Dissecting and ligature forceps Overholt 20.5cm delicate pattern |  | 2
| 53 | Dissecting and ligature forceps Overholt 22cm delicate pattern |  | 4
| 54 | Intestinal and tissue grasping forceps Babcock 23.5cm |  | 2
| 55 | Kidney clamp Guyon-De Bakey with non-traumatic serration 24cm curved |  | 3
| 56 | Needleholder with tungsten carbide inserts De Bakey serration 0.2mm, 23cm, golden handles |  | 2
| 57 | Needleholder with tungsten carbide inserts Mayo-Hegar heavy pattern serration 0.5mm, 20.5cm, golden handles |  | 2
| 58 | Needleholder Webster smooth jaws, 13cm |  | 2
| 59 | Dissecting scissors Metzenbaum 23cm standard pattern curved |  | 2
| 60 | Dissecting scissors Nelson-Metzenbaum 28.5cm curved |  | 2
| 61 | Dissecting scissors Toennis-Adson 17.5cm curved very delicate pattern |  | 2
| 62 | Dissecting scissors Metzenbaum 14.5cm curved |  | 1
| 63 | Vascular scissors Potts-Smith angled on side 60° , 18cm |  | 1
| 64 | Vascular scissors Potts-Smith angled on side 60° , 18cm, one probe-pointed blade |  | 1
| 65 | Kocher-Ochsner haemostatic forceps 1x2 teeth, curved 24 cm |  | 2
| 66 | Doyen polypus and ovum forceps 18.5cm |  | 2
| 67 | Buttoned probe diameter 2mm length 18cm |  | 4
| 68 | Buttoned probe diameter 2mm length 20cm |  | 2
| 69 | Intestinal clamp forceps Doyen 21cm straight, soft elastic |  | 2
| 70 | Intestinal clamp forceps Doyen 21cm curved, soft elastic |  | 2
| 71 | Stitch helper (retractor for wound closing), 200mm |  | 1
| 72 | Deaver retractor 30cm, 100x25mm |  | 2
| 73 | Deaver retractor 30cm, 100x38mm |  | 2
| 74 | Deaver retractor 30cm, 100x50mm |  | 2
| 75 | Deaver retractor 30.5cm, 100x75mm |  | 2
| 76 | Fritsch retractor 23,5cm/ 35x40mm, handle vertical to the blade |  | 2
| 77 | Fritsch retractor 26cm/ 60x85mm, handle vertical to the blade |  | 2
| 78 | Suction tube De Bakey diamater 9mm, length 28cm, with lateral slots |  | 2
| 1.1.11 | Sterilization containers and baskets, SET |  | 1 |  | 0
| 1 | WIRE BASKET 250X250X70 MM |  | 2
| 2 | WIRE BASKET 550X250X100 MM |  | 1
| 3 | CONTAINER 1/1 BASIC MODEL 600X280X200 MM GREY - BOTTOM NON-PERFORATED/LID PERFORATED with filter for min 2000 sterilization cycles |  | 1
| 4 | RACK FOR ENDOSCOPY 500X250X125 MM USED WITH 1/1 200 MM HEIGHT CONTAINER |  | 1
| 5 | CONTAINER 1/2 DENTAL MODEL 230X175X65 MM GREY - LID PERFORATED - BOTTOM NON-PERFORATED TEFLON FILTER FOR 1/2 DENTAL CONTAINERS 130X130 MM (2000 CYCLE) |  | 40
| 6 | CONTAINER 1/2 BASIC MODEL 280X310X200 MM GREY - BOTTOM PERFORATED/LID PERFORATED with filter for min 2000 sterilization cycles |  | 40
| 7 | CONTAINER 1/2 BASIC MODEL 280X310X250 MM GREY - BOTTOM PERFORATED/LID PERFORATED with filter for min 2000 sterilization cycles |  | 10
| 8 | KIDNEY TRAY STAINLESS STEEL |  | 100
| 9 | WIRE BASKET 250X250X70 MM |  | 1
| 10 | WIRE BASKET 250X250X100 MM |  | 1
| 11 | WIRE BASKET 550X250X100 MM |  | 1
| 1.1.12 | Pouch making machine |  | 1 |  | 0
| 1 | Width of seal seam and cutting lenghts, roll holder - 400 mm
| 2 | Suitable for use in medical practice, hospital and medical industry
| 3 | Adjustment pouch lengths 100 - 1,000 mm
| 4 | Adjustment quantity of pouches 1 - 500 pouches
| 5 | Sealing temperature max. 200 °C
| 6 | Contact pressure fix
| 7 | Have overheating protection
| 8 | LCD-Display
| 9 | Microprocessor controlled
| 10 | Interface USB A/B, RS 232, Ethernet connection
| 11 | Barcode reader
| 12 | Stainless steel AISI 304 casing cover
| 13 | Reel holder
| 14 | Sorting unit
| 15 | Sealing process automatic/reproducible
| 16 | Alarm and machine stop in case the monitored parameters exceed predetermined limits
1.2 | 1.2.1 | Telescopes for ENT (set) |  | 1 |  | 0
| 1 | HD Telescope, 4 mm, 0°, connector for light guide on bottom, autoclavable |  | 1
| 2 | HD Telescope, 4 mm, 30°, connector for light guide on bottom, autoclavable |  | 1
| 3 | Telescope,2.7 mm, 0°, autoclavable |  | 1
| 4 | Telescope,2.7 mm, 30°, autoclavable |  | 1
| 5 | Light-guide cable,3 m, CF type |  | 1
| 6 | LED technology light source, noise-free, homogeneous light distribution, low energy consumption (55 W) |  | 1
1.3 | 1.3.1 | Ultrasound surgical system |  | 1 |  | 0
| 1 | Ultrasound surgical system with accessories for surgical cuts of soft tissues when bleeding control and minimal thermal injury
| 2 | It can be used as a supplement or replacement for electrosurgery, laser surgery and traditional scalpels in general, gynecological and thoracic surgery and exposure to orthopedic structures (such as the hip joint)
| 3 | Foot control switch option
| 4 | LCD on the rear panel for displaying the frequency during activation, error code and probe lifetime when booting the system
| 5 | Microprocessor-controlled start scissors
| 6 | Production of ultrasonic energy at the resonant frequency of connected scissors in the range of 35 to 37 kHz
| 7 | Control the output of the generator, which ensures the automatic increase in the energy of the converter, which leads to higher power delivery to scissors
| 9 | Classification of generator isolation - Class I
| 10 | Classification of transducer isolation - type BF
| 13 | It has only internal fuses
| 14 | Main frequency 50/60 Hz
| 15 | Possesses a bond for equipotential bonding
| 18 | Maximum output power 70w +/- 30W
| 19 | Output frequency - 36 kHz
| 20 | Precision of the output frequency indicator – 1%
1.4 | 1.4.1 | Endoscopic ultrasound system with Linear ultrasound gastroscope. |  | 1 |  | 0
| 1 | Ultrasound gastrovideoscope Linear
| 2 | Field of view 100°/ forward oblique view 55°
| 3 | Depth of field 3-100 mm
| 4 | Insertion tube outer diameter 12,6 mm
| 5 | Channel inner diameter 3,7 mm
| 6 | Angulation range Up/Down 130°/90° left/right 90°/90°
| 7 | Scanning Direction: Parallel to the insertion direction
| 8 | Frequency: 5/6/7.5/10/12 MHz
| 9 | Scaning range 180°
| 10 | Contacting Method: Balloon method, Direct Contact method
| 11 | Operation mode: B-mode, FLOW mode, PW mode or equivalent
| 12 | Connection with US processor via detachable cable
| 13 | Ultra Sound video processor
| 14 | Mobile ulrasound system with 4 active connectors
| 15 | Monitor 21“
| 16 | Scanning methods:
| 17 | Electronic Convex
| 18 | Electronic Linear
| 19 | Electronic Phased Array Sector
| 20 | Electronic Radial
| 21 | Operating Modes:
| 22 | B-mode
| 23 | BiPlane-mode
| 24 | M-mode
| 25 | D: Spectral Doppler mode (PW, CW, HPRF-PW)
| 26 | Dual Gate Doppler mode
| 27 | Color Flow mode
| 28 | Power Doppler mode (Directional Power Doppler)
| 29 | eFLOW mode (Directional eFLOW)
| 30 | Real-time Tissue Elastography
| 31 | Compatible with wide range of ultrasound probes (abdominal, transrectal...)
| 32 | Xenon light source
| 33 | 300 W xenon lamp, 500h expected lamp life
| 34 | Emergency lamp
| 35 | Air feeding pump – 4 levels (OFF, low, medium, high)
| 36 | Enhancement of the visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light
| 37 | Automatic brightness adjustment
| 38 | Video processor
| 39 | Enhancement of the visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light
| 40 | Analog HDTV signal output: (RGB, YPbPr),
| 41 | Analog SDTV signal output: VBS composite; Y/C i RGB
| 42 | Digital signal output: HD-SDI (SMTPE 292M), SD-SDI (SMPTE 259M), DV (IEEE 1394), i DVI (WUXGA, 1080p or SXGA)
| 43 | Picture-in-picture and index function
| 44 | Image enhancement setting:
| 45 | Structural enhancement: Enhancement of contrast of the fine patterns in the image
| 46 | Edge enhancement: Enhancement of edges of the endoscopic image
| 47 | Automatic gain control
| 48 | Contrast: 3 levels (Normal, Low, High)
| 49 | Iris (auto, peak and average)
| 50 | Freeze and pre-freeze control
| 51 | Endoscope’s remote switches function settings
| 52 | Registration of patient data
| 53 | White balance
| 54 | Color tone adjustment: Red, Blue, Chroma
| 55 | Different user settings
| 56 | Portable Memory to store endoscopic images
| 57 | Trolley
| 58 | Sliding Keyboard Tray
| 59 | LCD Monitor arm
| 60 | Scope Pole Kit
| 61 | Integrated Separation Transformer
| 62 | LCD Monitor:
| 63 | Display size min 30inches
| 64 | 4K upscaling function
| 65 | TFT active matrix LCD
| 66 | Resolution 3840x2160 pixels (UHD) or higher
| 67 | Contrast ratio 1000:1
| 68 | Aspect ratio 16:9
| 69 | Luminance 450 cd/m2 or more
| 70 | Viewing angle 178° /178°
| 72 | Multiple image display: PIP/POP
| 73 | Input:
| 74 | 4K in: 12G-SDI ×2, Display Port ×1, HDMI ×1
| 75 | 2K in: 3G-SDI ×1, DVI-D x 1
| 76 | Output:
| 77 | 4K out: 12G-SDI ×2
| 78 | 2K out: 3G-SDI ×1
| 79 | Clone Out: 12G-SDI ×1
| 80 | Suction pump:
| 81 | Nominal vacuum: 95 kPa
| 82 | pre-set flow speed settings 40l/min, 50l/min and 60l/min
| 83 | Additional vacuum regulation knob
| 84 | Vacuum gauge
| 85 | Thermal cut-out protection
| 86 | Suction jar 2l
| 87 | Tubing
| 88 | Disposable bacteria filter 10 pcs
1.5 | 1.5.1 | Uretroton |  | 1 |  | 0
| 1 | Sheath, 22 Fr ,for urethrotome, 4 Fr. Channel, with obturator |  | 1
| 2 | Outer sheath, 26 Fr., for percutaneous access |  | 1
| 3 | Insertion sleeve for balloon catheter of max. 18 Fr. |  | 2
| 4 | Working element |  | 2
| 5 | Knife,semi-circular, 4 Fr. Channel |  | 2
| 6 | Knife,semi-circular |  | 2
| 7 | Knife, lancet type, straight |  | 1
| 8 | Telescope 4 mm, 12°,HD, autoclavable |  | 1
1.6 | 1.6.1 | Bipolar resection |  | 1 |  | 0
| 1 | Telescope,4 mm, 12°, HD, autoclavable |  | 1
| 2 | Outer sheath,27 Fr.,2 stopcocks, rotatable |  | 1
| 3 | Resection sheath, 24 Fr. |  | 1
| 4 | Working element, active, for resection in saline |  | 1
| 5 | HF cable, bipolar, 4 m |  | 2
| 6 | HF-resection electrode, loop, 24 Fr., 0.2 wire, large, 12°, 12 pieces, sterile, single use |  | 1
| 7 | HF-Resection electrode, button, 24 – 28 Fr.,12° and 30°, 12 pieces, sterile, single use |  | 1
| 8 | Light-guide cable,3 m, CF type |  | 1
| 9 | Electrosurgical Unit |  | 1
| 10 | Performs open, laparoscopic and endoscopic surgery procedures
| 11 | Tissue-adaptive power output applies the optimal amount of energy for controlled, precise and effective cutting and coagulation performance, helps you cut quickly, cleanly and precisely
| 12 | Touchscreen and streamlined user interface allows the user to choose settings easily
| 13 | Memory function allows the user to save preferred settings and bring them up fast
| 14 | Instrument recognition at universal socket allows “plug and play” operation
| 15 | Automatic saline detection makes sure that you are using the correct fluid
| 16 | Monopolar-4 modes cutting and 4 modes coagulation, Bipolar-2 modes cutting and 6 modes coagulation.
| 17 | Sockets: Monolopar, Bipolar and Universal
| 18 | Disposable split patient plate,100 pcs |  | 1
| 19 | Cord for disposable plates |  | 1
| 20 | Evacuator, according to Ellik |  | 1
1.7 | 1.7.1 | Lithotripsy device |  | 1 |  | 0
| 1 | Dual-Action Lithotripsy
| 2 | Constant ultrasonic wave energy with intermittent ballistic shockwave energy(300Hz)
| 3 | Transducer maintains low surface temperature throughout entire procedure
| 4 | Variable Suction Control
| 5 | Button Operation
| 6 | Plug-and-Play System-No setup or tuning required – just one socket for superior ease of use
| 7 | Lithotripsy Probe, Reusable, 1.50 mm (4.5 Fr), length 564 mm, 1 pc |  | 3
| 8 | Ureteroscope
| 9 | Ureteroscope,7° direction of view, angled ocular, 8.6/9.8 Fr. x 430 mm, 6.4 Fr. Channel, with attachment, two-way, with two irrigation ports |  | 1
| 10 | Biopsy forceps,5 Fr. x 570 mm, double action, restoring spring |  | 1
| 11 | Light-guide cable, 3 m, CF type |  | 1
1.8 | 1.8.1 | 4K laparoscopic tower |  | 1 |  | 0
| 1 | Video processor |  | 1
| 2 | Full 4K Image Reproduction
| 3 | Wide color gamut
| 4 | Touch panel enables adjustments during procedure
| 5 | Easy registrations and loading of user presets
| 6 | Signal output (resolution) 4096 × 2160, 3840 × 2160, or 1920 × 1080 can be selected
| 7 | Signal output (transmission method) 3G-SDI or HD-SDI can be selected
| 8 | White balance adjustment
| 9 | Color tone adjustment : red. blue, chroma
| 10 | Automatic gain control
| 11 | Contrast: Normal, High, Low
| 12 | Iris Area: Auto, Center
| 13 | Image enhancement: Structural Enhancement, Edge enhancement
| 14 | Optical-digital observation with narrow band light (enhances the visibility of capillaries and other structures on the mucosal surface) by special filters.
| 15 | Camera head |  | 1
| 16 | High sensitivity CMOS sensor
| 18 | No delay 4K high-speed transmission
| 19 | One-touch auto focus function
| 20 | Electronic zoom, ×1.0, ×1.2, ×1.4, ×1.6, ×1.8, ×2.0
| 22 | Focal length f = 23.5 mm
| 23 | Optical-digital observation with narrow band light (enhances the visibility of capillaries and other structures on the mucosal surface) by special filters
| 24 | Light source |  | 1
| 25 | 300 W xenon lamp
| 26 | Automatic light control
| 28 | Emergency Halogen lamp, automatic switching to the emergency lamp
| 29 | Automatic brightness control Automatic exposure
| 30 | Optical-digital observation with narrow band light (enhances the visibility of capillaries and other structures on the mucosal surface) by special filters
| 31 | Monitor |  | 1
| 32 | Resolution: 4096 x 2160 pixels
| 33 | Aspect 17:9
| 34 | Contrast ration: 1450:1
| 36 | Viewing angle: 90º
| 37 | Input : HDMI x 1, DVI-D x 1, SDI (BNC x 5, 3G/HD/SD-SDI)
| 38 | Output: DVI-D x 1, SDI (BNC x 5)
| 39 | Rotation image, Side by Side, Picture in Picture, Picture out Picture
| 40 | High contrast images by controlling reflection and minimizing light dispersion in the LCD panel
| 41 | Telescope |  | 2
| 42 | Telescope 10 mm, 0°
| 43 | Field of view 90°
| 44 | Working length 316mm
| 45 | LCD monitor arm VESA standard |  | 1
| 46 | Separation transformer
| 47 | Bottle holder 250 mm
| 48 | Side shelf
| 49 | Light-guide cable |  | 2
| 50 | Insufflator |  | 1
| 51 | Flow rate up to 45L/min
| 52 | 3 modes: LOW 0.1-1.0 L/min, MED 1.1-1.9 and 2.0-19 L/min and HIGH 20-45 L/min
| 53 | Smoke evacuation mode OFF- turned off, LOW - flow rate 5 L/min, HIGH – flow rate 10 L/min, automatic smoke evacuation available with compatible devices
| 54 | Cavity Modes-Normal and Small
| 55 | Alarms for abnormalities, excessive pressure, tube obstruction
| 56 | Delivery with tubing for insufflation and suction
| 57 | Tubing high pressure for DIN gas cylinders |  | 1
| 58 | Suction irrigation pump |  | 1
| 59 | Irrigation up to 2 l/min
| 60 | RFID technology
| 61 | Max pump pressure 450 mmHg
| 62 | Max negative suction pressure -60 kPA
| 63 | Telescope
| 64 | Telescope 4 mm, 12°, HD, autoclavable, extra low dispersion glass in the telescopes |  | 1
| 65 | Working element, passive, for resection in saline |  | 2
| 66 | Resection sheath,8 mm, including standard obturator |  | 1
| 67 | Outer sheath,8.5 mm/26 Fr.,2 stopcocks, rotatable |  | 1
| 68 | Light-guide cable, 3 m, CF type |  | 2
| 69 | HF cable, bipolar,4 m, for Saline mode |  | 2
| 70 | HF-Resection electrode, loop, medium, 12°, 12 pieces, sterile, single use |  | 2
| 71 | HF-resection electrode, roller, 24 – 28 Fr.,12° and 30, 12 pieces, sterile, single use |  | 1
| 72 | HF-resection electrode,45° needle, 24 – 28 Fr.,12° and 30, 12 pieces, sterile, single use |  | 1
| 73 | Dissection forceps, Maryland, length of jaws: 20 mm,5x330mm, handle monopolar |  | 1
| 74 | Scissors Metzenbaum, length of jaws: 19 mm, 5x330 mm, handle monopolar |  | 1
| 75 | Grasping forceps, claw type,10x330mmlength of jaws: 45 mm, handle ratchet |  | 1
| 76 | Grasping forceps, fenestrated, length of jaws: 16.5 mm, bipolar,5x330 mm, Ergo handle |  | 1
| 77 | Grasping forceps, Manhes, length of jaws: 14 mm, bipolar.5x330 mm, Ergo handle |  | 1
| 78 | Grasping forceps, Babcock length of jaws: 31 mm, 5x330 mm, handle ratchet |  | 1
| 79 | HF electrode, hook,5 x 330 mm, monopolar |  | 1
| 80 | HF electrode, button,5 x 330 mm, monopolar |  | 1
| 81 | Suction/irrigation tube, distal holes, 5.3 x 360 mm |  | 1
| 82 | Handle, set, with suction control lever,5 mm suction channel |  | 1
| 83 | Suction/irrigation tube, distal holes,10 x 360 mm |  | 1
| 84 | Handle, set, with suction control lever,8 mm suction channel |  | 1
| 85 | Grasping forceps, needle holder, left-curved,5 mmx330 mm, Inline handle, ratchet |  | 2
| 86 | HF cable, monopolar, 3.5 m length |  | 2
| 87 | HF cable, bipolar,3.5 m length, |  | 2
| 88 | HF Electrosurgical unit |  | 1
| 89 | For open, lap and endoscopic surgery procedures
| 90 | Tissue-adaptive power output
| 91 | Instrument recognition
| 92 | Maximum high frequency power 320 W
| 93 | Monopolar-4 modes cutting and 4 modes coagulation, Bipolar-2 modes cutting and 6 modes coagulation.
| 94 | Saline mode for Trans-Cervical Resection in Saline-plug & play connection
| 95 | Touchscreen, monitoring of neutral electrode, memory of saved procedure settings
| 96 | Automatic smoke evacuation with compatible insufflator
| 97 | Disposable neutral electrodes
| 98 | Cable for neutral electrodes
| 99 | Communication cable to insufflator for smoke evacuation
| 100 | HD Medical recorder |  | 1
| 101 | Recording Media Internal HDD (500GB), External USB Storage, Network (CIFS)
| 102 | Recording Video Format MPEG-4 AVC/H.264
| 103 | Recording Audio Format AC-3/AAC LC
| 104 | Recording File Format AVCHD/MP4
| 105 | Recording Resolution 1280 × 720/59.94p, 1280 × 720/50p, 720 × 480/59.94i, 720 × 576/50i
| 106 | Input Connectors HDMI (Type A) (1), DVI-D (DVI 19-pin) (1), S VIDEO (Mini DIN 4-pin type) (1), VIDEO (BNC type) (1)
| 107 | Output Connectors HDMI (Type A) (1), DVI-D (DVI 19-pin) (1), S VIDEO (Mini DIN 4-pin type) (1), VIDEO (BNC type) (1)
| 108 | Adapter SDI to HDMI
1.9 | 1.9.1 | Operating table – gynecology/urology |  | 2 |  | 0
| 1 | Made of acid proof stainless steel
| 2 | Movable table base with central blockage with foot lever
| 3 | Mattresses made of polyurethane foam, without seams, anti-static, resistant to disinfectants and machine washing, removable, attachment without Velcro.
| 4 | Table top is X-ray translucent and enables taking X-ray pictures and patient monitoring with C-arm
| 5 | Table top min. 5 segments
| 6 | Table top surface easy for maintenance and resistant to disinfectants
| 7 | Possibility of switching places of head-rest and leg-rest segments
| 8 | Wire hand controller
| 9 | Wire hand controller controls all electric functions of the table: height adjustment, Trendelenburg and reverse Trendelenburg, lateral tilt, back-rest segment adjustment, kidney elevator, “0” position, longitudinal shift.
| 10 | Additional control system in case of power failure
| 11 | Flex and reflex positions
| 12 | Table length min 2000 mm
| 13 | Table width without side rails max 550 mm
| 14 | Mattress dimensions is adequate to table top dimension
| 15 | Mattress thickness min 50 mm
| 16 | Table top height adjustment in minimal range 750 to 1100 mm
| 17 | Trendelenburg/Reverse Trendelenburg +40° / -40°
| 18 | Backrest segment adjustment min. -40° to +80°
| 19 | Kidney elevator min. 150 mm
| 20 | Head rest segment adjustment min. -50° to + 55°
| 21 | Leg-rest segment adjustment min. -90° to + 25°
| 22 | Aside movement of leg rest segments 180°
| 23 | Lateral tilt min. +/- 30°
| 24 | Longitudinal table top shift min. 400 mm
| 25 | Rechargeable battery 24V DC
| 26 | Electric power supply 230V/50Hz
| 27 | Permissible load min. 350 kg
| 28 | Accessories
| 29 | Arm support – 1 pc
| 30 | Drip bottle holder – 1 pc
| 31 | Anesthetic frame – 1 pc
| 32 | Wristlet – 2 pc
| 33 | Abdominal strap – 1 pc
| 34 | Leg strap – 2 pc
| 35 | Holder for anesthetic tubes – 1 pc
| 36 | Knee rest (Goepel) – 1 pair
| 37 | Side support. - 2 sets
| 38 | Chest support, - 2 sets
| 39 | Pubic support, - 1 set
| 40 | Gynecological bowl, 1 pc
| 41 | Gynecological seat section, 1 pc
| 42 | Urological bowl with strainer, 1 pc
1.10 |  | Surgical lights |  |  |  | 0
| 1.10.1 | Surgical lights |  | 1 |  | 0
| 1 | LED lights
| 2 | Number of LED lights:
| 3 | Main light 72 or more
| 4 | Satellite light 72 or more
| 5 | LED lifespan: › 60.000 hours
| 6 | Warranty for LED: min 3years
| 7 | Surgical light housing made of aluminum alloy
| 8 | Temperature increase at head height: less than 1°C
| 9 | Safety glass on the lamp, resistant to cleaning agents and scratch-proof
| 10 | Homogenous light beam that doesn’t create shadows.
| 11 | Illumination in standard mode: Main light: min. 140.000 Lux Satellite light: min. 140.000 Lux
| 12 | The shape of main light and satellite light should be triangular (for better laminar flow above the surgical field).
| 13 | Technical specification of the main light
| 14 | Electrical light dimming: 30%-100%
| 15 | Endoscopic mode
| 16 | Electrical field adjustment
| 17 | Light field diameter (d10) at distance of 1 meter = 200-280 mm or more
| 18 | Color temperature form 3.500 to 5.500 K or better
| 19 | Color rendering index Ra (1-8) not less than 95
| 20 | Red rendering index R9 not less than 94
| 21 | „Skin color” rendering index R13 not less than 98
| 22 | Illumination depth without refocusing (L1+L2) at 20% not less than 1.100 mm
| 23 | Light emitting surface not less than 750 cm2
| 25 | Power supply – primary voltage 100-240V, 50Hz
| 26 | Power consumption with 24 V max. 50 W
| 27 | Technical specification of the satellite light
| 28 | Electrical light dimming:30%-100%
| 29 | Endoscopic mode
| 30 | Electrical field adjustment
| 31 | Light field diameter (d10) at distance of 1 meter = 200-280 mm or more
| 32 | Color temperature form 3.500 to 5.500 K or better
| 33 | Color rendering index Ra (1-8) not less than 95
| 34 | Red rendering index R9 not less than 94
| 35 | „Skin color” rendering index R13 not less than 98
| 36 | Illumination depth without refocusing (L1+L2) at 20% not less than 1.100 mm
| 37 | Light emitting surface not less than 750 cm2
| 39 | Power supply – primary voltage 240V, 50Hz
| 40 | Power consumption with 24 V max. 50 W
| 41 | Other
| 42 | Delivery of two (2) sterile handles per lamp
| 43 | Warranty for components 24 months
| 1.10.2 | Surgical lights with camera |  | 1 |  | 0
| 1 | LED lights
| 2 | Number of LED lights:
| 3 | Main light 72 or more
| 4 | Satellite light 72 or more
| 5 | LED lifespan: › 60.000 hours
| 6 | Warranty for LED: min 3years
| 7 | Surgical light housing made of aluminum or better
| 8 | Temperature increase at head height: less than 1°C
| 9 | Safety glass on the lamp, resistant to cleaning agents and scratch-proof
| 10 | Homogenous light beam that doesn’t create shadows.
| 11 | Illumination in standard mode: Main light: min140.000 Lux Satellite light: min140.000 Lux
| 12 | The shape of main light and satellite light should be triangular (for better laminar flow above the surgical field).
| 13 | Technical specification of the main light
| 14 | Electrical light dimming: 30%-100%
| 15 | Endoscopic mode
| 16 | Electrical field adjustment
| 17 | Light field diameter (d10) at distance of 1 meter = 200-280 mm or better
| 18 | Color temperature form 3.500 to 5.500 K or better
| 19 | Color rendering index Ra (1-8) not less than 95
| 20 | Red rendering index R9 not less than 94
| 21 | „Skin color” rendering index R13 not less than 98
| 22 | Illumination depth without refocusing (L1+L2) at 20% not less than 1.100 mm
| 23 | Light emitting surface not less than 750 cm2
| 24 | Video camera , full HD resolution placed in the main light
| 26 | Power supply – primary voltage 100-240V, 50Hz
| 27 | Power consumption with 24 V max. 50 W
| 28 | Technical specification of the satellite light
| 29 | Electrical light dimming:30%-100%
| 30 | Endoscopic mode
| 31 | Electrical field adjustment
| 32 | Light field diameter (d10) at distance of 1 meter = 200-280 mm or better
| 33 | Color temperature form 3.500 to 5.500 K or better
| 34 | Color rendering index Ra (1-8) not less than 95
| 35 | Red rendering index R9 not less than 94
| 36 | „Skin color” rendering index R13 not less than 98
| 37 | Illumination depth without refocusing (L1+L2) at 20% not less than 1.100 mm
| 38 | Light emitting surface not less than 750 cm2
| 40 | Power supply – primary voltage 240V, 50Hz
| 41 | Power consumption with 24 V max. 50 W
| 42 | Other
| 43 | Delivery of two (2) sterile handles per lamp
| 44 | Warranty for components 24 months
1.11 | 1.11.1 | Anesthesia system with monitor |  | 1 |  | 0
| 1 | Mobile anesthesia system with central breaking of wheels;
| 3 | Must have minimum two drawers;
| 4 | System console must have work surface in the middle and upper shelf for additional devices and accessories;
| 5 | Besides main power supply must have backup internal battery supply;
| 6 | It has connections for pipeline gas supply with O2, N2O and air and on the back position for placement of back up cylinder. The cylinder should be delivered with machine;
| 7 | Must have safety solution that guarantee in in exceptional circumstances the ventilation can be performed using alternative oxygen supply including undisturbed vaporizer operation;
| 8 | Must have protection from hypoxia (hypoxyguard);
| 9 | Must have integrated suction system;
| 10 | Must have anesthestic gas scavenging system;
| 11 | Must have auxiliary O2 supply (for nasal or mask oxygenation).
| 12 | Must have active ventilation modes: Volume controlled, pressure controlled, pressure controlled with volume guarantee, SIMV in variant with pressure and with volume control, SIMV with pressure control volume guarantee, pressure supported ventilation (PSV), CPAP/PSV;
| 13 | System has ability for lung recruitment either as automated procedure (e.g. mechanical breath delivery with pre-set inspiratory pressure and duration) or on other, alternative way (e.g. manually);
| 14 | System has possibility of delivery several consecutive mechnical breaths with different ventilation parameter settings or at least has possibility of preset ventilation parametes before starting procedure;
| 15 | Minimal range of tidal volume adjustment: 20 – 1500 ml;
| 16 | Minimal range of flow trigger: 0.2 - 10 L/min
| 17 | It has ability of placement two vaporisers. Sevoflurane vaporiser should be delivered with machine;
| 18 | It has flow and pressure measurement with proximally placed sensor, right in front of airway tube;
| 19 | It has measurement of inspiratory and expiratory conc. O2, CO2, N2O and agent via movable gas module, compatibile with acompanying patient monitor;
| 20 | System has its own color display on mobile arm, of size minimum 15”, touchscreen, with resolution of 1024x768 higher, The display must have a rotary command for menu selection and settings;
| 21 | Display must show following measured data: three parametric curves at the same time (pressure, flow, CO2), respiratory mechanic curve (pressure-volume or flow-volume), airway pressure, breathing rate, tidal and minute volume, compliance, bargraf of flowmeters, digitally showed pipeline pressures, measured conc.of O2, CO2, N2O and anesthetic agent, MAC value;
| 22 | It has continuous calculation with graphical/numerical display of flow range in which user can adjust flow settings according to target FiO2 and maintain oxygenation most economically on that way. This view must have also anesthetic agent consumption expediture numerically.
| 23 | Breathing system should be small, compact, designed for low flow. It has reusable CO2 absorber canister;
| 24 | There must be an alarm system with adjustable limits;
| 25 | Acompanying patient monitor that is physically mounted on the anesthesia system;
| 26 | The patient monitor has main power and back up (battery) supply. The battery should allow minimum 1h of work without main power supply;
| 27 | Patient monitor has color LCD display of size minimum 15”, touchscreen;
| 28 | It must display 8 curves on screen;
| 29 | Active measuring parameters on monitor are: ECG, saturation (SpO2), non invasive pressure (NBIP), temperature, invasive pressure x 2, measurement of sensitivity on surgical stimuli during general anesthesia (level of analgesia), status of CNS during anesthesia (anesthesia depth), NMT;
| 30 | It has trend of 24 h;
| 31 | It has adjustable alarm system according to type of patients;
| 32 | It has networking ability (LAN, WLAN) and connection with central station;
| 33 | It has minimum integrated 2 USB ports for periferal devices connection (keyboard, mouse, barcod reader, remote control etc);
1.12 | 1.12.1 | Anaesthesia monitor |  | 3 |  | 0
| 1 | Anaesthesia monitor for monitoring hypnotic and analgesic effects of anaesthesia on patient
| 2 | Portable monitor
| 3 | qCON and qNOX indexes in one stand alone anaesthetic monitor
| 4 | Indexes range min 0-99
| 6 | Touch screen with direct access to monitor functions
| 8 | Rechargable battery
| 9 | Battery life min. 2 hours
| 10 | Bluetooth conectivity
| 12 | EEG sensitivityy 44.58 nV
| 13 | Noise <±6 µVp-p (BW 70 Hz)
| 14 | CMRR >100 dB
| 15 | Input impedance >1 MΩ (@ 10 Hz)