## Technical specification
Bidder:
LOT |  | Note: Line item No. 1, 2, 3 do not have to be from the same manufacturer, but must be mutually compatible in order to be used within the same procedure |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Device for stereotactic and tomosynthesis guided biopsy and excision of the breast lesions |  | 1 |  | 0
| 1 | Digital device for performing lesion localization procedures, stereotactic (2D) and tomosynthesis (3D) guided procedures
| 2 | The device has the functionality for performing vacuum assisted biopsies and excision of breast lesions
| 3 | Automatic, motorized movement of X-ray tube for stereo and tomo procedures
| 4 | Possibility to target lesion
| 5 | Motorized and manual compression
| 6 | Cartesian (Decart) coordinate system for needle guidance with an accuracy of ± 1mm
| 7 | Stereotaxis angle min ± 15°
| 8 | Angle of tomosynthesis min 15°
| 9 | DICOM services: Print, query, storage, storage commitment, worklist
2 |  | System for the vacuum biopsy and excision of breast lesions with imaging |  | 1 |  | 0
| 1 | A System powered with constant vacuum
| 2 | A system designed to guide biopsy needle by stereotaxy and tomosynthesis to a suspected lesion
| 3 | System designed for single incision and insertion for acquisition of multiple samples
| 4 | A system designed to deliver fluids directly into an extirpation cavity through a needle
| 5 | Ability to wash extirpation cavity with physiological solution during biopsy procedure
| 6 | Automatic and manual delivery of anesthetics into the extirpation cavity during biopsy procedure
| 7 | Automated collection of tissue samples during biopsy procedure
| 8 | Automatic separation of tissue samples for transfer to the pathology laboratory
| 9 | Fully enclosed biohazard prevention system
| 10 | A system that allows the successive placement of markers into the biopsy site
| 11 | Imaging of tissue samples
| 12 | 50μm microfocus
| 13 | Monitor color, LCD
| 14 | Continuous zoom
| 15 | DICOM services: Worklist, Print, Storage
| 16 | Tissue-separation filter
| 17 | Set of 50 disposable needles
| 18 | System should be compatible to work with device described in item 1
3 |  | System for vacuum biopsy and excision of breast lesions |  | 1 |  | 0
| 1 | A system powered by constant vacuum
| 2 | A system designed to guide biopsy needle by ultrasound, stereotaxy, tomosynthesis, and magnetic resonance to a suspect lesion
| 3 | A system for single incision and insertion for acquisition of multiple samples
| 4 | A system designed to deliver fluids directly into an extirpation cavity through a needle
| 5 | Ability to wash extirpation cavity with physiological solution (saline) during biopsy procedure
| 6 | Automatic and manual delivery of anesthetics to the extirpation cavity during biopsy procedure
| 7 | Automated collection of tissue samples during biopsy procedure
| 8 | Fully enclosed biohazard prevention system
| 9 | A system that allows the successive placement of markers in the biopsy site

## Price Schedule
Bidder:
ID | Equipment name | QTY | Model offered | Unit value | Total value per line item
1 | Device for stereotactic and tomosynthesis guided biopsy and excision of the breast lesions | 1
2 | System for the vacuum biopsy and excision of breast lesions with imaging | 1
3 | System for vacuum biopsy and excision of breast lesions | 1
|  |  | Total Price per Lot |  | 0

## Scoring
BID EVALUATION SCORING SYSTEM
| CRITERIA FOR EVALUATION | Maximum No. of ponders
1. | LOWEST OFFERED PRICE | 70
This criterion is expressed according to the formula: Number of points of the bidder = (Lowest offered price) x 70 ponders / (offered price)
2. | TECHNICAL AND TECHNOLOGICAL ADVANTAGES | 30
This criterion is evaluated so that the number of points for each of the selected technical characteristics is determined below
| TOTAL : | 100
SCORING SYSTEM FOR TECHNICAL AND TECHNOLOGICAL ADVANTAGES EVALUATION
No. | Parametar/Characteristic | Maximum No. of pponders | Proof (in form of manufacturer statement or brochure or technical/product data)
1. | Performing biopsy procedure in pronation, with breast positioned in the middle of the table, with c arm integrated in table, not add on system, without interruption of the biopsy procedure for adding accessories
| No - 0 points, Yes - 15 points | 15
2. | Full 360° access to lesions in breast, both lateral and standard
| No - 0 points, Yes - 5 points | 5
4. | Vacuum biopsy device and radiography specimen device as one system with direct transfer of specimen to verification device, without manual handling of specimen
| No - 0 points, Yes - 10 points | 10