## Price Schedule
Bidder:
ID | Equipment name requested | QTY | Technical Specification Offered(model) | Offered price per unit | Offered price total
1 | Linear Accelerator | 2
2 | Patient immobilisation package for VMAT/IMRT/3D CRT treatment | 2
3 | MRI Scanner for radiotherapy planning | 1
4 | Advanced visualisation system in server client architecture | 1
5 | Patient monitor for MRI enviorment | 1
6 | Portable color doppler ultrasound system for anesthesia purpose | 1
7 | MR compatible anesthesia device | 1
8 | Rack with infusion pumps for up to 4 pumps, mobile station, MRI compatible | 1
9 | Transport Ventilator | 1
10 | Patient immobilisation package for VMAT/IMRT/3D CRT treatment (MRI adapted) | 2
11 | Patient positioning and transfer system for Brachytherapy – MRI compatible | 1
12 | Adaptation works for placement of equipment (Design and works) | 1
0

## LINEAR ACCELERATOR
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 1 | LINEAR ACCELERATOR |  | 2 |  | 0
| 1.1 | Linear accelerators designed primarily for photon therapy
General requirements, other performance and safety parameter of the linear accelerator that are not specified in this tender specification shall fulfill the requirements of the international standards:
FDA clearance and(or) CE Marking according to 93/42/EEC
| 1.2 | Treatment techniques: - 3D Conformal therapy (3D CRT) - Intensity modulated radiotherapy (IMRT) - Volumetric modulated arc therapy (VMAT /Rapid Arc). - Image guided radiotherapy (IGRT),
| 1.3 | Minimum one Photon beam energy X-6MV FFF
| 1.4 | Dose rate for photon beam min. 600 MU/min
| 1.5 | Integrated Dose Monitors System with two dosimetry channels (according to IEC 60601-2-1)
Precision of the dosimetry measurement system, for every energy ±1% or ±1 MU, whichever is greater, at a fixed dose rate.
Dose repeatability / Reproducibility better than 0.5%.
Stability / dependence on angle not more than 1% for any gantry and radiation head angle
Beam Symmetry (%) ≤ 3.0 or better
| 1.6 | MV portal dosimetry system for verification of VMAT and IMRT therapy plans. Chosen bidder must deliver and install all hardware and software component (including licences) needed in order to integrate system with existing TPS (Eclipse) and Record and Verify (ARIA) systems
| 1.7 | Photon treatment field for IMRT and VMAT min. 25cm x 30cm
| 1.8 | Gantry Rotation Range not less than 360° (±180°)
| 1.9 | Gantry Rotation speed range from 0 to 6°/sec (1 RPM) or faster
Digital Readouts: Accuracy: ±0.5°, Resolution: 0.2° or better
| 1.10 | MLC system with: - max. leaf speed 2,0 cm/sec or better - leaf resolution at isoplane - 0,5cm or better - leaf interdigitation capability
| 1.11 | Collimator rotation range min. ±90º
Digital Readouts: Accuracy 0.5° Resolution 0.2°
Anti-Collision Detection System (if there is risk for patient)
| 1.12 | Imaging techniques: - MV CBCT - MV/MV - kV CBCT
| 1.13 | Fast (iterative) CBCT reconstruction for kV imaging
| 1.14 | MV/kV imaging field of view (25x25) cm or bigger
| 1.15 | MV/kV image and treatment coincidence ≤ 0,2 cm
| 1.16 | IGRT treatment couch , controllable from control room
| 1.17 | Couch-top: -Compatible with Varian Exact indexation - Lateral travel min ±15 cm - Weight limit of 180 kg or bigger
Assisted Set-Up - Remote from Control room, Automatic positioning of the gantry, collimator, beam geometric parameters and the table to the positions specified in the field prescription
| 1.18 | Image quality phantoms (to check contrast resolution, small objects detection)
| 1.19 | QA phantom to check CBCT performance
| 1.20 | QA phantom to check linac performance
| 1.21 | Linear Accelerator should be fully integrated in existing Record and Verify system at IORS – ARIA ver. 15.5
| 2 | LINEAR ACCELERATOR - OTHER REQUIREMENTS
| 2.1 | Patient audio-video communication system
| 2.2 | UPS along with battery should be provided to give back up for safe work completion and shut down of critical devices (servers, dedicated workstations) in the case of power failure
| 2.3 | Humidity & Temperature Data logger
| 2.4 | Voltage stabilizer for linear accelerator
| 2.5 | Patient positioning laser system
| 2.6 | The closed water cooling system (complete Chiller installation)
LINEAR ACCELERATOR - TRAINING, INSTALLATION, WARRANTY
| 2.7 | On-site (customer facility) 5 day applications training
| 2.8 | Clinical training at educational clinical center on modern Radiotherapy (customized) in Europe/USA. Expected duration: 2 weeks. Participants: Two teams (doctor, physicist, radiographer)
| 2.9 | Accelerator`s maintenance – training course (off site, 2 weeks)
| 2.10 | Installation
| 2.11 | Warranty for accelerator, 1 year
| 2.12 | Warranty for other components (Wave guide, Klystron /magnetron and all vacuum parts) , 3 years
| 2.13 | Additional Maintenance (after warranty time) , 3 year
| 2.14 | Desired uptime of 95% of 365 days (24 hours), from operational acceptance, trough defect liability period, till the end of the three years’ maintenance services, has to be maintained. If downtime exceeds 5% the Supplier shall be entitled to impose penalty to the Purchaser equal to an amount of 0.2% of the total cost of the medical equipment per day for the first seven days. For the subsequent weeks the penalty amount will be doubled (0.4% of the total cost of the equipment) per day, along with extension of the Maintenance period by the excess downtime period. Total penalty amount shouldn’t be more than 10% of the total cost of the medical equipment Scheduled preventive maintenance according to manufacturer recommendations is included. Content and schedule of preventive maintenance to be done in consultation with the end-user (according manufacturer recommendations) The vendor must undertake to supply all spares for optimal upkeep of the equipment for at least 10 years after warranty period. In order to proof compliance with technical specification bidder must submit original manufacturer documentation (Brochures, Technical characteristics, Instructions for use etc.)
OTHER REQUIREMENTS
| 2.15 | Linacs should be installed in bunkers of Siemens Primus and Elekta which should be decommissioned, together with control romms and corridors should be fully adapted (In accordance to ToR, attached). Bidders are invited to visit site and acquire all informations needed to prepare bid.
| 2.16 | Decommissioning of two linac’s to be replaced (Siemens Primus and Elekta Synergy) fully in accordance with local law.
| 2.17 | Delivery and installation off all RGS hardware and software, for existing GE CT Discovery 590 RT, fully compatible with Varian respiratory gating systems installed on three linac’s at IORS.
| 2.18 | Electron density calibration phantom for CT. Low density inserts (for lung) included and inserts at density level titanium and stainless steel. |  | 1
| 2.19 | Survey meter for Gamma & Beta detection |  | 1
| 2.20 | Humidity & Temperature Data logger |  | 1
| 2.21 | Printer – A3 and A4 Color LaserJet Enterprise, network enabled, monthly page volume, minimum 5000 - 6000 pages |  | 1
| 3 | RECORD & VERIFY SYSTEM
| 3.1 | Patient management - Demographic data, Diagnosis & Staging (ICD-10, AJCC), patient record available to all members of the clinical and administrative staff, allowing them to perform activities according level of competency. Licenses |  | 5
| 3.2 | Treatment prescriptions create and edit treatment plans (treatment dose, fractions, regime, reference point information...). Licenses. |  | 5
| 3.3 | Review of reference images, analysis, Offline Review, multi-modality image review. Licenses. |  | 5
| 3.4 | QA chart, pre-treatment audits that ensure treatment plan’s validity. Licenses. |  | 5
| 3.5 | Import, Capture and Review of Documents: written document, scanned documents, photos. Licenses. |  | 5
| 3.6 | Manage the tasks and appointments of department resources and staff. Licenses. |  | 5
| 3.7 | Advanced Appointment Management (time, resources, staff, activities, patients, reports...). Licenses. |  | 5
| 3.8 | Tools for Report management |  | 5
| 3.9 | Electronic Plan Approval |  | 5
| 3.10 | DICOM-standard design for extensive image modality support – CBCT (iCBCT)
| 3.11 | DICOM Bundle (in according the newest manufacturer’s DICOM conformance statements), DICOM management with maximum functionality (Export/Import and Query/Retrieve and other service if is available) |  | 5
| 3.12 | Data Management Automation
| 3.13 | Automatically find and retrieve data acquired during therapy and/or follow-up for specific patients |  | 5
| 3.14 | Automatically route DICOM objects between locations based upon configurable “triggers” |  | 5
| 3.15 | Automatically generate structure set and external body contour during import of volume |  | 5
| 3.16 | Automatic creation and sending of a session message notifying specific users of newly imported data |  | 5
| 3.17 | Adaptive Calculation and Tracking for Offline Review |  | 5
| 3.18 | User R&V workstations over minimum system requirements (mission critical devices, sequencer or 4D console for example) , for access any activities at same time from any member of the clinical and administrative staff ) |  | 5
| 3.19 | Warranty time, 1 year
| 3.20 | Additional Maintenance (after warranty time) , 3 years
| 4 | TREATMENT PLANNING SYSTEM (TPS)
| 4.1 | TPS is dedicated for 3D conformal therapy, intensity-modulated radiation therapy (IMRT) and VMAT (Rapid Arc)
| 4.2 | Multi-modality image registration
| 4.3 | Manual and Automatic Contouring
| 4.4 | 3D visualization incl. Multi Planar View Real-time coronal/sagittal sections
| 4.5 | Multi-Criteria optimization
| 4.6 | Experience based, automated, plan generation
| 4.7 | Multi Leaf Collimator Planning
| 4.8 | Easy-to-use templates and optimization tools that substantially sped up treatment planning for complex radiotherapy treatments like volumetric modulated arc therapy (VMAT), intensity-modulated radiotherapy (IMRT) and image-guided radiotherapy (IGRT).
| 4.9 | Fast treatment setup and calculation of VMAT plans
| 4.10 | Multiple Dose Displays
| 4.11 | Dose Volume Histograms (DVH) and interactive 3-D surface-rendering image
| 4.12 | Tools for beam data commissioning
| 4.13 | Portal Dosimetry analyze for VMAT (Pre-Treatment verification)
| 4.14 | DICOM Bundle (in according the newest manufacturer’s DICOM conformance statements), DICOM management with maximum functionality (Export/Import and Query/Retrieve and other service if is available)
| 4.15 | Treatment planning workstations, including licenses for 3D-CRT planning, IMRT planning, VMAT planning including advanced optimization algorithms for VMAT, QA & Treatment plan verification. All necessary licenses for OS and AVS |  | 2
| 4.16 | UPS along with battery should be provided to give back up for safe work completion and shut down of critical devices (servers, dedicated workstations) & all treatment planning workstations in case of power failure
| 4.17 | Treatment planning system, dosimetry & beam data commissioning, measurements, modelling and verification (of all algorithms and energies) for 3D CRT, IMRT, VMAT including FFF beams.
| 4.18 | FDA clearance and European certificate (CE)
| 4.19 | On-site (customer facility) 1 day applications training
| 4.20 | Clinical training at educational clinical center on modern RT (customized) in Europe/USA. Expected duration: 2 weeks. Participants: Two teams (doctor, physicist).
| 4.21 | Installation
| 4.22 | Measurement of beam data, creating and verifying model of the offered accelerator into the existing (Eclipse 15.6) and newly delivered treatment planning system (commissioning, modeling and model verification)
| 4.23 | Warranty time, 1 year
| 4.24 | Additional Maintenance (after warranty time) , 3 years
| 5 | CONTOURING & PLAN REVIEW
| 5.1 | Multimodality image registration (CT, MR…). deformable registratiom.
| 5.2 | Image multimodality fusion tools (CT-CT, CT-MR, CT-PET…)
| 5.3 | Image Visualization and contouring tools
| 5.4 | Contouring Tools Including Manual and Automatic Contouring, 3D Variable Margins
| 5.5 | Advanced contouring tools (Fully automatic segmentation of targets and organs at risk or Atlas-based Auto Segmentation or similarly)
| 5.6 | Imaging Tools including Window and Level controls, Transverse, Sagittal and Coronal viewing, Real-time coronal/sagittal sections
| 5.7 | Multi-Plan Viewing - Compare Dose distributions and DVH for minimum two alternative plans (same or different modalities)
| 5.8 | Multiple Dose Displays, Analyze tools
| 5.9 | View DVH, Ports
| 5.10 | Electronic Plan Approval
| 5.11 | DICOM Export Bundle (according to manufacturer’s DICOM statement)
| 5.12 | Workstations for Contouring & Plan Review including licenses and all mentioned feature. All necessary OS and AVS (if manufacturer recommend) |  | 3
| 5.13 | On-site (customer facility) 2 day applications training
| 5.14 | Installation
| 5.15 | Warranty time, 1 year
| 5.16 | Additional Maintenance (after warranty time) , 3 years
| 6 | EQUIPMENT FOR THE VMAT PLAN VERIFICATION
| 6.1 | Multifunctional detector system with possibility to adapt detector to isocenter for pre-treatment QA IMRT&VMAT plans, DVH analysis and gantry angle detection
| 6.2 | Flat panel with minimum 1400 detectors
| 6.3 | Minimal field size or Active area 24 x 24 cm2 or larger
| 6.4 | Small detector volume 0.1 cm3 or less
| 6.5 | Detector to detector “resolution” 7 mm or better
| 6.6 | Zero dead time, long-term radiation resistance of detectors
| 6.7 | Dose Resolution: 1mGy or better on dose rate 0.5 Gy/min or better
| 6.8 | Wide range of dose rate measurements
| 6.9 | Sampling time, acceptable for VMAT plan verification
| 6.1 | As low as possible angular dependency.
| 6.11 | Phantom, water equivalent compatible with flat panel device if need it for routine measurements.
| 6.12 | Data transfer via Ethernet and TCP/IP protocol or WiFi
| 6.13 | Fully integrated in manufacturer software
| 6.14 | Fast data acquisition in real time
| 6.15 | Provides comprehensive evaluation routines with gamma index analyses and other methods
| 6.16 | Advanced tools for data displaying, 1D, 2D and 3D data visualization, profiles, isodose, contours, 2D/3D dose distributions and histograms, etc.
| 6.17 | Fully automated fluency analysis
| 6.18 | Advanced software for comprehensive 3D patient anatomy based dose analysis
| 6.19 | Analyze 3D patient dose distributions based on DVH
| 6.20 | DICOM Bundle in according the newest manufacturer’s DICOM conformance statement, DICOM management with maximum functionality (Export/Import and Query/Retrieve and other service if is available)
| 6.21 | Capable of data import from all customary planning systems (from XiO, Monaco, Eclipse...), DICOM or own data formats, customary images, etc.
| 6.22 | PC Workstation, Monitor, Cables (if there is dedicated cable beetwen Workstation/ControlUnit and measured device, mandatory delivery 2 cables as spare cables).
| 6.23 | On-site (customer facility) 2 day applications training
| 6.24 | Installation
| 6.25 | Warranty time, 1 year
| 6.26 | Additional Maintenance (after warranty time) , 3 years
| 7 | EQUIPMENT FOR FAST QA ACCELERATORS
| 7.1 | Dedicated for fast radiotherapy machine QA – measures CAX dose, dose rate, checks flatness and symmetry, radiation quality with FFF 6X compensator
| 7.2 | Measurement volume 0.2 cm3 or less
| 7.3 | Vented ionization chambers
| 7.4 | Field sizes (20 x 20) cm and 10 x 10cm
| 7.5 | Energy range: (4-25) MV photons, (4-10)MV photons with FFF compensator
| 7.6 | Dose rate range:(0.5–10)Gy/min, (1–25)Gy/min with FFF compensator
| 7.7 | Long-term radiation resistance of ionization chambers
| 7.8 | Fully integrated in manufacturer software
| 7.9 | Immediate display of measurement results
| 7.10 | Automated archiving in database
| 7.11 | PC Workstation, Monitor, Cables (if there is dedicated cable beetwen Workstation/ControlUnit and measured device, mandatory delivery 1 cable as spare cables)
| 7.12 | Installation and on-site (customer facility), 1 day applications training
| 7.13 | Warranty time, 1 year
| 7.14 | Additional Maintenance (after warranty time) , 3 years
| 8 | EQUIPMENT FOR DAILY QA ACCELERATORS
| 8.1 | Dedicated for daily radiotherapy machine QA, checks flatness, symmetry, energy, output constancy, wedge angle of dynamic, virtual and fixed wedges as well as the congruency of light and radiation fields etc. MLC QA included.
| 8.2 | Hardware
| 8.3 | Flat panel with minimum 450 ionization chambers
| 8.4 | Small ionization chamber volume 0.1 cm3 or less
| 8.5 | Vented ionization chambers
| 8.6 | Gantry mount holder for positioning
| 8.7 | One build-up plate each for photon and electron energies
| 8.8 | Minimal field size 25 x 25 cm2
| 8.9 | Energy range: Co60, 4 MV–20 MV Photons, 4 MeV-21 MeV Electrons
| 8.10 | Dose rate range: 0.1–20 Gy/min
| 8.11 | Signal to noise ratio: better than 1 % with 10 cGy integrated dose
| 8.12 | Dose Resolution: 1mGy or better on dose rate 0.5 Gy/min
| 8.13 | Sampling time acceptable for daily QA (minimum 200 ms)
| 8.14 | Readout: parallel and synchronous with no dead time
| 8.15 | Long-term radiation resistance of ionization chambers
| 8.16 | Interface to PC: Ethernet RJ-45 (direct connection or via network)
| 8.17 | Fully integrated in manufacturer software
| 8.18 | Software |  | 1
| 8.19 | Real time measurement
| 8.20 | Automated archiving in database
| 8.21 | Advanced grouping, filtering & sorting for analysis
| 8.22 | Different QA protocols for analysis of penumbra, symmetry, flatness and field size for simplified measurements
| 8.23 | Checks flatness, symmetry, output, wedge angle of dynamic, virtual and fixed wedges as well as the congruency of light and radiation fields. Compares all parameters with a reference data record and displays deviations.
| 8.24 | PC Workstation, Monitor, Cable (if there is dedicated cable beetwen Workstation/ControlUnit and measured device, mandatory delivery 1 cable as spare cables) |  | 1
| 8.25 | On-site (customer facility) 1 day applications training
| 8.26 | Installation
| 8.27 | Warranty time, 1 year
| 8.28 | Additional Maintenance (after warranty time) , 3 years
| 9 | ABSOLUTE DOSIMETRY
| 9.1 | High precision reference class electrometer for absolute dose measurements and calibration of the linear accelerators. |  | 2
| 9.2 | Small ionization chamber. Chamber volume 0.01-0,015cc, brackets/adapters for scanning water phantom and 1D water phantom |  | 2
| 9.3 | Small ionization chamber, waterproof. Chamber volume less than 0.005cc, brackets/adapters for scanning water phantom and 1D water phantom |  | 2
| 9.4 | Small ionization chamber. Chamber volume 0.03-0,04cc, brackets/adapters for scanning water phantom and 1D water phantom |  | 2
| 9.5 | Dosimetry diode for dose measurements in small fields |  | 2
| 9.6 | Extension triaxial shielded cable 15m |  | 5
| 9.7 | Waterproof Farmer ionization chamber 0.6cc-0.7cc, with holder for 1D and scanning water phantom |  | 2
| 9.8 | Triaxle shielded chamber cable (length >15m). |  | 4
| 9.9 | Set triaxiale adapters: TNC Male to BNC Female, TNC Female to BNC male TNC Male to TNC Female, BNC Female to BNC male |  | 5
| 9.10 | Set of build-up caps for beam measurements for treatment planning system, commissioning – 2 sets. |  | 1
| 9.11 | Manual 1D water phantom for absolute dose measurements with chamber holders for all chambers |  | 1
| 9.12 | Measurement depth can be adjusted manually in < 0.5 mm steps.
| 9.13 | Measurement depth can be read out with digital display.
| 9.14 | Digital Thermometer with, Resolution 0.1°C, Accuracy ±0.2°C Barometer or digital Barometer, Accuracy ±1 mbar Acceptable option: Digital Barometer with thermometer |  | 1
| 9.15 | Dosimetry PC Workstation, Monitor, cables, |  | 1
| 9.16 | Installation |  | 1
| 9.17 | Warranty time, 1 year
| 9.18 | Additional Maintenance (after warranty time) , 3 years

## PATIENT IMMOBILISATION
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
2 | 2 | PATIENT IMMOBILISATION PACKAGE FOR VMAT/IMRT/3DCRT TREATMENT |  | 2 |  | 0
| 2.1 | Treatment base plate for Head & Neck made of low density materials. Full technical solution for use to prone & supine. |  | 2
| 2.2 | Head & Neck thermoplastic masks for 3D CRT |  | 50
| 2.3 | Head & Neck thermoplastic masks dedicated for VMAT/IMRT ( more rigid ) |  | 50
| 2.4 | Treatment base plate for Neck & Shoulders made of low density materials. Full technical solution for use to prone & supine. |  | 2
| 2.5 | Neck & Shoulders, thermoplastic masks for 3D CRT |  | 50
| 2.6 | Neck & Shoulders, thermoplastic masks dedicated for VMAT/IMRT (more rigid) |  | 50
| 2.7 | Headrests, min. 5 different sizes & shapes |  | 2
| 2.8 | Wing Board |  | 2
| 2.9 | Breast board with adjusting inclination. Full technical solution for use to supine |  | 1
| 2.10 | Breast board. Full technical solution for use to prone |  | 1
| 2.11 | Treatment base plate for Breast (Thorax)
| 2.12 | Thermoplastic masks for breast (Thorax) |  | 50
| 2.13 | Treatment base plate for Pelvis & Abdomen made of low density materials. Full technical solution for use to prone & supine. |  | 2
| 2.14 | Pelvis & Abdomen, thermoplastic masks for 3D CRT |  | 50
| 2.15 | Pelvis & Abdomen, thermoplastic masks dedicated for VMAT/IMRT (more rigid) |  | 50
| 2.16 | Pads for the legs and knees, min. 5 different sizes & shapes |  | 2
| 2.17 | Carbon fiber Baseplate for fixation of extremities |  | 2
| 2.18 | Extremity masks, small |  | 50
| 2.19 | Vacuum Bag Cushion small, medium, large sizes (with arm rest adapter for use with vacuum bag & indexing possibilities). |  | 2
| 2.20 | Vacuum pump |  | 1
| 2.21 | Pediatric small base blate for head, neck and shoulder. Full technical solution for use to prone & supine. |  | 2
| 2.22 | Pediatric long base plate for head, neck and shoulder. Full technical solution for use to prone & supine. |  | 2
| 2.23 | Pediatric headrests, min. 4 different sizes & shapes |  | 2
| 2.24 | Pediatric Head & Neck small thermoplastic masks |  | 30
| 2.25 | Pediatric Head & Neck medium thermoplastic masks |  | 30
| 2.26 | Pediatric Head, Neck & Shoulders small thermoplastic masks |  | 30
| 2.27 | Pediatric Head, Neck &Shoulders medium thermoplastic masks |  | 30
| 2.28 | Pediatric Supine, Cushion for shoulder, Breast, board Carbon fiber with all required |  | 2
| 2.29 | Pediatric vacuum bag cushion small, medium, large sizes (with arm rest adapter for use with vacuum bag & indexing possibilities). |  | 2
| 2.30 | CT Simulation Skin Markers, non-metallic pellets, three diameter sizes (~2mm - 100 pcs, ~3mm -100 pcs , 3-5 mm 100 pcs) |  | min. 300 pcs
| 2.31 | Skin Marking Labels, CT Mark Wire Labels, 2.0 mm, 300 cm long, |  | 50
| 2.32 | Medical Precision Pen Marking |  | 2
| 2.33 | Tattoo kits (ink, capsule, needle pkg) |  | 3

## MR MON UZ
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
3 | 3 | MRI Scanner for radiotherapy planning |  | 1 |  | 0
| 1 | МAGNET 1.5Т
| 1.1 | Superconductive magnet with active shielding
| 1.2 | Tunnel bore diameter in isocenter – min. 70 cm
| 1.3 | Guaranteed homogeneity values of magnetic field in accordance with V.R.M.S. method: - DSV, for spheric volume with diameter of 10 cm: not bigger then 0,02 ppm - DSV, for spheric volume with diameter of 40 cm: not bigger then 1 ppm
| 1.4 | Magnet length including covers: no longer than 180 cm
| 1.5 | Field Of View (FOV)– not less then 50 cm in each direction X,Y,Z
| 1.6 | Zero boíl off technology
| 1.7 | Noise reduction in tunnel
| 2 | GRADIENT SYSTEM
| 2.1 | Gradient performance (Amax x SRmax)- minimum 4.000 mT²/m²/ms per axis, where Amax is the maximum gradient amplitude for each axes and SRmax is the highest value of slew rate for each axes)
| 2.2 | Maximum power of gradient amplifier, min. 1.2MW
| 3 | RF SYSTEM
| 3.1 | Number of independent receiver channels that can be used simultaneously in one single scan and in one single FOV without patient table repositioning, each generating an independent partial image, minimum 32
| 3.2 | Number of physical coil elements that can be simultaneously connected during one exam, minimum 128
| 3.3 | The resolution of the signal receiver is not less than 32 bits
| 3.4 | RF transmitter power not less than 16 kW
| 4 | RF COILS with requested number of coil elements- number of requested coil elements must correspond to the number of physical elements of particular requested coil (the sum of elements of multiple coils/ combination of coils is not acceptable- except for coil under point 4.2.)
| 4.1 | Latest generation of integrated coil technology (GEM, Tim 4G, dStream, etc.) which supports simultaneous acquisition from different coils
| 4.2 | Separate, dedicated anterior body coil (one or combination of two) for body scanning in range of min. of 60 cm (from neck to pelvic) and with minimum 24 coil elements in total (one or combination of two).
| 4.3 | Separate, dedicated multi-channel coil for the imaging of head and neck with minimum 19 coil elements
| 4.4 | Spine coil with minimum 24 coil elements
| 4.5 | Set of two different size multi-channel flex coils (2 pcs small and 2 pcs large) with minimum 4 elements each for imaging of the head for radiotherapy planning purpose
| 4.6 | Separate, dedicated multi-channel flex coil with minimum 16 elements for imaging of the head, abdomen and pelvic for radiotherapy planning purpose
| 4.7 | Separate, dedicated multi-channel coil for the imaging of breast, with minimum of 16 coil elements
| 5 | PATIENT TABLE
| 5.1 | Dockable table which enables easy patient preparation outside the examination room. Dockable table should have all coil connectors integrated on it, so complete patient preparation including coil positioning and connecting to the table could be done outside the examination room
| 5.2 | Maximal scan range- not less than 200cm
| 5.3 | In bore ventilation and lightening
| 5.4 | In room loudspeaker adjustment and bidirectional voice communication
| 5.5 | Patient video monitoring system
| 5.6 | Max. Patient weight for vertical and horizontal table movement, not less then 200 kg
| 5.7 | Coil storage cart or shelf
| 6 | ACQUISITION WORKPLACE (OPERATOR CONSOLE)
| 6.1 | DICOM standard for communication between devices from different with following functionalities:
DICOM STORAGE, DICOM QUERY/RETRIVE, DICOM PRINT, MODALITY WORKLIST
| 6.2 | Two (2) Color LCD monitors with resolution not less than 1.3MP with diagonal size not less than 19“ or one (1) color LCD monitor with diagonal size not less than 24“ with 1920x1200 resolution
| 6.3 | CD or DVD for DICOM archiving
| 6.4 | Acquisition packages for basic Neurology, Angiography, Cardiology, Radiology, Oncology, Musculoskeletal and Pediatric examination
| 6.5 | Basic 3D post-processing techniques – VRT, Multi Planar Reconstruction (MPR), Maximum Intensity Projection
| 6.6 | Integration (composing) of several diagnostic images in 1 “Full Format” (longer anatomical structures and blood vessels)
| 7 | SEQUENCES AND IMAGING TECHNIQUES
| 7.1 | Turbo Dark Fluid (FLAIR)
| 7.2 | Susceptibility-Weighted Imaging
| 7.3 | Magnetization Transfer Contrast
| 7.4 | Diffusion Tensor Imaging
| 7.5 | Neuro perfusion
| 7.6 | High resolution diffusion- RESOLVE, FOCUS or equivalent
| 7.7 | Respiratory (breathing) and patient movement artifact elimination- PROPELLER or BLADE or MultiVane or JET or equivalent)
| 7.8 | 2D and 3D MRCP
| 7.9 | Dynamic Liver Imaging (THRIVE, VIBE, LAVA-Flex or equivalent
| 7.10 | Ultrafast Dynamic Contrast-enhanced Imaging (TWIST Vibe, DISCO or equivaent)
| 7.11 | 3D Contrast angiography in real time (TRICKS or TWIST or 4D-TRACK or Freeze Frame or equivalent)
| 7.12 | Angiography with synchronized bolus tracking (SmartPrep or BolusTrak or Care Bolus or equivalent)
| 7.13 | Non-contrast angiography (for peripheral blood vessels, abdomen and neuro exams- NATIVE or Inhance or TRANCE or Time-SLIP or equivalent)
| 7.14 | Fat and water separation technique (using DIXON or similar technique)
| 7.15 | Technique to reduce artifacts caused by metal implants- (O-MAR, Advanced WARP, MAVRIC or equivalent)
| 7.16 | Single voxel spectroscopy, Chemical Shift Imaging (2D and 3D)
| 8 | ADDITIONAL EQUIPMENT, WORKS AND SERVICES
| 8.1 | RF cabin for electromagnetic protection (damping value of min 90dB). RF cabin contains complete cabin with antistatic floor, door, window, antimagnetic lightening, all necessary filters, professional airconditioning, and all final works, for running requirements of MR scanner for work with the patients
| 8.2 | Chiller for MR gradient system cooling
| 8.3 | Electrical distribution cabinet for MR system
| 8.4 | Injector for MR studies
| 8.5 | Antimagnetic stretcher for patient transport
| 8.6 | Indexing plate (the same indexation as on linacs) with body coil holder, head coil holder
| 8.7 | Laser simulation system- Laser system should include a movable sagital plane, a movable coronal plane which is projected from both lateral sides and a fixed transverse plane. The software should be operated via touch screen interface. The adjustment of the laser should be remote-controlled
| 8.8 | 10 days on site training Training for 4 physicians, technicians or physicist for min 5 days in Hospital with same or similar system
4 | 4 | Diagnostic workstations in client – server architecture, with server and 4 clients computer, with 2D, 3D basic and advanced visualization applications from different modalities MR, CT, X-ray system, ultrasound |  | 1 |  | 0
| 1.1 | 2D and 3D viewing of images from several modalities (CT, MR, PET/CT, CR, X-ray images);
| 1.2 | VRT, MIP thin/thick, MPR thin / thick, interactive slice thickness change
| 1.3 | 3D SSD technique
| 1.4 | MR Radiology Workflows: predefined layouts for Head, C-Spine, T-Spine, L-Spine, Whole Spine, Breast, Prostate, Abdomen, Hip and Knee scans, Composed Images
| 1.5 | MR Evaluation: Subtraction, MeanCurve, Image Filter, 2D/3D Distortion Correction.
| 1.6 | MR Spectroscopy evaluation (Single voxel spectroscopy, Chemical Shift imaging)
| 1.7 | MR brain perfusion evaluation
| 1.8 | MR 3D lesion segmentation (semi-automated volumetric evaluation of tumors, lymph nodes and metastases
| 1.9 | MR application for quantitative evaluation of functional response to treatment
| 1.1 | Advanced MR application which create intuitive parameter maps for efficient reading of dynamic series of prostate imaging
| 1.11 | Image review and contouring on CT, MR, PET, PET-CT, CBCT, 4D CT, time-resolved CT / MR images
| 1.12 | Direct4D for creation of tMIP, tminIP, AverageCT, phase splitting, 4D contouring, tumor motion analysis, semi-automatic generation of ITV / Mid Ventilation phase
| 1.13 | Multiple contouring tools, incl. auto contouring for organs-at-risk (brain, heart, lungs, liver, kidneys, femoral heads) and parallel contouring on multiple images
| 1.14 | Deformable Image Registration including contour warping and advanced evaluation of registration (spyglass, deformation vector map, colorwash)
| 1.15 | CT like image reconstruction- CT free approach enabling a straight forward workflow for target definition, patient marking, beam placement and patient positioning.
Option provides density information for dose calculations. Eliminates the needs of registration between CT and MR to avoid registration errors.
| 1.16 | LCD high resolution monitor 2MP min. 24“, 2 pcs. for each client, 8 pcs in total plus one monitor min. 19” for administration for each client, 4 pcs in total
5 | 5 | Patient monitor for MR enviorment |  | 1 |  | 0
6 | 6 | Portable color doppler ultrasound system with height adjustable cart, printer and with convex and linear probes for anesthesia purpose |  | 1 |  | 0

## MR compatible anesthesia device
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
7 | 7 | MR compatible anesthesia device |  | 1 |  | 0
| 1.1 | MR compatible anesthesia device
| 1.2 | Work with all patient categories: Neonatal, Pediatric and Adults
| 1.3 | Breathing (circuit) system
| 1.4 | The device should be capable of low-flow anesthesia
| 1.5 | Work with the gas type: O2, N20, Air
| 1.6 | Device should be on the mobile console, with brake
| 1.7 | Device should have three drawers for storing supplies
| 1.8 | Electricaly driven mechanical ventilator
| 1.9 | Device should be designed and tested to work in a magnetic field of MRI device strength of 1.5 Tesla and 3 Tesla
| 1.10 | The possibility of operating in the field strength 40mTesla
| 1.11 | Device must have a strong indicator alarms that are visible from different angles and with greater distances (ie. Control room)
| 1.12, | Integrated battery with 120 minutes capacity
| 1.13 | Ventilaton modes:
| 1.14 | Manual
| 1.15 | Spontaneous
| 1.16 | Volume controlled
| 1.17 | Pressure controlled
| 1.18 | Pressure support
| 1.19 | SIMV-PS
| 1.20 | Apnea ventilation
| 1.21 | Possibility of mechanical ventilation in case of interruption of gas supply from the central gas supply system
| 1.22 | Pressure limit: 15 do 70 cm H2O
| 1.23 | PEEP: 0 to 20 cm H2O
| 1.24 | Breathing frequency: 4 – 60
| 1.25 | Minimal O2 concentration in O2/N2O mixture: 23 %
| 1.26 | Digital display of gas flow (O2, Air, N2O)
| 1.27 | Integrated security function that provides protection against hypoxia
| 1.28 | Color screen for display of ventilation parameters and curves
| 1.29 | Measured parameters:inspiratory O2 concentration, breathing frequency,expiratory tidal volume, minute volume, peak inspiratory pressure (Ppeak),
| 1.30 | PEEP, mean pressure and plateau pressure
| 1.31 | Device has integrated endotracheal suction unit
| 1.32 | Possibility for connection two vapors,
| 1.33 | Device is supplied with one sevoflurane vapor
| 1.34 | Device should be delivered with all neccessary accessories: supply hoses, patient hoses, connectors
| 1.35 | One dimension of station(wide or depth) must be less than 90cm because of space limitation

## Infusion pump Ventilator
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
8 | 8 | Rack with infusion pumps for up to 4 pumps, mobile station, MRI compatible |  | 1 |  | 0
| 1.1 | Device type - MR workstation compatible with infusion pumps, for up to four pumps, mobile station
| 1.2 | Upgrading the composition by compulsory delivery of a compatible workstation with the offered pumps
| 1.3 | A magnetic field meter that ensures that the MRI workstation is always placed at a safe distance from the magnet
| 1.4 | Maximum magnetic field density 20 mT = 200 G. Vector field measurement in three axes during operation of the device
| 1.5 | Radio interference suppression in accordance with IEC EN 60601-1-2 and IEC EN 60601-2-24
| 1.6 | Power supply: 100/240 V; 56/60 Hz, power up to max. 80 VA
| 1.7 | Integrated battery, battery life approx. 10 hours
| 1.8 | Central Beep Alarm for the Workstation and Magnet Indicator
| 1.9 | Optical alarm - minimum 3 LEDs; Magnetic field (green, yellow, red), Error (red), Empty battery (yellow)
| 1.10 | Operating elements and display status on the cover:
| 1.11 | Battery status display, Battery / operating network display, Switching station MRI on battery power off, Error indicator, Battery maintenance starter
| 1.12 | Integrated magnetic field monitoring system independent of the MRI workstation with special status indicator and warning device
| 1.13 | Vector recognition of the magnetic field in three axes
| 1.14 | Supports safe scanning of the scanner in the following conditions:
| 1.15 | Static magnetic field MR scanner max. 3 Tesle
| 1.16 | MR spacer gradient max. 20 mT / 200 Gaus
| 1.17 | A maximum average specific full body (SAR) rate of 4W / kg for 30 minutes of scanning
| 1.18 | An increase in temperature of less than 1 ° C on an MRI workstation
| 9 | Transport ventilator |  | 1 |  | 0
| 1.1 | Work with all patient categories: Pediatric and Adults
| 1.2 | Used for transporting patinent
| 1.3 | maximum weight 5±1 kg
| 1.4 | At least 4 inch screen display with possibility to rotate view
| 1.5 | Ventilation modes: NIV, PS, VC-CMV, VC-AC, VC-SIMV/PS, CPR
| 1.6 | Insp flow 100 l/min
| 1.7 | PEEO 0-20 mbar, flow trigger 3-15 l/min, psupp 0-35 mbar
| 1.8 | Battery
| 1.9 | storing conditions: temp -20 to +50, humidity 5-95%, pressure 620-1100 hPA

## PATIENT IMMOBILISATION MRI
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
10 | 10 | PATIENT IMMOBILISATION PACKAGE (MRI ADAPTED) FOR VMAT/IMRT/3DCRT TREATMENT |  | 2 |  | 0
| 1.1 | Treatment base plate for Head & Neck made of low density materials. Full technical solution for use to prone & supine. |  | 2
| 1.2 | Head & Neck thermoplastic masks for 3D CRT |  | 25
| 1.3 | Head & Neck thermoplastic masks dedicated for VMAT/IMRT ( more rigid ) |  | 25
| 1.4 | Treatment base plate for Neck & Shoulders made of low density materials. Full technical solution for use to prone & supine. |  | 2
| 1.5 | Neck & Shoulders, thermoplastic masks for 3D CRT |  | 25
| 1.6 | Neck & Shoulders, thermoplastic masks dedicated for VMAT/IMRT (more rigid) |  | 25
| 1.7 | Headrests, min. 5 different sizes & shapes |  | 2
| 1.8 | Wing Board |  | 2
| 1.9 | Breast board with adjusting inclination. Full technical solution for use to supine |  | 1
| 1.10 | SBRT immobilisation - full technical solution (with two bridges) |  | 1
| 1.11 | Treatment base plate for Breast (Thorax) |  | 1
| 1.12 | Thermoplastic masks for breast (Thorax) |  | 25
| 1.13 | Treatment base plate for Pelvis & Abdomen made of low density materials. Full technical solution for use to prone & supine. |  | 2
| 1.14 | Pelvis & Abdomen, thermoplastic masks for 3D CRT |  | 25
| 1.15 | Pelvis & Abdomen, thermoplastic masks dedicated for VMAT/IMRT (more rigid) |  | 25
| 1.16 | Pads for the legs and knees, min. 5 different sizes & shapes |  | 2
| 1.17 | Fiducial markers for abdomen and pelvis, 0.8-1mm x 4.5-5mm. |  | 30
| 1.18 | Functional kit for inserting fiducial abdominal and pelvis markers. |  | 3
| 1.19 | Possibility of CT and MRI imaging of a patient with inserted fiducial markers.

## PATIENT POSITIONING AND TRANSFE
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
11 | 11 | PATIENT POSITIONING AND TRANSFER SYSTEM FOR BRACHYTHERAPY |  | 1 |  | 0
| 1.1 | Patient Positioning and Transfer System for Brachytherapy that should allow for the quick and easy transportation and transfer of patients from the brachytherapy room to the imaging device.
| 1.2 | Advanced technological solution for raising, lowering and transferring patient from The Positioning and Transfer System to imaging device table.
| 1.3 | Patient Positioning and Transfer System for Brachytherapy, with possibility that the user can only use MRI, CT alone, or both modalities as needed for same patient.
| 1.4 | Positioning and Transfer System for Brachytherapy have abillity to use both imaging modalities regardless of the type of MRI (CT) table i.e. whether the table is detachable or in a fixed position
| 1.5 | The transfer and transportation of the patient should be easy and require as less as posible of humans resources.
| 1.6 | The minimum risk of unwanted patient movement during brachytherapy procedures, i.e. displacement of inserted applicators and needles in the patient.
| 1.7 | Holder pair (left/right). Possibility of mounting the side holders by which the patient retains the same position during the imaging procedure and during brachytherapy treatment. (1 set + 1 additional set)
| 1.8 | Ability to insert a brachytherapy applicator during brachytherapy procedures, in a gynecological position i.e. possibility of table use as a gynecological table.
| 1.9 | Suggested revision: Multipurpose (universal) clamping bracket available for various GYN post-loading clamps.
| 1.10 | Ability to insert needles for brachytherapy and perform interstitial prostate brachytherapy without having to transfer the patient to another table.
| 1.11 | The possibility of transferring the patient to the imaging table laterally, a longitudinal option is desirable
| 1.12 | Accessory set for for better and more reliable patient positioning (cushions, foot rest, attachment of patients arm boards, …) a.) Items to be included with transfer system: Cushions, Foot rest, Patient Safety Straps. 1 set +1 additional set b.) Anesthesia Accessory: Patients arm boards Accessory item: Arm Board w/ cushion (single arm board). 1 set + 1 additional set
| 1.13 | Instalation and training