## Price Schedule
Bidder:
ID | Equipment name | QTY | Technical Specification Offered(model) | Offered price per unit | Offered price total
1 | Open bore 3T MRI Scanner | 1 |  |  | 0
2 | Multislices CT scanner for diagnostics and RT procedures with CT Simulation Laser System | 1 |  |  | 0
3 | Diagnostic workstation in client | 1 |  |  | 0
4 | 3D Digital Mammography System 1 | 1 |  |  | 0
5 | 3D Digital Mammography System 2 | 1 |  |  | 0
6 | Ultrasound system 1 | 2 |  |  | 0
7 | Ultrasound system 2 | 1 |  |  | 0
8 | Ultrasound system 3 | 1 |  |  | 0
9 | Conference system | 1 |  |  | 0
10 | Adaptation works for placement of equipment (Design and works) | 1 |  |  | 0

## MRI
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 1 | Open bore 3T MRI Scanner |  | 1 |  | 0
| 1 | МAGNET 3Т
| 1.1. | Superconductive magnet with active shielding
| 1.2. | Tunnel bore diameter in isocenter – min. 70 cm
| 1.3. | Guaranteed homogeneity values of magnetic field in accordance with V.R.M.S. method:
- DSV, for spheric volume with diameter of 10 cm: not bigger then 0,01 ppm
DSV, for spheric volume with diameter of 50 cm: not bigger then 4 ppm
| 1.4. | Magnet length including covers: no longer than 210 cm
| 1.5. | Field Of View (FOV)– not less then 50 cm in each direction X,Y,Z
| 1.6. | Zero boíl off technology
| 1.7. | Noise reduction in tunnel
| 2 | GRADIENT SYSTEM
| 2.1. | Gradient performance (Amax x SRmax)- minimum 7.000 mT²/m²/ms per axis, where Amax is the maximum gradient amplitude for each axes and SRmax is the highest value of slew rate for each axes)
| 2.2. | Maximum power of gradient amplifier, min. 1.6MW
| 3 | RF SYSTEM
| 3.1. | Number of independent receiver channels that can be used simultaneously in one single scan and in one single FOV without patient table repositioning, each generating an independent partial image, minimum 32
| 3.2. | Number of physical coil elements that can be simultaneously connected during one exam, minimum 122
| 3.3. | The resolution of the signal receiver is not less than 32 bits
| 3.4. | RF transmitter power not less than 30 kW
| 4 | RF COILS with requested number of coil elements- number of requested coil elements must correspond to the number of physical elements of particular requested coil (the sum of elements of multiple coils/combination of coils is not acceptable-except for coil under point 4.2.)
| 4.1. | Latest generation of integrated coil technology (GEM, Tim 4G, dStream, etc.) which supports simultaneous acquisition from different coils
| 4.2. | Separate, dedicated anterior body coil (one or combination of two) for body scanning in range of min. of 60 cm (from neck to pelvic) and with minimum 30 coil elements in total (one or combination of two).
| 4.3. | Separate, tiltable, dedicated multi-channel coil for the imaging of head and neck with minimum 19 coil elements
| 4.4. | Spine coil with minimum 24 coil elements
| 4.5. | Separate, multi-channel flex coils (two of them, different sizes) with minimum 4 elements each for the imaging of smaller joints
| 4.6. | Separate, dedicated multi-channel coil for the imaging of knee, with minimum of 16 coil elements, flex coil is not acceptable
| 4.7. | Separate, dedicated multi-channel coil for the imaging of shoulder, with minimum of 16 coil elements, flex coil is not acceptable
| 4.8. | Separate, dedicated multi-channel coil for the high resolution imaging and spetroscopy examinations of breast, with minimum of 16 coil elements
| 4.9 | Separate, dedicated multi-channel coil only for breast biopsy procedures, with minimum of 4 coil elements
| 5 | PATIENT TABLE
| 5.1. | Maximal scan range- not less than 181cm
| 5.2. | In bore ventilation and lightening
| 5.3. | Existence of volume control either in room or on the system console and bidirectional voice communication
| 5.4. | Patient video monitoring system
| 5.5. | Max. Patient weight for vertical and horizontal table movement, not less then 200 kg
| 5.6. | Coil storage cart or shelf
| 6 | ACQUISITION WORKPLACE (OPERATOR CONSOLE)
| 6.1. | DICOM standard for communication between devices with following functionalities:
DICOM STORAGE, DICOM QUERY/RETRIVE, DICOM PRINT, MODALITY WORKLIST
| 6.2. | Two (2) Color LCD monitors with resolution not less than 1.3MP with diagonal size not less than 19“ or one (1) color LCD monitor with diagonal size not less than 24“ with 1920x1200 resolution
| 6.3. | CD or DVD for DICOM archiving
| 6.4. | Acquisition packages for basic Neurology, Angiography, Cardiology, Radiology, Oncology, Musculoskeletal and Pediatric examinations
| 6.5. | Basic 3D post-processing techniques – VRT, Multi Planar Reconstruction (MPR), Maximum Intensity Projection
| 6.6. | Integration (composing) of several diagnostic images in 1 “Full Format” (longer anatomical structures and blood vessels)
| 7 | SEQUENCES AND IMAGING TECHNIQUES
| 7.1. | Turbo Dark Fluid (FLAIR)
| 7.2. | Susceptibility-Weighted Imaging
| 7.3. | ASL (2D or 3D)
| 7.4. | Magnetization Transfer Contrast
| 7.5. | Diffusion Tensor Imaging
| 7.6. | Neuro perfusion
| 7.7. | Non-invasive blood flow quantification
| 7.8. | High resolution diffusion- RESOLVE, FOCUS or equivalent
| 7.9. | Simultaneous Multi-Slice/ Multi-Band imaging (HyperBand, MultiBand SENSE, SMS or equivalent)
| 7.10. | Respiratory (breathing) and patient movement artifact elimination- PROPELLER or BLADE or MultiVane or JET or equivalent)
| 7.11. | 2D and 3D MRCP
| 7.12. | Dynamic Liver Imaging (THRIVE, VIBE, LAVA-Flex or equivalent
| 7.13. | Ultrafast Dynamic Contrast-enhanced Imaging (TWIST Vibe, DISCO or equivalent)
| 7.14. | 3D Contrast angiography in real time (TRICKS or TWIST or 4D-TRACK or Freeze Frame or equivalent)
| 7.15. | Angiography with synchronized bolus tracking (SmartPrep or BolusTrak or Care Bolus or equivalent)
| 7.16. | Non-contrast angiography (for peripheral blood vessels, abdomen and neuro exams- NATIVE or Inhance or TRANCE or Time-SLIP or equivalent)
| 7.17. | Fat and water separation technique (using DIXON or similar technique)
| 7.18. | Technique to reduce artifacts caused by metal implants- (O-MAR, Advanced WARP, MAVRIC or equivalent)
| 7.19. | Single voxel spectroscopy, Chemical Shift Imaging (2D and 3D)
| 7.20. | Software for guiding breast biopsy procedures
| 7.21. | Cartilage evaluation of joints- T2 mapping (T2 Map(CartiGram), MapIt or equivalent)
| 8 | ADDITIONAL EQUIPMENT AND SERVICES
| 8.1. | RF cabin for electromagnetic protection (damping value of min 90dB). RF cabin contains complete cabin with antistatic floor, door, window, antimagnetic lightening, all necessary filters, professional airconditioning also with connections for gases supply of the anesthesia system, and all final works, for running requirements of MR scanner for work with the patients
| 8.2. | Chiller for MR gradient system cooling
| 8.3. | Electrical distribution cabinet for MR system
| 8.4. | Injector for MR studies compatible with 1.5T and 3T MR systems-2 pcs.
| 8.5. | Antimagnetic stretcher for patient transport
| 8.6. | 10 days on site training
| 8.7. | Training for 4 physicians, technicians or physicist for min 5 days in Hospital with same or similar system

## CT
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
2 | 2 | Multislices CT scanner for diagnostics and RT procedures with CT Simulation Laser System |  | 1 |  | 0
| 1 | Gantry
| 1.1 | Gantry aperture – min 78 cm
| 1.2 | Scan Field of View (FOV) min 50cm
| 1.3 | /
| 1.4 | The shortest rotation time for one single full rotation of scanning (3600) max. 0.3 sec
| 2 | Patient table
| 2.1 | Motorized vertical table movement
| 2.2 | Max. table load with guaranteed positioning accuracy, min. 200kg
| 2.3 | The lowest position of the table max 55 cm
| 2.4 | Horizontal scan range is up to 160 cm
| 2.5 | Accessories: Mattress protector, head-arm support, non-tiltable head holders with positioning cushion set, patient restraining system for head fixation, restraining-strap set with body fixation strap that can be directly connected to the patient table top, headrest, knee-leg support
| 3 | X-ray generator
| 3.1 | Max. Generator power min. 80 kW
| 3.2 | Maximal tube curent min. 600 mA.
| 3.3 | Tube voltage range min. 80 – 130kV
| 4 | X-ray tube
| 4.1 | Heat storage capacity min. 6.5MHU or min. equivalent 50 MHU comparing to the performance of a conventional tube
| 4.2 | Anode cooling rate, min. 1.3 MHU/min
| 4.3 | Number of tube focal spots, min. two, please specify the size
| 4.4 | Possibility of the automated real-time tube current adjustment for optimal image quality at lowest possible dose, depending on patient size and anatomy
| 4.5 | Indexing plate, same indexed as existing plate on linac
| 5 | Detector system
| 5.1 | Detectors should be the latest technology (Gemstone Clarity, Stellar, Matrix Array, Nano Panel, PureVision or equivalent)
| 5.2 | Min. 128 acquisition slices, independent on number of detector rows or producer’s technology
| 5.3 | Min. 256 reconstructed slices, independent on number of detector rows or producer’s technology
| 5.4 | Free selection of the pitch without reduced image quality, please specify the pitch range
| 6 | Acquisition parameters
| 6.1 | Scan Field of View (FOV) min 50cm
| 6.2 | Spatial resolution in all three directions x/y/z for 2% MTF min. 14,5 lp/cm
| 7 | Operator’s console:
| 7.1 | 3D Real Time MPR (Multiplanar reconstruction in real time of the different slices thickness;
| 7.2 | 3D SSD (shaded surface display)
| 7.3 | CT – angiography (evaluation of the spiral images and display of the vessels, vascular anomality)
| 7.4 | Automatically Start of spiral triggered via contrast medium enhancement
| 7.5 | 4D workflow with direct generation of axial, sagittal, coronal, or double-oblique images from standard scanning protocols or workflow with direct generation of axial, sagittal or coronal from standard scanning protocols.
| 7.6 | Tool for contrast enhancement without increasing in noise or dose for better gray-white matter differentiation.
| 7.7 | Distance and angle measurement, image annotation and labeling
| 7.8 | Iterative Reconstruction for significant dose reduction and image quality improvement (iDose, AIDR 3D, ASiR, SAFIRE or equivalent.)
| 7.9 | Metal artefact reduction algorithm, for redustion artefact coused by spine implants, pacemakers, dental filings, neuro coils etc.
| 7.1 | Software for patient dose reduction
| 7.11 | High resolution flat screen monitor min.19“, resolution min. 1280 x 1024
| 7.12 | Software feature that enables user to obtain electron densities directly from the CT images, allowing patients to be scanned at any kV setting or Elector Density Phantom for precise correlation CT data to electron density of various tissues and software for analysis of the measurement achieved with the Phantom for CT numbers conversion to Electron density
| 7.13 | 4DCT/respiratory gating package for prospective or retrospective acquisitions, (Acquire, review/edit wave form, reconstruct, generate MIP, minIP)
| 7.14 | Hardware for respiratory gating
| 7.15 | CD/DVD
| 8 | Networking and images archiving
| 8.1 | DICOM storage (send/receive), DICOM Query/Retrieve, DICOM Basic Print, DICOM Get worklist (HIS/RIS), MPPS
| 8.2 | DICOM adapter for networking and connection with other modalities.
| 9 | Automatic CT injector-2pcs.
| 10 | External laser system
| 10.1 | CT Simulation Laser System with three red movable lasers for identifying the isocenter location: One Ceiling -mounted Sagittal Laser, and two side mounted Lasers
| 10.2 | Isocentre definition and data transfer to CT Simulation Laser system for direct patient marking
| 11 | 10 working days application training on-site

## DW
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
3 | 3 | Diagnostic workstation in client – server architecture, with server and 10 clients computer(workstations), with 2D, 3D basic and advanced visualization applications from different modalities MR, CT, X-ray system, ultrasound |  | 1 |  | 0
| 1.1 | 2D and 3D viewing of images from several modalities (CT, MR, PET/CT, CR, X-ray images);
| 1.2 | Browser functionality for patient and data access
| 1.3 | Mirror, rotate, invert, windowing, pan/zoom, annotations, distance and angle measurement, pixel lens, ROI / VOI evaluation;
| 1.4 | Time-curve evaluation
| 1.5 | VRT, MIP thin/thick, MPR thin / thick, interactive slice thickness change
| 1.6 | 3D SSD technique;
| 1.7 | Bone removal for fast segmentation and removal of bony structures;
| 1.8 | Parallel, curved & radial ranges;
| 1.9 | 2D & 3D reference lines, 3D reference point;
| 1.1 | Region growing and quantification for interactive segmentation of anatomical structures;
| 1.11 | Automatic spine labeling;
| 1.12 | Automatic rib labeling for CT thorax scans;
| 1.13 | Automatic landmark registration for accurate anatomical alignment of multiple timepoint cases;
| 1.14 | Automatically and standardized reconstruction, with automatically range and projection initialization, depending on anatomy and efficient and standardized findings documentation by reports.
| 1.15 | Cinematic VRT for photorealistic 3D images of anatomical details, obtained with photon simulation.
| 1.16 | Oncology application for automated segmentation and evaluation of lesions in lung, liver, lymph nodes and other organs, with simultaneous two time point visualization and comparison, follow up of multiple time-points, 3D evaluation of lesion measurements.
| 1.17 | Software for virtual colonoscopy, for non-invasive, diagnostic evaluation of lesions in the colon.
| 1.18 | Computer-aided detection tool for automated detection of colon polyps.
| 1.19 | Computer-aided detection tool for automated detection of solid pulmonary nodules, partial-solid, and Ground-Glass Nodules (GGN) during review of CT examinations of the chest.
| 1.2 | Vascular analysis for automatically evaluation and quantification of the angiography images of the general vessels.
| 1.21 | Automatically table and bone removal and automatic generation of VRT/MIP Radial Ranges.
| 1.22 | Vessel analysis tools for obtaining of all relevant information, e.g. stenosis diameter and area, curved length, profile curve, minimum lumen identification.
| 1.23 | Center line creation for a quick and robust vessel segmentation and CPR display.
| 1.24 | Dedicated tool for bone assessment in a way that all ribs and the spine should be labeled and displayed spread out in one plane, for time efficient reporting of potential lesion locations.
| 1.25 | Software application for liver segmentation and semi-automated segmentation of arterial, portal venous, and venous vascular tree along with visualization of vascular supply of the hepatic parenchyma for planning of surgical interventions, functionality for semi-automated segmentation of liver lesions, and calculation of resulting partial liver volumes
| 1.26 | CT Neuro Perfusion for visualization of the blood perfusion in the brain
| 1.27 | CT Body Perfusion for quantitative evaluation of dynamic CT data of organs and tumors, following the injection of contrast media
| 1.28 | Dedicated Radiation Therapy (RT) software that is designed to make multi-modality (CT, MRI, PET, Cone Beam CT) simulation, image assesment and contouring
| 1.29 | MR Radiology Workflows: predefined layouts for Head, C-Spine, T-Spine, L-Spine, Whole Spine, Breast, Prostate, Abdomen, Hip and Knee scans, Composed Images
| 1.3 | MR Evaluation: Subtraction, MeanCurve, Image Filter, 2D/3D Distortion Correction.
| 1.31 | MR Spectroscopy evaluation (Single voxel spectroscopy, Chemical Shift imaging)
| 1.32 | MR brain perfusion evaluation
| 1.33 | MR 3D lesion segmentation (semi-automated volumetric evaluation of tumors, lymph nodes and metastases)
| 1.34 | MR application for quantitative evaluation of functional response to treatment
| 1.35 | Advanced dedicated MR breast studies evaluation (Real-time display and analysis of kinetic parameters, Graphical volume statistics of lesion enhancement, Calculation of lesion volumes
| 1.36 | Advanced MR application which create intuitive parameter maps for efficient reading of dynamic series of prostate imaging
| 1.37 | Advanced MR application for diffusion tensor imaging (DTI) evaluation and Tractography
| 1.38 | Advanced MR application for calculation of the Perfusion-Diffusion difference for brain imaging
| 1.39 | LCD high resolution monitor min. 2MP and diagonal size min. 24“, 2 pcs. for each client, 20 pcs in total plus one monitor min. 19” for administration for each client, 10 pcs in total
| 1.4 | CD / DVD Robot, Automatic Burning System, 2 pcs Protocol: DICOM store, HTTP WEB server Speed: min 30 CD per hour and min 15 DVD per hour (simultaneous studies reception, burning on both recorders) Optical recorder: Two CD / DVD optical recorders DICOM connection: "DICOM-Send" from PACS and all other modalities DICOM Qeury / Retrieve from PACS from one or more PACS
| 1.41 | 10 days on site application training
| 1.42 | 10 days training for technical engineer in training center of producer
| 1.43 | Server hardware and clients hardware which support following functionalities (All applications listed above are installed on server and are accessible on all 10 clients)

## Mammo
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
4 | 4 | 3D Digital Mammography System with Thomosynthesis, Mammography Contrast Software (CE2D), with addition for stereotactic and tomosynthesis-guided breast biopsy |  | 1 |  | 0
| 1 | GENERATOR
| 1.1 | High frequency generator of a constant potential of minimum power of 7kW
| 1.2 | Output voltage in the minimum range of 25 - 45 kV or wider in 1 kV steps
| 1.3 | mAs range: 4 to 480 mAs or wider
| 2 | X-RAY TUBE
| 2.1 | Minimum two focuses (for standard imaging and magnification), Focal spot size: 0,1/ 0,3 mm.
| 2.2 | Anode heat storage capacity: min. 300.000 HU.
| 2.3 | Anode material: Tungsten or Molybdenum (Mo) Rhodium (Rh)
| 2.4 | Filter: Rhodium (Rh) and Aluminum (Al) ) or Titanium (Ti).
| 2.5 | Anode rotation speed min.9000r.p.m.
| 2.6 | Automatic collimation according to the type and size of the compression paddle
| 2.7 | Automatic selection of filtration depending on the percentage of glandular tissue of the breast
| 3 | GANTRY AND C-ARM
| 3.1 | Motorized C-arm rotation with numerical angle indication
| 3.2 | SID: min. 65 cm.
| 3.3 | Motorized C-arm rotation: min 350°
| 3.4 | Automatic and manual compresion.
| 3.5 | Scattered radiation grid, linear
| 3.6 | Double foot control for vertical movement and compression.
| 3.7 | Patient face protector.
| 3.8 | The vertical movement and rotation control buttons shall be accessible on both sides both on the C-arm
| 3.9 | Possibility of selecting automatic or manual decompression after exposure
| 3.1 | Magnification factor for diagnostic procedures min. 1.5x
| 3.11 | Compression paddles:
| 3.12 | for small and large breast
| 3.13 | compression paddle for spot magnification
| 3.14 | for a total magnification
| 3.15 | for axilla radiography
| 3.16 | Compression paddle for small breast has automatic motorized shift in the X axis for correct collimation in MLO projections
| 3.17 | Compression paddle with the ability to adjust to the shape of the breast, in order to equalize compression by volume for all breast sizes
| 3.18 | Automatic exposure control (AEC) depending on the type of breast tissue
| 3.19 | Wall mounted compression plates holder
| 4 | FLAT PANEL DETECTOR
| 4.1 | Direct-conversion method.
| 4.2 | The pixel size max. 70µm.
| 4.3 | Efective image area min. 23x29cm.
| 4.4 | Depth of image gray scale at least 14 bits
| 5 | TOMOSYNTHESIS
| 5.1 | Allowing projection images from multiple angles to be acquired.
| 5.2 | Scan angle min. 15°.
| 5.3 | Number of acquired images: min. 15
| 5.4 | Synthesized 2D image that is composed of tomosynthesis images (synthesized several sliced images).
| 5.5 | Possibility of 2D and 3D images under the same compression
| 5.6 | Thickness of a single layer of tomosynthesis maximum 1 mm
| 5.7 | Tomosynthesis biopsy software providing functions to calculate the necessary coordinates for performing a biopsy study using slice images acquired from a tomosynthesis exposure.
| 5.8 | Fixed protection for the patient's face during tomosynthesis.
| 5.9 | Speed of Tomosynthesis scan less than 5 seconds
| 6 | SYSTEM FOR STEREOTACTIC AND TOMOSINTHESIS GUIDED BIOPSY AND EXCISION OF BREAST LESIONS
| 6.1 | Biopsy Positioner with needle holder, guide and lateral approach kit. Providing functions to calculate the necessary coordinates of the sample for performing a biopsy study and to move the needle to the calculated position accurately.
| 6.2 | The capability to perform wire localization procedures on breast lesions
| 6.3 | The capability to perform stereotactic biopsy and excision procedures (2D) of the breast lesions
| 6.4 | The capability to perform tomosynthesis guided biopsy and excision procedure (3D) of the breast lesions compatible with ATEC Hologic vacuum biopsy system
| 6.5 | Standard compression paddle for biopsy
| 6.6 | Special compression biopsy paddle
| 6.7 | Special compression paddle with a wide opening for biopsy or compression paddle for tomobiopsy
| 6.8 | System for biopsy QC
| 7 | ACQUISITION WORKSTATION
| 7.1 | Hardware features validated by the manufacturer.
| 7.2 | Dedicated height adjustable operation console (acquisition workstation) with protective operator schield min. 0,5 mm Pb, integrated in the housing of the acquisition station.
| 7.3 | Diagnostic medical monitors min. 3MP for image preview.
| 7.4 | Acquisition Workstation allows for entering patient information, selecting various study conditions and exposure dose, executing image processing.
| 7.5 | Display of a tomosynthesis image (scan, reconstruction, and display on AWS
| 7.6 | Enter data manually or via DICOM worklist
| 7.7 | DICOM Compatibility (Send / Receive, Query / Retrieve, Storage Commitment, Print, Worklist, MPPS, Radiation Dose SR)
| 7.8 | Retake Analysis- Cause analysis of reexposure executed by tehnician
| 7.9 | The previous image as well as all relevant information about the image (kV, mAs, radiation dose, compression thickness, compression force, etc.) are available after each exposure
| 7.1 | Opportunity to edit projections in the protocol
| 7.11 | Special processing algorithm for implant imaging
| 7.12 | Possibility of individual work under the user's credentials (name and password)
| 7.13 | UPS for the acquisition workstation
| 8 | DIAGNOSTIC WORKSTATION
| 8.1 | Workstation with dedicated software for 2D and tomosynthesis images.
| 8.2 | Workstation with diagnostic monitor features validated by the manufacturer.
| 8.3 | Two diagnostic medical TFT monitors, min. 21 inch, resolution min. 5MP with technology which enable to render 15 mega sub-pixel image, max. luminance min. 1000 cd/m² dicom calibrated, contarst ration min. 1:1400 or one dual screen 12MP diagnostic monitor, max. luminance 2100, min. 1000 cd/m² dicom calibrated.
| 8.4 | One LCD monitor for patient administration, min. 21 inch.
| 8.5 | UPS - Uninterruptible Power Supply.
| 8.6 | Special, dedicated keyboard with special buttons for mammography.
| 8.7 | Database search by multiple criteria, patient name, patient ID, type of modality, study date, key word etc.
| 8.8 | Image processing tools: contrast, brightnes, inversion, zoom, rotation, annotations, meaurements.
| 8.9 | User-configurable "hanging protocols" for the mammography.
| 8.1 | Divide mammograms according to breast sizes and display all regions of breasts in pixel-to-pixel zoom
| 8.11 | Support for other modalities: US, CT, MRI
| 8.12 | Computer Assisted Detection (CAD) software for the detection of malignant tissue densities and calcifications.
| 9 | FUNCTIONAL CONTRAST ENHANCED 2D BREAST IMAGING
| 9.1 | Software for CE2D (Contrast Enhanced 2D)
| 9.2 | The software should eliminate structural noise from the mammogram and show a clear contrast dynamics
| 9.3 | Software such that CE2D can be combined with 3D images to obtain a co-registered image
| 9.4 | Possibility to work with contrast on any compression paddle
5 | 5 | 3D Digital Mammography System with Thomosynthesis and addition for stereotactic and tomosynthesis-guided breast biopsy |  | 1 |  | 0
| 1 | GENERATOR
| 1.1 | High frequency generator of a constant potential of minimum power of 7kW
| 1.2 | Output voltage in the minimum range of 25 - 45 kV or wider in 1 kV steps
| 1.3 | mAs range: 4 to 480 mAs or wider
| 2 | X-RAY TUBE
| 2.1 | Minimum two focuses (for standard imaging and magnification), Focal spot size: 0,1/ 0,3 mm.
| 2.2 | Anode heat storage capacity: min. 300.000 HU.
| 2.3 | Anode material: Tungsten or Molybdenum (Mo) Rhodium (Rh)
| 2.4 | Filter: Rhodium (Rh) and Aluminum (Al) ) or Titanium (Ti).
| 2.5 | Anode rotation speed min.9000r.p.m.
| 2.6 | Automatic collimation according to the type and size of the compression paddle
| 2.7 | Automatic selection of filtration depending on the percentage of glandular tissue of the breast
| 3 | GANTRY AND C-ARM
| 3.1 | Motorized C-arm rotation with numerical angle indication
| 3.2 | SID: min. 65 cm.
| 3.3 | Motorized C-arm rotation: min 350°
| 3.4 | Automatic and manual compresion.
| 3.5 | Scattered radiation grid, linear
| 3.6 | Double foot control for vertical movement and compression.
| 3.7 | Patient face protector.
| 3.8 | The vertical movement and rotation control buttons shall be accessible on both sides both on the C-arm
| 3.9 | Possibility of selecting automatic or manual decompression after exposure
| 3.1 | Magnification factor for diagnostic procedures min. 1.5x
| 3.11 | Compression paddles:
| 3.12 | for small and large breast
| 3.13 | compression paddle for spot magnification
| 3.14 | for a total magnification
| 3.15 | for axilla radiography
| 3.16 | Compression paddle for small breast has automatic motorized shift in the X axis for correct collimation in MLO projections
| 3.17 | Compression paddle with the ability to adjust to the shape of the breast, in order to equalize compression by volume for all breast sizes
| 3.18 | Automatic exposure control (AEC) depending on the type of breast tissue
| 3.19 | Wall mounted compression plates holder
| 4 | FLAT PANEL DETECTOR
| 4.1 | Direct-conversion method.
| 4.2 | The pixel size max. 70µm.
| 4.3 | Efective image area min. 23x29cm.
| 4.4 | Depth of image gray scale at least 14 bits
| 5 | TOMOSYNTHESIS
| 5.1 | Allowing projection images from multiple angles to be acquired.
| 5.2 | Scan angle min. 15°.
| 5.3 | Number of acquired images: min. 15
| 5.4 | Synthesized 2D image that is composed of tomosynthesis images (synthesized several sliced images).
| 5.5 | Possibility of 2D and 3D images under the same compression
| 5.6 | Thickness of a single layer of tomosynthesis maximum 1 mm
| 5.7 | Tomosynthesis biopsy software providing functions to calculate the necessary coordinates for performing a biopsy study using slice images acquired from a tomosynthesis exposure.
| 5.8 | Fixed protection for the patient's face during tomosynthesis.
| 5.9 | Speed of Tomosynthesis scan less than 5 seconds
| 6 | SYSTEM FOR STEREOTACTIC AND TOMOSINTHESIS GUIDED BIOPSY AND EXCISION OF BREAST LESIONS
| 6.1 | Biopsy Positioner with needle holder, guide and lateral approach kit. Providing functions to calculate the necessary coordinates of the sample for performing a biopsy study and to move the needle to the calculated position accurately.
| 6.2 | The capability to perform wire localization procedures on breast lesions
| 6.3 | The capability to perform stereotactic biopsy and excision procedures (2D) of the breast lesions
| 6.4 | The capability to perform tomosynthesis guided biopsy and excision procedure (3D) of the breast lesions compatible with ATEC Hologic vacuum biopsy system
| 6.5 | Standard compression paddle for biopsy
| 6.6 | Special compression biopsy paddle
| 6.7 | Special compression paddle with a wide opening for biopsy or compression paddle for tomobiopsy
| 6.8 | System for biopsy QC
| 7 | ACQUISITION WORKSTATION
| 7.1 | Hardware features validated by the manufacturer.
| 7.2 | Dedicated height adjustable operation console (acquisition workstation) with protective operator schield min. 0,5 mm Pb, integrated in the housing of the acquisition station.
| 7.3 | Diagnostic medical monitors min. 3MP for image preview.
| 7.4 | Acquisition Workstation allows for entering patient information, selecting various study conditions and exposure dose, executing image processing.
| 7.5 | Display of a tomosynthesis image (scan, reconstruction, and display on AWS
| 7.6 | Enter data manually or via DICOM worklist
| 7.7 | DICOM Compatibility (Send / Receive, Query / Retrieve, Storage Commitment, Print, Worklist, MPPS, Radiation Dose SR)
| 7.8 | Retake Analysis- Cause analysis of reexposure executed by tehnician
| 7.9 | The previous image as well as all relevant information about the image (kV, mAs, radiation dose, compression thickness, compression force, etc.) are available after each exposure
| 7.1 | Opportunity to edit projections in the protocol
| 7.11 | Special processing algorithm for implant imaging
| 7.12 | Possibility of individual work under the user's credentials (name and password)
| 7.13 | UPS for the acquisition workstation
| 8 | DIAGNOSTIC WORKSTATION
| 8.1 | Workstation with dedicated software for 2D and tomosynthesis images.
| 8.2 | Workstation with diagnostic monitor features validated by the manufacturer.
| 8.3 | Two diagnostic medical TFT monitors, min. 21 inch, resolution min. 5MP with technology which enable to render 15 mega sub-pixel image, max. luminance min. 1000 cd/m² dicom calibrated, contarst ration min. 1:1400 or one dual screen 12MP diagnostic monitor, max. luminance 2100, min. 1000 cd/m² dicom calibrated.
| 8.4 | One LCD monitor for patient administration, min. 21 inch.
| 8.5 | UPS - Uninterruptible Power Supply.
| 8.6 | Special, dedicated keyboard with special buttons for mammography.
| 8.7 | Database search by multiple criteria, patient name, patient ID, type of modality, study date, key word etc.
| 8.8 | Image processing tools: contrast, brightnes, inversion, zoom, rotation, annotations, meaurements.
| 8.9 | User-configurable "hanging protocols" for the mammography.
| 8.1 | Divide mammograms according to breast sizes and display all regions of breasts in pixel-to-pixel zoom
| 8.11 | Support for other modalities: US, CT, MRI
| 8.12 | Computer Assisted Detection (CAD) software for the detection of malignant tissue densities and calcifications.
| 8.13 | Mammography examination chair with following funcionalities:Height adjustable head rest, adjustable back support, drop down shoulder segements and arm rests, safety switch to lock height adjustment, mothorized height adjustment

## Ultrasound
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
6 | 6 | Ultrasound system 1 |  | 2 |  | 0
| 1.1 | Maximum imaging depth, min. 40cm
| 1.2 | Integrated high resolution flet panel with LCD/LED/OLED technology monitor, screen size of min. 22“, resolution min. 1920x1080
| 1.3 | Integrated Touch screen diagonal size min. 13“( +/-1’’) on control panel
| 1.4 | Height adjustable control panel and swivel
| 1.5 | Min. 4 active transduser ports plus additional port for pencil probe
| 1.6 | Easy manevaurable system on wheels. System weight max 130kg
| 1.7 | Imaging modes: B (2D) mod, Tissue Harmonic Imaging, PW doppler, color-doppler, M mode, anatomical M mode
| 1.8 | Automatic optimization thru overall field of view (FOV) of image brightness uniformity by changing the depth gain compensation and overall gain (in 2D-mode and Pulsed Wave Doppler)
| 1.9 | Trapezoid imaging
| 1.1 | Advanced speckle reduction technology
| 1.11 | Panoramic imaging
| 1.12 | Technology for fully automated detection of lesions and anatomical structures in 2D mode (based on automatic border detection technology or equivalent) and automatic measurments of area and diameter
| 1.13 | Special technology for slow flow quantification, B-flow, Clarify or equivalent
| 1.14 | Shear wave elastography for measuring tissue shear velocity (Vs) and elasticity (E)for a selected region of interest using acoustic radiation force impulse (ARFI) technology. Option should provide measurement labels for lesions and liversegments. Available on offered convex and at least one offered linear probe
| 1.15 | Strain elastography which provides a qualitative representation of relative tissue stiffness for the region of interest. Available on both offered linear probes
| 1.16 | Image archiving in DICOM and PC file formats (AVI and JPEG or similar) on CD/DVD and USB
| 1.17 | USB port, min. 2
| 1.18 | 2D covex multifrequency probe with frequency range of min 1-5 MHz with min FoV of 70 degrees and minimum 160 elements
| 1.19 | 2D linear multifrequency probe with frequency range of min 3-9 MHz with minimum 320 elements
| 1.2 | 2D linear multifrequency probe with frequency range of min 5-17 +/-3MHz with min FoV of 50mm and min 576 elements
| 1.21 | B/W digital video printer
| 1.22 | Integrated gel warmer on the system
| 2 | Table for US examination
Table for ultrasound examination with following functionalities: width: min 65 cm, length: min 195 cm, electric height adjustment, head section adjustment, paper roll holder
7 | 7 | Ultrasound system 2 |  | 1 |  | 0
| 1.1 | Maximum scanning depth is at least 40 cm.
| 1.2 | Minimum four active probe connectors for electronic scanning
| 1.3 | Integrated high resolution flet panel with LCD/LED/OLED technology monitor, screen size of min. 22“, resolution min. 1920x1080
| 1.4 | Control console with a minimum 10" TFT LCD touch panel, height adjustment and swivel
| 1.5 | It is possible to assign measuring functions to the alphabet keys on the keyboard or touch screen
| 1.6 | Easy manevaurable system on wheels.
| 1.7 | Frequency range of the system is 1-22 MHz or wider.
| 1.8 | Hard disk capacity is minimum 1TB.
| 1.9 | DICOM: Storage, MWM and Print.
| 1.1 | Min. 2 Digital video input / output
| 1.11 | USB connectors: min. 2.
| 1.12 | Wired and wireless connection to hospital network.
| 1.13 | At least 12,000 images in the cineloop sequence
| 1.14 | Integrated gel warmer.
| 1.15 | B/W digital thermal printer.
| 1.16 | OPERATING MODES
| 1.17 | B, M, Color Doppler, Power Doppler, Triplex mode, Pulse Wave Doppler
| 1.18 | Advanced Color Doppler technique for better visualization of micro vascularity and slow flows, with information about flow direction, with greater spatial resolution.
| 1.19 | B-flow, DFI (Detective Flow Imaging), or an equivalent technique for visualization of small blood vessels that cannot be visualized by Doppler techniques.
| 1.2 | Shear Wave elastography available on the offered convex probe and compression elastography available on all offered probes. Possibility of combined display of compression elastography and share wave elastography, as well as evaluation of the stage of liver fibrosis by both methods and obtaining a joint report.
| 1.21 | Synchronized display of volume data sets imported from MRI, CT or PETCT simultaneously in MPR reconstruction (on split screen) with real-time ultrasound. Following functions are supported: Combined display with compression elastography; Is possible to import a minimum 3 different data sets for the same patient and change them during the examination; Displays the puncture guide line on virtual images, with a 3-axis display; Body Motion Tracking; A color map superimposed on the CT image simulates the distribution of electric current from the given location of the multiple electrodes; Tracks and displays the needle tip location in real time during RFA procedures.
| 1.22 | A program for enhanced visualization of the needle during biopsy.
| 1.23 | Program for contrast agents.
| 1.24 | Automatic angle correction in PW Doppler.
| 1.25 | Spatial compound imaging with the possibility of adjusting the angle of wave propagation.
| 1.26 | Automatic Doppler spectrum outline in real time with measurement display.
| 1.27 | Automatic image optimization- Optimization of gain and sound velocity in the B mode and gain, velocity range and baseline shift in PW mode.
| 1.28 | Trapezoidal image display when using linear probe.
| 1.29 | Technique for generating the optimal image by changing the filter depending on the input signal.
| 1.3 | Focus through the entire depth of the image without need to set individual focal points.
| 1.31 | The ability to store raw data (Raw Data) on external media and internal hard disk, and to process basic parameters.
| 1.32 | Linear probe with wide frequency range min.8-22 MHz.+/-2MHz
| 1.33 | Linear probe for soft tissue, frequency range at least from 5-18 MHz, wide max. 40mm.
| 1.34 | Linear probe for soft tissues, frequency range 5-12 MHz+/-2MHz or wider, wide min. 50mm
| 1.35 | Abdominal Convex probe frequency range min. 1-5 MHz+/-1MHz
| 2 | Table for US examination
Table for ultrasound examination with following functionalities: width: min 65 cm, length: min 195 cm, electric height adjustment, head section adjustment, paper roll holder
8 | 8 | Ultrasound system 3 |  | 1
| 1.1 | Maximum imaging depth, min. 40cm
| 1.2 | The system shall have minimum 4 active pinless probe conectors
| 1.3 | Beamformer Technology: The system shall utilize a last-generation beamformer technology which supports high resolution and contrast
| 1.4 | Integrated high resolution flet panel with LCD/LED/OLED technology monitor, screen size of min. 22“, resolution min. 1920x1080
| 1.5 | Integrated Touch screen diagonal size min. 13“( +/-1’’) on control panel
| 1.6 | The system shall have full-size QWERTY keyboard with backlight
| 1.7 | Control panel with height adjustment and swivel
| 1.8 | Integrated Gel Warmer
| 1.9 | B/W thermal printer
| 1.1 | Easy manevaurable system on wheels
| 1.11 | Min 4 active transducer ports (pinless connectors)
| 1.12 | Scan modes and imaging features: B-mode Imaging, Color Doppler, PW Doppler, Hypersensitive Color Doppler, Power Doppler, Triplex Mode, ElastoScan, Contrast Imaging, Fusion Imaging
| 1.13 | Frequency Range: The system shall have frequency range of 1 – 15 MHz
| 1.14 | /
| 1.15 | CINE Function: The system shall perform CINE Function, 12.000 cine images minimum
| 1.16 | The system shall include at least 500GB of dedicated hard drive.
| 1.17 | DVD Drive
| 1.18 | Exported Image formats: The system shall export images and clips in PC compatible JPG format, AVI format and DICOM format
| 1.19 | USB connectors: min. 2.
| 1.2 | DICOM Connectivity and DICOM Services
| 1.21 | Transducers:
| 1.22 | Convex Probe: Frequency range: 1 – 5 MHz+/-1MHz or wider, number of elements min. 160, single crystal or matrix technology
| 1.23 | Linear Probe: Frequency range: min. 4 – 12 MHz+/-1MHz number of elements min.256, min., FOV max. 50mm
| 1.24 | Linear Transducer: Frequency range: min. 3 – 9 MHz+/-1MHz, number of elements min.160, min., FOV max. 50mm
| 1.25 | Advanced features:
| 1.26 | Real-time Elastography with quantification for Breast and Thyroid exams
| 1.27 | Shearwave elastography for Liver and Breast exams
| 1.28 | Contrast Imaging Studies available on the system
| 1.29 | The system provides projection of simulated needle path for accurate interventional procedures
| 1.3 | Automatic image optimization in B and color mode and the range of velocity and scale in PW Doppler by pressing one button.
| 1.31 | Available active Fusion imaging for realtime imaging in conjunction with other volumetric imaging modalities (CT, MRI)
| 1.32 | System supports visualization of slow flow microvascular structure for detailed examination of blood flow in relation to surrounding tissue
| 2 | Table for US examination
Table for ultrasound examination with following functionalities: width: min 65 cm, length: min 195 cm, electric height adjustment, head section adjustment, paper roll holder

## Conference system
Bidder:
LOT |  |  |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
9 | 9 | Conference system |  | 1 |  | 0
| 1 | Video and web conferencing services should allow the organization of multi-party meetings to be included in the meeting using a SIP or H.323 video conferencing device or through a web interface in order to organize consiliums, share video material and images from different modalities and diagnostic workstations Video and web conferencing services should also allow for the registration of video conferencing devices through the SIP protocol, with the possibility to define a unique SIP URI for the video conferencing device. It should always be possible to organize at least one meeting that can host at least 25 video conferencing devices and at least 200 participants accessing through the web interface
| 2 | Ability to start a meeting from a Microsoft Outlook application.
| 3 | The service should be provided as a cloud service without the need to implement any on-premises infrastructure
| 4 | Video and web conferencing services should provide the following functionality to video conferencing devices: o Support for video conferencing with a resolution of at least 720p and 30 frames per second and displaying content from a computer with a resolution of at least 720p and 5 frames per second o Video aspect ratio: 4: 3 and 16: 9 o The ability to have video conferencing devices that have one, two or more screens o Advanced Encryption Standard (AES) 128-bit encryption o Support for signaling protocols: SIP, SIP Transport Layer Security, H.323, H.323 with H.235, H.225.0 and H.245 o Support for media protocols: Real-Time Transfer Protocol, Secure RTP and Real-Time Control Protocol o The ability to place the video conferencing device mute or unmute during a meeting using the DTMF control
| 5 | Video and web conferencing services should provide the following functionality to users who access the meeting through the web interface: o The ability to show the other participants in the meeting an application, a document, or a computer screen o The ability to display the screen of your Android device o The ability to display the contents of an iOS device o Ability to display audio and video of the current speaker in the conference o Ability to define an alternative meeting host o Ability to send private and public text messages to other meeting participants o Organizing surveys during the meeting o Ability to start recording a meeting o Support for web access to meetings through: Chrome, Firefox, Safari, Internet Explorer and Egde browsers o Ability to define a password for users who access the meeting through the web interface
| 6 | Ability to organize a meeting and attend the meeting through mobile devices based on iOS and Android operating systems
| 7 | There should be a minimum of 10GB of memory in the cloud where the recorded meetings and files will be stored
| 8 | Video and web conferencing services should allow all participants to be shown during the meeting, there should be an option to mute or unmute the audio, to exclude the participant from the meeting and to close the meeting to prevent unplanned participants from attending
| 9 | There should be an admin portal that, through web access, should enable: adding new users, devices and services, monitoring and administering video conferencing devices, reporting, assigning different user access rights, and checking the status and performance of Video and Web conferencing services.
| 10 | It should be also enabled the use of a client application with the following functionalities: o The ability to make audio and video calls to video conferencing devices that use the standard SIP signaling protocol. o Option for whiteboard function within the application. o The ability to send text messages. o Communications with internal and external users through a single system. o The ability to schedule appointments through the app. o Application connectivity to Outlook to help you organize meetings. o The possibility of conference calls. o Ability to initiate video conferences for up to 25 participants in one group conversation o Ability to send files up to 2GB in size through the application. o Creating different teams of users within the application o Starting different group conversations within one or more different teams. o The ability to use the video conferencing device as the main screen. o The ability to use the video conferencing device as an auxiliary screen. o Ability to use a shared company directory. o Ability to run the application from different operating systems (Android, Windows and iOS) o Traffic encrypted with Advanced Encryption Standard (128-bit AES).
| 11 | A video conferencing device with the following features is also required: o Connect wirelessly to a client on a PC or phone. o 55-inch multi-touch 4K touch screen. o 4K camera built into the device. o Fixed camera with min 80 degree viewing angle (to get full room coverage). o Minimum of 12 high quality integrated microphones covering a room of 8 to 10 people. o The possibility of conference calls. o Ability to zoom in on one user from the entire conference. o Ability to store content in a secure cloud. o Ability to access files created from the client application and stored on the cloud directly from the device. o Ability to annotate files created from a client application using a video conferencing device. o Ability to use the device as a Whiteboard. o Possibility for two users to use whiteboard at the same time. o Share Whiteboard content to all conference participants. o Permanent storage of Whiteboard content as well as chat history. o Register user accounts directly in coud. o HDMI input. o Extended Display Indetification Data (EDID). o Multilingual interface (English, French, German or Spanish).
| 12 | 31.5" diagnostic monitor Pixel Pitch max 0.18159mm x 0.18159mm 8MP (3840x2160) IPS display with changeable resolution to 6 and 4MP Auto luminance calibration HDR10, DICOM complient Brightness 450cd/m2 Contrast ratio 1300:1 Response Time 5ms HDMI x 1/ Display Port x 2 2PBP, dual controller