## Auxiliray devices
Bidder: |  | Offered price:(EUR)
LOT 4 | Auxiliary surgical devices
ID | Equipment name | QTY | Technical Specification Offered(model) | Total Price per line item
4.1 | Electrosurgical generator
4,1,1 | Generator | 1
4,1,2 | Plasma Generator | 1
4.2 | Audiometer | 1
4.3 | Retromolar Intubation Endoscopes with Camera | 1
4.4 | Platform for invasive hemodynamic monitoring of the patients | 1
4.5 | Flexible intubation fiberscope | 1
4.6 | Intensive care unit Hemodynamic monitor | 1
4.7 | Nebulizer based on vibrating mesh | 10
4.8 | Device for Mechanical prevention of venous thromboembolism with active intermittent, sequential and graduated pneumatic compression of lower limbs | 3
4.9 | ODM+ - Oesophageal Doppler monitor | 1
4.1 | Multifunctional volumetric infusion pump with software for enteral nutrition | 5
4.11 | Laryngoscope sets for special purposes | 1
4.12 | Laryngoscopes with replaceable lights
4,12,1 | MACINTOSH Laryngoscope Blade- cold light with with replaceable fiber optic light carrier | 4
4,12,2 | MACINTOSH Laryngoscope Blade- warm light fiber optic light carrier incorporated | 4
4.13 | Video Laryngoscope for visual endotracheal intubation -set with three video spatula | 1
4.14 | Sets for therapeutic hypothermia / normothermia | 10
4.15 | Temporary single chamber exteranl pacemaker, 2 units and 4 patient cables | 2
4.16 | Instruments for abdominal surgery -set | 10
4.17 | Neonatal Resuscitation and Jaundice
4,17,1 | Resuscitation table for newborns | 1
4,17,2 | Jaundice meter | 1
4.18 | HOSE HOLDER for anesthetic systems
4,18,1 | Hose retainer plate | 10
4,18,2 | Hinged arm | 10
4.19 | Mobile bathroom | 3
4.2 | Coasters medical-grade silicone for operating tables -set (Equipment for Pressure care Management in the OR and ICU) | 5
Total Price per Lot

## 1.Generator
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
| 4.1 | Electrosurgical generator |  |  |  | 0
1 | 4,1,1 | Generator |  | 1 |  | 0
| 1 | Electrosurgical generator with a rotary peristaltic pump, delivers bipolar RF power with adjustable settings up to 200W. The generator senses the high resistance tissue and reduces the RF power output to a level which prevents arcing or cutting
| 2 | · Simultaneously delivers RF power and saline provides hemostatic sealing of soft tissue and bone at the surgical site; using the saline results in:
| 3 | · Conducting electrical energy from electrode to tissue, even when the electrode is moving
| 4 | · Preventing over heating of the tissue surface (approximately 100°C)
| 5 | · Allowing electrodes to move smoothly across targeted tissue without sticking, carbonization, smoke, char or odor
| 6 | · Enlarging the contact area and allows a painting motion to hemostatically seal bleeding tissue
| 7 | · Ability to quickly stop bleeding from large oozing surfaces
| 8 | · Ability to quickly spot coagulate spurting vessels
| 9 | · Ability to treat tissue deeply if desired
| 10 | · The three possible saline flow rates for each power setting are preset automatically (low, medium, high) to provide the optimal saline flow for a given power setting
| 11 | · Flow rate speed from up to 36mL/min depending on power setting and flow rate setting
| 12 | · The saline flow is calibrated to keep the device tips cooled to approximately 100°C as the power is moved up and down
| 13 | · Output waveform is up to 370kHz
| 14 | o Trolley for holding and transport of system components
2 | 4,1,2 | Plasma Generator |  | 1 |  | 0
| 1 | · Operational frequency up to 1MHz in all 3 cutting modes and 2 coagulation modes with the optional continuous and discontinuous work cycle of the outbound signal (cutting mode: 10 seconds ON, 30 seconds OFF / coagulation mode: 10 seconds ON, 30 seconds OFF)
| 2 | · 3 cutting modes with automatic power adjustments for each mode
| 3 | o Low cutting mode, with a maximum power of 160W, enables precise cutting with some hemostasis and minimal collateral damage
| 4 | o Medium cutting mode, with a maximum power of 60W, enables precise cutting with elevated hemostasis
| 5 | o High cutting mode, with a maximum power of 80W, enables precise cutting with significant hemostasis
| 6 | · 2 coagulation modes with automatic power adjustments for each mode
| 7 | o Low coagulation mode, with a maximum power of 80W
| 8 | o High coagulation mode, 60W
| 9 | · Features the possibility of handling a hand operated or a foot operated switch
| 10 | · Color LCD monitor with a diode screen i 4 buttons for memorizing the operational values; detected error message display
| 11 | · Detected connectivity function and a patient feedback diode impedance with a sound warning about the connection error
| 12 | · Automatic “power leakage“ detection (“draining current“) and automatic generator operating power reduction up to 90% high draining currents, and returning to normal operational power after the “draining currents“ reduction, thanks to the integrated operational safety hardware and software mechanism
| 13 | Wireless foot operated switch
| 14 | · A wireless receiver enables a connection between the generator and the foot operated switch
| 15 | Trolley for holding and transport of system components

## 2.Audiometer
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.2 | Audiometer |  | 1 |  | 0
| 1 | 125Hz do 20kHz
| 2 | Available Intensity Steps is 1, 2 or 5dB
| 3 | Manual or Reverse. Single or multiple pulses
| 4 | Narrow band noise
| 5 | White noise
| 6 | Speech Noise
| 7 | Basic audiometry
| 8 | Békésy Test.
| 9 | SISI
| 10 | Langenbeck (tone in noise)
| 11 | Stenger
| 12 | Modified Hughson-Westlake
| 13 | ABLB
| 14 | Speech testing with CD/Mic
| 15 | Build-in wave files
| 16 | Binaural Speech
| 17 | Hearing Loss Simulator (HLS)
| 18 | DD45 Audiometric headset
| 19 | HDA300 Phone headset for HF
| 20 | B71 Bone conductor
| 21 | Monitor Headset with boom
| 22 | Talk Forvard(TF) i Talk Back(TB)

## 3.Endoscope
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.3 | Retromolar Intubation Endoscopes with Camera |  | 1 |  | 0
| 1 | Retromolar Intubation Endoscope, autoclavable, outer diameter 5 mm, for ETT > 5.5 mm, usable sheath length 40 cm, distal bending 40°, with movable eyepiece, with Tube
| 2 | Holder for tube fixation and O2 application:
| 3 | o Battery Light Source LED for Endoscopes, rechargeable, with click connection, brightness > 110 lm / >150 klx, color temperature 5500 K, lithium-ion batteries, charging time 60 min, burning time at 100% brightness 40 min, ready for use, suitable for wipe disinfection
| 4 | o Charging Unit, for LED battery lightsources, with fixed integrated power supply and adaptor for EU, power supply 110 – 240 VAC, 50/60 Hz, suitable forwipe disinfection
| 5 | o Wire Tray for Cleaning, Sterilization and Storage of one rigid endoscope
| 6 | Camera Head, 8-pin, one-chip CMOS camera head, resolution 640 x 480, focal length f = 20 mm,
| 7 | Monitor for CMOS Endoscopes, screen size 7", with resolutionof 1280 x 800 pixels, documentation can be stored directly on SD card, rechargeable Li-Ion batteries,additional standards: RTCA/DO-160F, EMI Test Report (German air rescue service DRF Luftrettung) i MIL-STD-461F (Fixed Wing) power adaptor for EU, power supply 110 – 240 VAC, 50/60 Hz, suitable for wipe disinfection
| 8 | Mobile Stand for Accessories for Intubation Fiberscopes and Endoscopes:
| 9 | o Shelf, incl. installation accessories, max. capacity 12 kg,
| 10 | o Fixation Device, for Holder 10330 BC/BD to Standard Equipment Rail 25 x 10 mm
| 11 | o Holder, for Retromolar Intubation Fiberscope,made of plexiglass

## 4.Platform monitoring
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.4 | Platform for invasive hemodynamic monitoring of the patients |  | 1 |  | 0
| 1 | Operating system Windows XPe,
| 2 | Touch screen,
| 3 | Monitor must have RS232 Serial Port, USB port, LAN port, RJ-45 Ethernet connector and VGA port
| 4 | Monitor and data box must have the possibility to be fixed on a pole, or IV stand,
| 5 | Electromagnetic compatibility, Reference: EN 60601-1-2:2001,
| 6 | Operating temp. 0 - 37°C, humidity 10% - 95%,
| 7 | Storage temp. -25 - 70°C,
| 8 | Intermittent measuring of CO, SV, SVR, GEDV, ITBV, EVLW, PVPI, CFI, GEF
| 9 | Continious measuring of CCO, SV, SVR, SVV,
| 10 | Calculation of index values for measured parameters,
| 11 | Measuring of pressures in femoral artery ABP and central vein CVP,
| 12 | Continious measuring of oxygenation in mixed vein blood, or upper vein cava SvO2/ScvO2),
| 13 | Possibility of display of calculated values (DO2, VO2, CPO,..),
| 14 | Possibility of display of four coloured separated parameters in graphical and numerical way,
| 15 | Possibility of display of parameters in optimisation algorithm,
| 16 | Possibility of display of parameters in physiological state of small and large circulation,
| 17 | Possibility of comment entry (events),
| 18 | Tabelar screen of parameters continiously measured,
| 19 | Monitor operation via “touch screen” technics
| 20 | Monitor must be able to import and export data via analog, serial and USB ports,
| 21 | Possibility of screen snapshot and export via USB port in jpg. form,
| 22 | Possibility of communication with patient monitors, printers and IT system,
| 23 | Data storage for the last 72 hours.

## 5.Fiberscope
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.5 | Flexible intubation fiberscope |  | 1 |  | 0
| 1 | Intubation Fiberscope 3.7 mm x 65cm, with the suction valve
| 2 | Dflection up/down: 140°/140°
| 3 | Intubation Fiberscope 3.7 mmx 65cm, with the suction valve
| 4 | Direction of view: 0°
| 5 | Angle of view: 120°
| 6 | Working length: 65 cm
| 7 | Total length: 93 cm
| 8 | Working channel inner diameter: 1.5 mm
| 9 | Distal tip outer diameter:3,7mm
| 10 | With the following accessories:
| 11 | Suitcase, Lid for pressure compensation, leakage tester, LIPP holder tube body, The cleaning brush, The irrigation adapter, Via suction valve tube for placing, a bronchoscope, size 4, tube for the placement of the bronchoscope, size 2, Plug connector for cleaning LUER

## 6.Intnesive care monitoring
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.6 | Intensive care unit hemodynamic monitor |  | 1 |  | 0
| 1 | Continuous assessment of the hemodynamic status of critical care and surgery patients
| 2 | Continuous arterial waveform analysis system (pulse power) coupled to a single point lithium indicator dilution calibration system.
| 3 | Calculates continuous beat-to-beat cardiac output by analysis of the arterial blood pressure trace
| 4 | Touch screen size minimum 12”
| 5 | Arterial pressure waveform is slaved from conventional blood pressure monitor via an analogue cable link
| 6 | Minimally invasively continuous measurement of hemodynamic patients with the existing arterial line
| 7 | Calibration of the hemodynamic monitor using lithium chloride
| 8 | Provides access to preload response values or volume status indicators of: Systolic Pressure Variation (SPV), Pulse Pressure Variation (PPV%), Stroke Volume Variation (SVV%) as well as Heart Rate Varation (HRV%).
| 9 | Provides real time, beat-to-beat cardiac output and oxygen delivery
| 10 | Provides real time preload and afterload values
| 11 | Provides Intra-thoracic blood volume (ITBV)
| 12 | Displays the actual and indexed values for Cardiac Output, Systemic Vascular Resistance, Mean Arterial Pressure, Heart Rate and Left Ventricle Stoke Volume.
| 13 | Oxygen Delivery and Venous Oxygen Saturation can also be displayed
| 14 | Display allows the user to view up to 2 hemodynamic variables allowing evaluation of the patient’s response to targeted interventions such as a fluid challenges or changes in inotrope therapy.
| 15 | Showing the relationship between pressure, flow and resistance in an integral bar chart display.
| 16 | The graph screen
| 17 | Oxygen delivery and ScVo2/SvO2 display
| 18 | Event marker

## 7.Nebulizer
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.7 | Nebulizer based on vibrating mesh |  | 10 |  | 0
| 1 | USB controller
| 2 | Nois level: < 35 dB at a distance of 1m
| 3 | Mods duration of therapy: 30min, 6h
| 4 | AC/DC Adapter
| 5 | Nebuliser
| 6 | It does not affect on ventilator settings (Benet 840, Hamilton C2)
| 7 | The implementation of all drugs (solution/suspensions, antibiotics)
| 8 | Compatible with drugs: Ipratropium Bromide,Terbutaline,Ipratropium Bromide + Salbutamol,Acetylcysteine,Salbutamol,Levosalbutamol, Sevofluran
| 9 | The possibility of administering drugs continuously and intermittent
| 10 | The possibility of administering medication to a patient while the MV or via facial masks or mouthpieces
| 11 | A vibrating palladium mesh
| 12 | Flow: >0.2 ml/min ( ≈ 0.38 ml/min)
| 13 | Particle size: 1.5-6.2 μm (depending on the type of the drug)
| 14 | Aerosol Output rate: 0.30 ml/min
| 15 | Minim. Residual Volume: <0.1 ml for 3 mL dose
| 16 | The temperature of the medication will not rise more than 100C
| 17 | Nebuliser Weight : max 25g
| 18 | Nebuliser Capacity max.6 ml-10ml
| 19 | T pece adult

## 8.Prevention thromboembolism
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.8 | Device for Mechanical prevention of venous thromboembolism with active intermittent, sequential and graduated pneumatic compression of lower limbs |  | 3 |  | 0
| 1 | Mechanical prevention of deep vein thrombosis (VTE) and pulmonary embolism. with active intermittent, sequential and graduated pneumatic compression of lower limbs in the ICU and OR.
| 2 | The garments, both leg sleeves and foot cuffs, compress the limbs to enhance venous blood movement. After the compression cycle has reached set pressure, the Controller measures the time it takes for the limbs to refill with blood and waits that period of time before the next compression is initiated.
| 3 | Device for intermittent, sequential and graduated pneumatic compression of lower limbs
| 4 | Compact, lightweight device (max. 2.5 kg), portable
| 5 | Capable of operating on battery power min. 6 h
| 6 | Integrated carrier for hospital bed
| 7 | LCD color display with graphical information about current operational functions, settings and alarms
| 8 | The compression of the legs and feet, simultaneously or separately
| 9 | The compression of the lower extremities using compression cushions inflatable air associated with the device with transparent silicone hoses
| 10 | Graduated compression profile for compressing the legs from 45mmHg to 30 mmHg above the article in the thigh / leg femoral part
| 11 | Compression profile for foot: 130mmHg
| 12 | Compression cycle for leg compression: 10-12seconds
| 13 | Compression cycle for foot compression: 3-5 seconds
| 14 | The possibility of detecting and measuring venous filling time, and interval adjustment for the next compression cycle
| 15 | The device is registered in ALIMS (Agency for Medicines and Medical Devices) Serbia

## 9.ODM
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.9 | ODM+ - Oesophageal Doppler monitor |  | 1 |  | 0
| 1 | Apparatus for intra- and peri-operative minimal invasive hemodynamic monitoring through the esophageal Doppler probe (1A level of evidence - clinical records)
| 2 | Monitoring the flow of blood through the descendent aorta in real time, "beat-by-beat"
| 3 | Precise individualized calculation of hemodynamic parameters over the nomogram (patient age, sex, height and weight)
| 4 | LCD color monitor with audio component (the sound of blood flow through the descendent aorta)
| 5 | Display graphic of blood flow through the descendent aorta in the form of waves, the ratio of the speed and flow of time, with the option of stopping and painting of the current situation - "snap shot"
| 6 | Display Numeric: SV-stroke volume, SVI-index stroke volume, CO-output, CI-cardiac index, FTc- flow time , PV maximum flow rate, HR-heart rate, SVV-variations in stroke volume, as well as additional options calculations-SVR systemic vascular resistance and DO2-supply of oxygen to tissues
| 7 | Ability to connect to a monitor for the invasive monitoring of blood pressure, the calibration of parameters of the flow of the parameters analysis of the pulse wave curve and obtaining the additional values: PPV- pulse pressure variation, the CPO-effect cardiac power, and use of the device in the modality-independent of doppler flow detected, with the 12h period of the calibration
| 8 | The possibility of recording the entire course of monitoring and storing patient data
| 9 | The option networking and data transfer on USB
| 10 | The device is registered in ALIMS (Agencija za lekove i medicinska sredstva) (Agency for Medicines and Medical Devices) Serbia

## 10.Volumetric pump
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.1 | Multifunctional volumetric infusion pump with software for enteral nutrition |  | 5 |  | 0
| 1 | BASIC FUNCTIONS
| 2 | o Multi-functional, portable, volumetric infusion pump for periodic or continuous delivery of drugs, parenteral, enteral food or blood.
| 3 | o Delivery rate range: 0.1ml/h up to min. 1.200 ml/h.
| 4 | o Adjusting the flow interval:
| 5 | 0.1 – 99.99 ml/hin 0.01 mlincrements
| 6 | 100.0 – 999.9 ml/hin 0.1 ml increments
| 7 | 1.000 ml.- 1.200 ml/h in 1ml increments
| 8 | o Accuracy of set delivery rate: ± 5% (with Original Infusomat Line)
| 9 | o Bolus delivery:
| 10 | o Bolus on demand (manual bolus)
| 11 | o Bolus with volume/dose pre-selection (predetermined volume bolus)
| 12 | o Bolus over time (calculated flow rate)
| 13 | SELECTION OF FUNCTION
| 14 | o Volume pre-selection: 0,1-9.999 ml
| 15 | 0,1 – 99.99 ml in min. 0.1 ml increments
| 16 | 100.0– 999.9 ml in 0.1 ml increments
| 17 | 1.000– 9.999 ml in 1ml increments
| 18 | o Time pre-selection: 00h 01minup to min. 99h 59min.in 1 min increments
| 19 | o Automatic flow rate calculation by calculating volume over time(ml/h) based on dose entries in (µg, mg, ../kg/min)
| 20 | o Automatic flow rate calculation depending on:
| 21 | o concentration or amount of drug in µg, mg, IU, mmol, ilimEq/ml
| 22 | o amount of drug per ml, and the prescribed dose
| 23 | o patient weight
| 24 | o Displaying data on the pump screen during the operation of the pump:
| 25 | o flow and flow visualization
| 26 | o the therapy profile
| 27 | o battery status
| 28 | o the pressure in the line / limit pressure
| 29 | o time and volume by the end of therapy
| 30 | o the volume dosed to patient
| 31 | o The ability to create lists of drugs (min. 1,200 drugs) using the software program with the possibility of grouping of drugs in the category of at least 15 (with the parameters of the dose, a bolus dosing limits the speed / flow rate).
| 32 | o Mod of secondary infusio, temporarily stopping primary infusion with the purpose of application of secondary infusion
| 33 | o Mod- Patient Controlled Analgesia (PCA). Creating a list of PCA medicines, using a software program, in accordance with user needs
| 34 | o PCA joystick must be included in the equipment cost.
| 35 | o Target Controlled Infusions (TCI) mode with min. 2 or more of the pharmacokinetic models (Propofol, Remifentanyl, Sufentanyl, Alfentanyl)
| 36 | o Mode of automatically increase the flow rate of infusion to continuous application, depending on the volume and the time of the application of the infusion, application and maintenance infusion at a given speed and in a certain time, an automatic reduction in the flow rate.
| 37 | Use when applying parenteral and enteral food
| 38 | o Sequential infusion - the possibility of programming the min. 12 infusions sequences in a single therapy
| 39 | o Mod- programmed continuous delivery of therapy with occasional programmed bolus deliveries in a given period of time
| 40 | SAFETY SYSTEMS
| 41 | o Displaying the status of the pump (alarm) via 3 LED
| 42 | o Visual and acoustic early warning of :volume deviation, the expiry time, low battery, KVO active
| 43 | o Visual and audible alarms: supplied volume, the expiry time, low battery, unexpected pressure increase above the limit, the air in the line, KVO completed, calibration, data invalidated therapy, locking data
| 44 | o Setting the limit pressure in the system at 9 levels (0.1 to 1.2 bar), the automatic reduction of the bolus volume in the system after the occlusive alarm
| 45 | o The possibility of subsequent connection of the sdropsensor
| 46 | o Alarm at least 3 minutes before the end of treatment with the possibility of setting / adjustment time in service mode
| 47 | o KVO (Keep Vane Open) - option of maintaining a patent venous access for a period of 30 minutes after dosing program:
| 48 | o for flow < 10ml/h = min. 1ml/h
| 49 | o for flow < 1ml/h = adjusted rate (min. 0.1 ml/h)
| 50 | o Detection of air in line: volume of air in the line (the total value of the air, bubble size of 0.01 ml for 1 h)
| 51 | o Two detectors of air in lines ("upstream" and "downstream" from the pump)
| 52 | o Two safety clamps (one inside the pump and the other integrated in the original infusion line) for automatically stopping the flow of free fall when changing infusion line
| 53 | o Programmed washing with repeating the process to expel the air from the line without removing the line in case of an air alarm
| 54 | o Locking the data in stations in three levels depending on the patient
| 55 | o Event History for the last min. of 1,000 entries with the option of viewing the data by authorized service
| 56 | CONNECTIONS
| 57 | o Connection lead staff call systems
| 58 | o The possibility of the power supply in vehicles
| 59 | POWER SUPPLY
| 60 | o Power supply: Mains power supply 230/110 V, 50/60 Hz
| 61 | o External low voltage 12 V DC (the possibility of connecting in the ambulatory vehicle)
| 62 | o Rechargeable battery pack
| 63 | Battery type : NiMH (rechargeable)
| 64 | Battery operating time: min 4 hat 25 ml/h
| 65 | Battery recharging: approx 6 h
| 66 | o The possibility of replacing the battery without using tools
| 67 | ADDITIONAL PROPERTIES
| 68 | o Softver in Serbian
| 69 | o The ability to connect up to 2 pumps, one to another out of workstations and fixation to the rail, horizontal and vertical infusion stand and bedside
| 70 | o The ability to connect multiple pumps in the "infusion workstation" still in the meetings of the more "infusion workstations". Connectivity min. 12 pumps in a single unit. A set of "workstations" must be supplied with one power cord.
| 71 | o The configuration according to the user's needs
| 72 | o The ability to connect to the system for monitoring and glycose control in a certain range
| 73 | o Upgradeable for operation in magnetic radiation environment (MRI)
| 74 | o Moisture protection: Waterproof: IP22
| 75 | o Delivery workstation / s for storage and organization of volumetric pumps.
| 76 | o The workstations are included in the price of equipment
| 77 | Infusion workstation
| 78 | BASIC FUNCTIONS
| 79 | Workstations must be compatible with the offered volumetric infusion pumps
| 80 | Modular workstation setup for 4 volumetric pump in any combination
| 81 | Rationalization and systematization of the working area with the "line holders"
| 82 | Workstations cover with 3 LEDs (green, orange, red) - status indicator, centralization of visual and audible alarm in order to speed up the response of medical staff
| 83 | Adjustment the volume of the alarm
| 84 | Integrated handle on the top of the workstation for transporting workstations both inside and outside the hospital.
| 85 | Workstations housing has connector to connect a volumetric pump with workstation in order to provide power to the device and data transmission
| 86 | Automatic pumps safety lock when setting up the workstation.
| 87 | Simply unlock and pull the device from a workstation
| 88 | Connectivity ports for staff call, PCA joystick
| 89 | The ability to connect 6 work stations in the 'set' workstations, in two "pillars", the central power supply
| 90 | Equipment rack of workstation that is used to set up the workstation to the infusion stand or on a horizontal rail
| 91 | The possibility of cross-linking "workstation" in the computer Ethernet network (UTP)
| 92 | The ability to connect to the system for central monitoring of infusion therapies in the department
| 93 | Possibility USB port for data transfer and connection accessories
| 94 | Possibility USB port for bar-code reader.
| 95 | Power supply 240 V, 50/60 Hz
| 96 | Moisture protection: Waterproof: IP22

## 11.Laryngoscope
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.11 | Laryngoscope sets for special purposes |  | 1 |  | 0
| 1 | Tapered blade near end the handle for easy intubation even in terms of anatomical anomalies/4 units
| 2 | MACINTOSH
| 3 | Blades with replaceable fibre optics
| 4 | · tapered blade end near the handle for easy intubation, even for difficult anatomic conditions
| 5 | · matte stainless steel – prevents back reflection
| 6 | · replaceable light guide
| 7 | · easy to clean
| 8 | · environmentally-friendly and cost-effective
| 9 | · fibre optic illumination Ø 4 mm, 5500 single fibres minimum
| 10 | · can be treated in autoclave up to 134°C – approx. 4000 times
| 11 | The kit contains:
| 12 | · Xenon Fiberoptic battery handle medium-sized 2.5V
| 13 | · Macintosh fiberopticblades with replaceable fibre optic illumination, light guide Ø 4 mm, size. 2
| 14 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, light guide Ø 4 mm, size. 3
| 15 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, light guide Ø 4 mm, size. 4
| 16 | · Laryngoscope box for 1 handle, 3 blades
| 17 | Laryngoscope blade with large open angle / 4 units
| 18 | MACINTOSH Laryngoscope
| 19 | Laryngoscope blade with large open angle towards the handle for easy intubation of patients with cervical spine injury and obese
| 20 | Blades with replaceable fibre optics
| 21 | · matte stainless steel – prevents back reflection
| 22 | · replaceable light guide
| 23 | · easy to clean
| 24 | · environmentally-friendly and cost-effective
| 25 | · fibre optic illumination Ø 4 mm, 5500 single fibres minimum
| 26 | · can be treated in autoclave up to 134°C – approx. 4000 times
| 27 | The kit contains:
| 28 | · Xenon Fiberoptic battery handle medium-sized 2.5V
| 29 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, light guide Ø 4 mm, size. 3 (130 x 22 mm)
| 30 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, light guide Ø 4 mm, size. 4 (155 x 22mm)
| 31 | · Laryngoscope box for 1 handle, 2 blades
| 32 | Laryngoscope with movable blade tip (can be bent up to 60°) / 4 units
| 33 | Fibreoptic Laryngoscope
| 34 | The movable blade tip (can be bent up to 60°) facilitates lifting of the epiglottis and enables good visibility of the intubation process even if the anatomic conditions are difficult. Improved visibility and less lever action reduce the risk of injury to the larynx and teeth. The fibre optic illumination cannot be replaced by the customers themselves.
| 35 | Blades with replaceable fibre optics
| 36 | · matte stainless steel – prevents back reflection
| 37 | · replaceable light guide
| 38 | · easy to clean
| 39 | · environmentally-friendly and cost-effective
| 40 | · fibre optic illumination, light guide 3 x 7 mm, 5500 single fibres minimum
| 41 | · can be treated in autoclave up to 134°C – approx. 4000 times
| 42 | The kit contains:
| 43 | · Xenon Fiberoptic battery handle medium-sized 2.5V
| 44 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, size. 2 (115 x 17mm)
| 45 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, size. 3 (135 x 18mm)
| 46 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, size. 4 (155 x 18mm)
| 47 | · Laryngoscope box for 1 handle, 3 blades
| 48 | Laryngoscope with very large fibre optics – cross section ( > 6.000 lux) / 4 units
| 49 | Fibreoptic Laryngoscope with vary large fibre optics
| 50 | · matte stainless steel – prevents back reflection
| 51 | · with integrated fibre optics
| 52 | · no dirt-collecting recesses (easy to clean)
| 53 | · very large fibre optics – cross section (3.0 x 7.0 mm), 8000 single fibres minimum for optimal light emission
| 54 | · can be treated in the autoclave up to 134°C – approx. 4000 times
| 55 | The kit contains:
| 56 | · Xenon Fiberoptic battery handle medium-sized 2.5V
| 57 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, size. 2 (115 x 17mm)
| 58 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, size. 3 (135 x 18mm)
| 59 | · Macintosh fiberoptic blades with replaceable fibre optic illumination, size. 4 (155 x 18mm)
| 60 | · Laryngoscope box for 1 handle, 3 blades

## 12.Laryngoscope replaceable
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
| 4.12 | Laryngoscopes with replaceable lights |  |  |  | 0
1 | 4,12,1 | MACINTOSH Laryngoscope Blade- cold light with with replaceable fiber optic light carrier |  | 4 |  | 0
| 1 | · MACINTOSH Laryngoscope Blade,with replaceable fiber optic light carrier,size4, pieces 1
| 2 | · Spare Fiber Optic Light Carrier, for 8542 A, pieces 1
| 3 | · MACINTOSH Laryngoscope Blade,with replaceable fiber optic light carrier,size 3,pieces 1
| 4 | · Spare Fiber Optic Light Carrier, for 8542B, pieces 1
| 5 | · MACINTOSH Laryngoscope Blade,with replaceable fiber optic light carrier,size 2, pieces 1
| 6 | · Spare Fiber Optic Light Carrier, for 8542C, pieces 1
| 8 | · MACINTOSH Laryngoscope Blade,with replaceable fiber optic light carrier,size 1,pieces 1
| 9 | · Spare Fiber Optic Light Carrier, for 8542D, pieces 1
| 10 | · Handle Sleeve, ISO 7376, autoclavable, length 12 cm, foruse with Battery Inserts 8546 A,8546 LD, 8549 LD and Xenon cold light laryngoscopes, pieces 4
| 11 | · Battery Insert, Xenon, length 12 cm, with 2 Batteries, pieces 4
| 12 | · Batteries, Baby-Cell, LR 14, for Battery Inserts, package of 2, pieces 4
2 | 4,12,2 | MACINTOSH Laryngoscope Blade- warm light fiber optic light carrier incorporated |  | 4 |  | 0
| 1 | · MACINTOSH Laryngoscope Blade, size 5, with warm light and Fiber Optic Light Carrier Incorporated, pieces 1
| 2 | · MACINTOSH Laryngoscope Blade, size 4, with warm light and Fiber Optic Light Carrier Incorporated, pieces 1
| 3 | · MACINTOSH Laryngoscope Blade, size 3, with warm light and Fiber Optic Light Carrier Incorporated, pieces 1
| 4 | · MACINTOSH Laryngoscope Blade, size 2 with warm light and Fiber Optic Light Carrier Incorporated, pieces 1
| 5 | · MACINTOSH Laryngoscope Blade, size 1, with warm light and Fiber Optic Light Carrier Incorporated, pieces 1
| 6 | · MACINTOSH Laryngoscope Blade, size 0, with warm light and Fiber Optic Light Carrier Incorporated, pieces 1
| 8 | · Spare Lamp, warm light,for laryngoscope blades with Fiber Optic Light Carrier Incorporated, sizes 5 – 2,pieces 10
| 9 | · Spare Lamp, warm light,for laryngoscope blades with Fiber Optic Light Carrier Incorporated, sizes 1 – 0,pieces 10

## 13.Videolaryngoscope
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.13 | Video Laryngoscope for visual endotracheal intubation -set with three video spatula |  | 1 |  | 0
| 1 | · Video Laryngoscope, CMOS technology, with MACINTOSH laryngoscope blade, size 2, for use with Electronic Modules for Video Laryngoscope, CMOS technology, processing of the electronic module: Suitableand validated for the following low-temperature, reprocessing methods up to bis max. 60 °C:manual/machine cleaning and disinfection, sterilization with Steris AMSCO V-PRO 1,Sterrad (50S, 100S, 200S, NX, 100NX) and EtOgas. pieces 1
| 2 | · Video Laryngoscope, CMOS technology, with DÖRGES laryngoscope blade, for difficult intubation, with catheter introduction sizes 16 – 18 Fr., for use with Electronic Modules Processing of the electronic module: Suitable and validated for the following low-temperature reprocessing methods up to bis max. 60 °C: manual/machine cleaning and disinfection, sterilization with Steris AMSCO V-PRO 1, Sterrad (50S, 100S, 200S, NX, 100NX) and EtO and EtOgas. pieces 1
| 3 | · MILLER Video Laryngoscope, CMOS technology, size 0, for use with Electronic Modules,Processing of the electronic module: Suitable and validated for the following low-temperature reprocessing methods up to bis max. 60 °C: manual/machine cleaning and disinfection, sterilization with Steris AMSCO V-PRO 1,Sterrad (50S, 100S, 200S, NX, 100NX) and EtO, gas. pieces 1

## 14.Hypothermia
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.14 | Sets for therapeutic hypothermia / normothermia |  | 10 |  | 0
| 1 | · Special sets, flexible panels for non-invasive patient cooling
| 2 | · Sets must have the possibility of hospital and out- of -hospital use
| 3 | · Sets must be made of material that can be applied to the body of the patient and allows quickly reaching the desired temperature.
| 4 | · Cooling of the patient minimum 3⁰C / h
| 5 | · Sets must have the self-adhesive surface for easier application to the patient
| 6 | · Sets should be compatible in case of need for recording patient CT / MRI units
| 7 | · Simple and quick application sets the body of the patient
| 8 | · Maintaining the temperature set for a patient, consists of 1 flexible cooling plate A4 size

## 15.Pacemaker
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.15 | Temporary single chamber exteranl pacemaker, 2 units and 4 patient cables |  | 2 |  | 0
| 1 | Pacingmodes: AAI, AOO, VVI, VOO
| 2 | Basic pacing rates: 30-200ppm
| 3 | Rapid atrial pacing rates: 80-800ppm
| 4 | Output amplitude: 0,1 – 25mA
| 5 | Pulse width: 1,5ms
| 6 | Sensitivity: 0,4-20mV
| 7 | Weight˂500g
| 8 | Batteries : 2 x 1,5V
| 9 | Digitaldisplay

## 16.Instruments
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.16 | Instruments for abdominal surgery -set |  | 10 |  | 0
| 1 | Roux retractor length 15,5 cm ; 27x26 i 32x37 mm, 2 pieces
| 2 | Parker-Langenbeck retractor set, length 21cm; 26x15 mm /40x15, 2 pieces
| 3 | Volkmann sharp retractor ,length 22,5 cm, 4 prongs,; 10 x 18 mm, 2 pieces
| 4 | Harrington retractor length 32 cm, 150 x 62 mm , 1pc
| 5 | Hosel retractor length 25 cm,140 x 30 mm, 1pc
| 6 | Mikulicz liver retractor length 26cm, 125x50mm, 1pc
| 7 | Thorlakson retractor for pelvis, 33cm, 170x40mm, 1pc
| 8 | Fritsch retractor length 25 cm, 45 x 75 mm, 1pc
| 9 | Spatula Reverdin 280x41/65 mm, 1pc
| 10 | Pincette (russ model) length 15 cm, 2pcs
| 11 | Pincette (russ model) length 25,5 cm, 2pcs
| 12 | Pincette tissue 1x2 teeth, standard, straight,length 130 mm, 3 pcs
| 13 | Pincette dissecting (anatomic) standard, straight, length130 mm, 2pcs
| 14 | Pincette dissecting (anatomic) standard, straight, length180 mm, 2pcs
| 15 | Pincette dissecting (anatomic) standard, straight, length250 mm, 2pcs
| 16 | Pincette potts-smith straight, seratted 0,4 mm, length 160 mm, 1 pcs
| 17 | Pincette potts-smith straight, seratted 0,4 mm, length 185 mm,1 pc
| 18 | Pincette potts-smith straight, seratted 0,4 mm, length 210 mm, 1pc
| 19 | Pincette DeBakey straight, width 2 mm, length 155 mm 2pcs
| 20 | Pincette DeBakey straight, width 2 mm, length 200 mm 2pcs
| 21 | Pincette DeBakey straight, width 2 mm, length 240 mm 2pcs
| 22 | Pincette DeBakey straight, width 2 mm, length 305 mm
| 23 | Scissors surgical straight ,sharp/blunt ,length 130 mm 1pc
| 24 | Scissors dissecting Metzenbaum curved,length 180 mm 1pc
| 25 | Scissors dissecting Metzenbaum curved ,length 205 mm 1pc
| 26 | Scissors dissecting Metzenbaum curved, length 230 mm 1pc
| 27 | Scissors dissecting Nelson-Metzenbaum curved ,length 285 mm 1pc
| 28 | Scalpel handle no. 4 length 130 mm 1pc
| 29 | Scalpel blade remover, 14cm, 1pc
| 30 | Needle holder Mayo-Hegar heavy length 160 mm, 2pcs
| 31 | Needle holder De Bakey length 150 mm, 2pcs
| 32 | Needle holder De Bakey length 180 mm, 2pcs
| 33 | Needle holder De Bakey length 205 mm, 2pcs
| 34 | Needle holder De Bakey length 265 mm, 2pcs
| 35 | Mixter forceps curved length 23 cm,1pc
| 36 | Pean Baby-Mosquito length 100 mm,curved, full gripp,delicate, 4 pcs
| 37 | Pean length 130 mm, curved,full gripp,delicate, 6pcs
| 38 | Rochester-Pean length 185 mm, curved, full grip, 6 pcs
| 39 | Rochester-Pean length 240 mm, curved, full grip, 2pcs
| 40 | Rochester-Pean length 240 mm, straight, full grip, 2pcs
| 41 | Gemini-Mixter length 160 mm, curved, full grip, 2pcs
| 42 | Gemini-Mixter length 185 mm, curved, full grip, 2pcs
| 43 | Gemini-Mixter length 205 mm, curved, full grip, 2pcs
| 44 | Pean (Overholt) length 205 mm,curved, fine, full grip, 2pcs
| 45 | Pean (Overholt) length 215 mm,curved, fine, full grip, 2pcs
| 46 | Pean (Overholt) length 220 mm,curved, fine, full grip, 2pcs
| 47 | Mikulicz forceps length 185 mm,curved, 1x2 teeth , 2pcs
| 48 | Babcock length 170 mm, 2pcs
| 49 | Maier tupfer forceps length 260 mm, straigth with ratchet, 2pcs
| 50 | Maier tupfer forceps length 260 mm, curved with ratchet , 2pcs
| 51 | Backhaus length 110 mm, 14pcs
| 52 | Foerester-Ballenger tupfer forceps length 245 mm, serrated,straight, 4 pcs
| 53 | Foerester-Ballenger tupfer forceps 185 mm, serrated,straight, 1 pc
| 54 | Ulrich tupfer forceps length 220 mm, heavy serrated,straight, 1 pc
| 55 | Round bowl volume 0,45 l, 1 pc
| 56 | Round bowl volume 0,70 l, 1 pc
| 57 | Round bowl volume 1,10 l, 1 pc
| 58 | Round bowl volume 1,60 l, 1 pc

## 17.Neonatal
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
| 4.17 | Neonatal Resuscitation and Jaundice |  |  |  | 0
1 | 4,17,1 | Resuscitation table for newborns |  | 1 |  | 0
| 1 | Examination light>0.11 lumens/cm2 (100 ft. Candles)
| 2 | Heating alarms:
| 3 | o Check patient
| 4 | o Baby temperature
| 5 | o High temperature
| 6 | o Probe
| 8 | o System fail
| 9 | o Power fail
| 10 | Set temp range34-38° C (>37.0° override)
| 11 | Heater output0-100%, 10% increments
| 12 | Skin temp display18-43° C
| 13 | Patient gas supplyFlow control range0-15 LPM
| 14 | suction0-150 mmHg (0-20 kPa)
| 15 | Autobreath infant resuscitator
| 16 | Adjustable breath rate range18-60 BPM (+/-10% of setting)
| 17 | I:E ratioFixed at 1:2 Nomina
| 18 | Pressure (PEEP)0-18 H20 (0-1.77 kPa)
2 | 4,17,2 | Jaundice meter |  | 1 |  | 0
| 1 | Screens infants as young as 24 weeks gestational age
| 2 | Reusable probe just wipe with alcohol and reuse.
| 3 | Battery Internal number of measurements (when fully charged) up to 250
| 4 | Touch screen display
| 5 | easy to navigate through menu settings;
| 6 | choose to see results in mg/dL or μmol/L;
| 7 | adjust contrast and sensitivity of display
| 8 | Internal memory
| 9 | Holds up to 100 patient files

## 18.Hose Holder
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
| 4.18 | HOSE HOLDER for anesthetic systems |  |  |  | 0
1 | 4,18,1 | Hose retainer plate |  | 10 |  | 0
| 1 | Hose holder plate, for ball mounting, for all cable and hose dimensions
| 2 | As well as for coaxial and pediatric hoses.
| 3 | Applicable for quickstop hinged arm
| 4 | Hinged arm 240° and hinged arm Infinity ACS
2 | 4,18,2 | Hinged arm |  | 10 |  | 0
| 1 | Hinged arm Quickstop short, with universal tube holder plate
| 2 | For various cable and tube dimensions and coaxial hoses.
| 3 | Specifically designed for OR tables.
| 4 | Contains:
| 5 | o 1 x Universal tube holder plate for various cable and tube dimensions
| 6 | o Length: 403 mm
| 7 | o Load capacity: max. 4 kg

## 19.Mobile bathroom
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.19 | Mobile bathroom |  | 3 |  | 0
| 1 | Unit is fully autonomous
| 2 | Fresh water tank capacity 22 litres
| 3 | Used water tank capacity 22 litres
| 4 | 1 diaphragm pump ( 40 W, 220 Volts IP44) provides water to showerhead
| 5 | Power output pump: 4 bars, shower entry 2,5 bars to 2,8 bars depending on water flow
| 6 | Sealed pressure reducer + water pressure, either 1 bar to 0,5 bar at the output depending on water flow
| 7 | Adjustable water flow - 5 to 7 l/min
| 8 | 1 diaphragm pump ( 40 W, 220 Volts IP44 ) collects water
| 9 | Suction power- 9 l/min
| 10 | Temperature setting (digital controller)- 0°C to 40°C
| 11 | Drainage time of the recovery tank - 4 min
| 12 | Length of power cable (moulded plug) - 4,50 m
| 13 | Length of the suction hose with filter -1,80 m
| 14 | Length of the shower hose-2,00 m
| 15 | Noise level -50 Decibels
| 16 | Double wheel bearings, swivel on double bearing seal
| 17 | 220 Volts and 2700 Watts resistance
| 18 | Dimensions: obstruction: 0,29 m² - L. 0,80 m, W. 0,37 m, H. 0,97 m, Weight 40 kg
| 19 | Height of dump tank (on raised toilet): 0,51 m
| 20 | Security: Automatic temperature cut out at 40°C,visual control system
| 21 | Device while stainless steel
| 22 | This equipment complies with European standards and meets CE/EU norms
| 23 | Possibility of cleaning with detergents, surface disinfectants
| 24 | Suction hose (for a used water) of the silicone can be disassembled, cleaned and disinfected by immersion in detergent and disinfectant agent
| 25 | If used for infectious patients: the possibility of application disinfectant solution through a shower system

## 20. Coasters for ot
Bidder:
LOT 4 |  | Auxiliary surgical devices |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 4.2 | Coasters medical-grade silicone for operating tables -set (Equipment for Pressure care Management in the OR and ICU) |  | 5 |  | 0
| 1 | Positioning device used to reduce pressure on patient's body parts and prevent tissue damage (Pressure Ulcer, Pressure Neuropathy)
| 2 | Should have the documented ability to reduce capillary interface pressures to 32mmHg or less
| 3 | All products can be easily cleaned using standard OR environment detergents
| 4 | All products are latex-free and non-allergenic.
| 5 | Biocompatible
| 6 | Anatomically shaped individual elements
| 7 | Medical Grade Silicone
| 8 | The gel used in the products is a cured silicone and does not flow.
| 9 | It is an inert material and does not support the growth of bacteria.
| 10 | x-ray translucent, MRI Compatible, Reusable
| 11 | The products can be cooled to -12ºC, and heated to a maximum of 40ºC
| 12 | Can be safely used in conjunction with heated blankets.
| 13 | One kit should contain:
| 14 | · Supine Head Rest Adult (275mm x 237mm x 90mm)
| 15 | · Two Contoured Armboard Pad (510mm x 150mm x 35mm)
| 16 | · Sacral Protector Large (450mm x 445mm x 30mm)
| 17 | · Two Leg Protector (515mm x 180mm x 40mm)
| 18 | · Leteral Leg Positioner (650mm x 400mm x 230mm)
| 19 | · Pair Crutch Stirrup Pads (395mm x 245mm x 10mm)
| 20 | · Pair Heel Pads (18mm x 100mm x 70mm)