## STERILIZATION
Bidder: |  | Offered price:
LOT 2 | STERILIZATION
ID | Equipment name | QTY | Technical Specification Offered(model) | Total Price per line item
2.1 | Substerilization equipment
2,1,1 | Rotary Heat Sealers for bonding sterilization foil 1 | 3
2,1,2 | Wire basket, 1 STU | 40
2,1,3 | Roll holders with cutting device | 4
2,1,4 | Trolley for transport sterile material | 2
2,1,5 | System for tracking and tracing in the system of sterile instruments supply | 1
2,1,6 | Chairs without back rest | 10
2.2 | Automatic steam sterilizer | 1
2.3 | Air/space desinfection based on Ag/H2O2 solution | 3
2.4 | Central sterilization
2,4,1 | Automatic steam sterilizer | 2
2,4,2 | Plasma sterilizer | 1
2,4,3 | Duplex softener | 1
2,4,4 | Washer – disinfector | 2
2,4,5 | Cloth cutting machine | 1
2,4,6 | Trolley for transport sterile material | 5
2,4,7 | System for tracking and tracing in the system of sterile instruments supply | 1
2,4,8 | General table | 2
2,4,9 | Height adjustable packing table | 3
2,4,10 | Instrument washing sink | 1
2,4,11 | Linen inspection table | 1
2,4,12 | Paper trolley | 2
2.5 | Auxillary equipment for sterilization
2,5,1 | Sterilization box | 60
2,5,2 | Wire basket, 1 STU | 60
2,5,3 | Wall brackets s | 5
2,5,4 | Fixed solid shelf system | 2
2,5,5 | Magnifying lamp | 2
2,5,6 | Spray gun | 1
2,5,7 | Overhead shower | 1
2,5,8 | Ultrasonic cleaning unit | 1
Total Price per Lot

## 1.Subster.equip.
Bidder:
LOT 2 |  | STERILIZATION |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no.in technical documentation
| 2.1 | Substerilization equipment |  |  |  | 0
1 | 2,1,1 | Rotary Heat Sealers for bonding sterilization foil 1 |  | 3 |  | 0
| 1 | Rotary Heat Sealers for bonding sterilization foil 1
| 2 | Housing made of stainless steel
| 3 | Microprocessor control
| 4 | The regulation is done by means of LCD or "touch-screen monitors diagonal not less than 4”
| 5 | Auto start transport, reverse transport and stand-by function
| 6 | Security protection measures, which cuts off the welding process if the difference between the set and actual temperature is more than 5 ° C
| 7 | Timer with audible alarm
| 9 | The integrated printer capable of printing (date, deadline, codes and text) on the package, width min. 30 characters
| 10 | Ability to connect to an external printer
| 11 | Possibility of programming min. 2000 texts
| 12 | Rotating Text 180 °
| 13 | Constant monitoring of the welding parameters (temperature, speed, pressure); audible warning and termination of welding in the event that a parameter is incorrect.
| 14 | PC connection via RS232 or Ethernet or USB port.
| 15 | The ability to connect USB drives and storage of log files on it
| 16 | Features calendar and clock
| 17 | Seal width min. 15 mm
| 18 | Min welding speed 10 m / min
| 19 | Temperature settingin a range50-199° C
| 20 | Maximum power consumption 365 W
| 21 | Dimensions maximum 530x320x170 mm
| 22 | Easy opening for servicing without tools
The device must comply with the following standards:
| 23 | EN ISO 11607-2
| 24 | DIN 58953.7: 2010/2006 / 42CE / 2006/95 / CE
2 | 2,1,2 | Wire basket, 1 STU |  | 40 |  | 0
| 1 | Baskets are intended for storage and transportation of non-sterile and sterile material
| 2 | Baskets are made of wire mesh;raster prevents standard elements from falling out.
| 3 | Square shape
| 4 | Made of stainless steel AISI 304
| 5 | Type: 1 STU,
| 6 | Dimensions: approximately 600x300x300 mm (LxWxH)
3 | 2,1,3 | Roll holders with cutting device |  | 4 |  | 0
| 1 | Housing made of stainless steel AISI 304 except knife and roller on which rolls slide
| 2 | Self-sharpening knife
| 3 | Desk model that can be attached to the wall
| 4 | Capacity (roll width) min. 900 mm
| 5 | Dimensions max. 1120x420x130 mm
4 | 2,1,4 | Trolley for transport sterile material |  | 2 |  | 0
| 1 | Trolley for transport of sterile containers and baskets for 6 sterilization unit
| 2 | 2 pushing handles
| 3 | Two doors on hinges, lockable
| 4 | Door seal around the whole door and frame
| 5 | Wheels diameter min. 160 mm that do not leave traces, 2 fixed and 2 craft item. brakes on 2 wheels
| 6 | External dimensions max. 980x710x1500 mm
5 | 2,1,5 | System for tracking and tracing in the system of sterile instruments supply |  | 1 |  | 0
| 1 | A system that provides flow control of sterile instruments or instruments to be sterilized in more positions
| 2 | The system has the ability to make document at each stage of the sterile instruments supply
| 3 | The system is connected via computer to a hospital computer network and it can be supervised in accordance with individual access permissions
The system is located at the following locations in the sterile supply chain:
| 4 | o return of instruments from operating theatres
| 5 | o washing and disinfection of instruments
| 6 | o washing and disinfection process control (data are generated automatically by the machine for each batch)
| 7 | o checking and packing of instruments
| 8 | o sterilization
| 9 | o sterilization process control (data are generated automatically by the machine for each batch)
| 10 | o delivery of sterile instruments
| 11 | o the place of use of sterile instruments (operating theatres)
| 12 | System that controls the flow of sterile instruments or instruments that are sterilized in 6 positions: the process of returning used instruments, cleaning and disinfection, checking and packaging, sterilizing, issuing and handing over to the operating room or other user.
| 13 | The system integrates all devices in the chain of sterilization: sterilizers, washer-disinfectors, heat sealing machines
| 14 | The system recognises instruments according to the identification located on the instrument and package (bar code, RFID or the corresponding)
| 15 | The system has a database of instruments and, on-demand, it shows the current status of sterile supplies, position of the instruments in the chain
| 16 | The ability to administratively update set data when there is a change in their content by completely ejecting sets or combining new sets.
| 17 | Database with limited access. The modification of data only by authorized persons
6 | 2,1,6 | Chair without backrest or with backrest |  | 10 |  | 0
| 1 | For entire sterilization required centers
| 2 | Made of stainless steel AISI 304 or better
| 3 | Anti-static and anti bacterial seat
| 4 | Aluminum injection cast foot system, with 5 castors
| 5 | Moving by height by foot or hand command
| 6 | Dimensions: 890x700x500 mm ±5%
ALIMS registration is not necessary for equipment 1-6

## 2.Steam sterilization
Bidder:
LOT 2 |  | STERILIZATION |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no.in technical documentation
1 | 2.2 | Automatic steam sterilizer |  | 1 |  | 0
| 1 | Automatic steam sterilizer with one horizontal opening door, with electromechanical locking
| 2 | It is used for steam sterilization at temperatures of 121 ° C and 134 ° C
| 3 | The chamber is horizontal, square cross-section, capacity of 1 STU
| 4 | Chamber volume min 70 l
| 5 | Sterilizing chamber and double jacket material: CrNiMo steel (AISI 316L) or better
| 6 | The device has 1 transport trolley and 1 loading trolley
| 7 | Outer paneling made of chrome-nickel steel sheet or better
| 8 | The chamber insulated with high-quality insulating materials covered with galvanized metal sheets or better.
| 9 | Sterilizer has a safety system against door opening during the cycle
| 10 | The sterilizer has a vacuum pump with water ring
| 11 | The control panel with touch screen display diagonal larger than 8 " on the front side
| 12 | Showing process, any errors or alarms and self- diagnostics on the screen
| 13 | Sterilizer is driven by a microprocessor
| 14 | Audio visual alarm
| 15 | Process control sterilization with dual Pt100 sensor or better
Cycles:
| 16 | o Instruments 134 ° C
| 17 | o Textile 134 ° C
| 18 | o Rubber 121 ° C
| 19 | o The Bowie and Dick test 134 ° C
| 20 | o Vacuum test
| 21 | o Programs to customers’ demands
| 22 | Possibility of programs input by chip card or through PC.
| 23 | Sterilizer has its own steam generator with heaters submerged in water max.power 7.5 kW
| 24 | Power supply 380V / 50Hz
| 25 | Dimensions (HxWxD) max 2500 x 700 x 900 mm

## 3.Desinfection
Bidder:
LOT 2 |  | STERILIZATION |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no.in technical documentation
1 | 2.3 | Air/space desinfection based on Ag/H2O2 solution |  | 3 |  | 0
| 1 | Based on hydrogen peroxide stabilized by a colloidal silver complex (30 ppm)
| 2 | Ready for use aqueous solution
| 3 | Bactericidal, virucidal, sporicidal and fungicidal
| 4 | Neutral fragrance
| 5 | Weight less than 6 kg
| 6 | Based on electric turbine with speed at least 20 000 rpm
| 7 | No pump for the injection of fluids
| 8 | Fog spray by Venturi effect up to 15 metres away from the appliance
| 9 | Fog output speed at least 80 m/s
| 10 | Automatic stop, delayed start
| 11 | Manufacturers standard ISO 9001 or ISO 13485
| 12 | Treatment with one filling of biocides min 800 m3

## 4.CSSD
Bidder:
LOT 2 |  | STERILIZATION |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no.in technical documentation
| 2.4 | Central sterilization |  |  |  | 0
1 | 2,4,1 | Automatic steam sterilizer |  | 2 |  | 0
| 1 | Automatic steam sterilizer with horizontal electromechanical or pneumatic door
| 2 | The sterilizer should be pass-through, with two doors
| 3 | The sterilizer must be intended for water-steam sterilization (autoclaving) on the temperatures 121°C and 134°C
| 4 | The chamber is pass-through, horizontal, square cross-section. The chamber is adapted for easy loading, loading height max. 500 mm
| 5 | The capacity of minimum 650 l for 9 STU (standard sterilizing unit 600x300x300mm)
| 6 | Chamber and jacket material:
| 7 | Chamber made of stainless steel AISI 316L (CrNiMo) or better
| 8 | Jacket made of stainless steel AISI 316L or AISI 316Ti
| 9 | Construction and paneling made of stainless steel
| 10 | The chamber is insulated with high quality insulating material covered with galvanized metal sheet or better
| 11 | Automatic horizontal sliding door open with an electric motor or pneumatic cylinder
| 12 | Sterilizer is equipped with door security systems:
| 13 | o System against door opening during the cycle
| 14 | o System against opening both sterilizer doors at the same time
| 15 | Sterilizer has a double automatic system for protection of the user form door crushing
| 16 | Piping and valves are made of copper or stainless steel, the pipes are heat insulated
| 17 | Sterilizer is equipped with two-step vacuum pump with water ring. Recirculation water saving system with Pt100 probe or better
| 18 | On the non-sterile side of the touch screen diagonal min 10", and on the sterile side of the appliance, a panel for opening and closing the door and light signaling for the status of the sterilizer is sufficient
| 19 | Sterilizer is guided by microprocessor that is programmable and has a possibility of PC and web connection (Ethernet).
| 20 | Sterilization process control by double Pt100 probe and double pressure transducer.
| 21 | Printer that is showing the cycle parameters and graphs of the sterilization process.
Cycles:
| 22 | o Instruments (134°C),
| 23 | o Textile (134°C),
| 24 | o Rubber (121°C),
| 25 | o Bowie & Dick Test (134°C),
| 26 | o Vacuum Test,
| 27 | o Programs on customer’s demand (defined by user and programmed upon delivery and installation).
| 28 | The sterilizer has integrated electronic Bowie&Dick test, in accordance with DIN EN ISO 11140 standard.
| 29 | The device must have its own steam generator total power maximum 54кW
| 30 | Power: 380V/50Hz, three phase
| 31 | Sterilizer conforms to the following standards
| 32 | DIN EN ISO 9001
| 33 | EN ISO 13485
| 34 | CE sign
| 35 | EN 285
| 36 | Directive MDD 93/42/EEC
| 37 | Directive PED 97/23/EC
| 38 | Dimensions of the sterilizer should be maximum
| 39 | HeightxWidthxDepth=1920x1600x1300mm
| 40 | Sterilizer must be equipped with 1 loading cart and 2 transport wagons
2 | 2,4,2 | Plasma sterilizer |  | 1 |  | 0
| 1 | Sterilizer is intended for the sterilization of thermolabile materials at temperature of 50° C to 55° C
| 2 | The chamber is pass-through, square section, of the total volume min. 140 liters
| 3 | A system with two sliding doors
| 4 | Sterilant agent H₂O₂ (hydrogen peroxide)
| 5 | Bottle with sterilant / cartridge amounts are sufficient for min. 8 cycles
| 6 | The device is mobile with 4 wheels with brakes which ensure stability during operation
| 7 | Built in printer
| 8 | Cycle time for sterilizing lumens at three sterilization phases max.60 minutes, in one sterilization phase max. 70 min.
| 9 | Three sterilization programs:
| 10 | - Surface
| 11 | - Standard
| 12 | - Advanced
| 13 | Power supply 220V / 50Hz
| 14 | The device does not require any additional supplys or installations except electricity.
| 15 | Managing and control via touch-screen monitor diagonal min. 10 "
| 16 | Sterilization of lumens:
| 17 | o Flexible: - Ø1x1.000 - min. 10 pcs
| 18 | o -Ø1x4.000 – 1 pc
| 19 | o Rigid: - Ø1x400 – min. 10 pcs
Norms and standards:
| 20 | o 93/42 / EEC
| 21 | o ISO 9001: 2008
| 22 | o ISO 13485: 2003
| 23 | o ISO 14971: 2001
| 24 | Sterilizer and bottle with sterilant agent/ cartridge must be registered in ALIMS
| 25 | Dimensions: max. 2500x1100x1100 mm (WxHxD)
3 | 2,4,3 | Duplex softener |  | 1
| 1 | Flow of min1000 l / h
| 2 | Column( 2 pieces)
| 3 | Column filling 20Lx2
| 4 | Type of column - composite 8x35
| 5 | Fulfill type-cation (regeneration is done using salt tablets)
| 6 | Number of solanks 40L - 1 piece
| 7 | Received wather - 0dh
| 8 | Water entry ¾”
| 9 | Wather exit ¾”
| 10 | Drain for sewage ½”
| 11 | Digital control unit with automatic control unit, volumetric control, the possibility of manual launching regeneration process, parameter readings on display
| 12 | Tank 200L , closed type, receives RO water and controls the RO system
| 13 | Reverse osmotic module with housing min 200 lit
| 14 | Entrance ½”
| 15 | Exit for permeat ½”
| 16 | Exit for concentrate ½”
| 17 | Inlet valve
| 18 | Outlet valve
| 19 | Booster pump for supplying RO membranes
| 20 | Protective circuit high-pressure
| 21 | Protective circuit low pressure
| 22 | Power switch on/off
| 23 | The housing 10 "with filter cartridge 10" 2 pieces
| 24 | Microprocessor control for automatic managing
| 25 | ALIMS Registration is not necessary
4 | 2,4,4 | Washer – disinfector |  | 2
| 1 | The capacity of the device min 10 DIN trays
| 2 | The device for washing and thermal disinfection and drying of instruments
| 3 | The device should be pass-through, with two doors.
| 4 | The chamber is made of high quality stainless steel AISI 316L or better
| 5 | Flow control for dosing units
| 6 | Wash pump, capacity min. 620 l/min
| 7 | Foam control
| 8 | Pump pressure monitoring
| 9 | High-capacity drain valve, min 5 l/s
| 10 | Electric heating of the chamber (water) power max 20 kW
| 11 | Ventilator for drying with HEPA filter
| 12 | Control with touch panel with LCD screen or LCD screen with dirty buttons, and LCD display on the clean side
| 13 | Disinfection by A0 - value
| 14 | Temperature monitoring during complete cycle
| 15 | The device must have possibility for PC and bar code reader connection by different interfaces
| 16 | Acoustic signal at program end
| 17 | Maintenance reminder
| 18 | Front covers made of stainless steel AISI 304 or better, with glass doors
| 19 | Chamber volume min. 220 l
| 20 | Dimensions (HxWxD): max. 1850x680x720 mm
Equipment that is to be delivered with washer-disinfector:
| 21 | o Transport cart, 2 pc
| 22 | o Instrument rack for loading 10 standard instrument trays, 1 pc
| 23 | o Rack for AN material, 1 pc
| 24 | o Rack for containers for sterilization, 1 pc
| 25 | o Rack for OR shoes, 1 pc
| 26 | o MIS Rack
| 27 | EN ISO 15883
| 28 | ALIMS Registration
5 | 2,4,5 | Cloth cutting machine |  | 1
| 1 | Cutting knife diameter max 125 mm
| 2 | Power supply 220-230 V, 50Hz or 380-400 V, 50Hz
| 3 | Power of the device max 400W
| 4 | Capacity 25 mm / 1 "(seconds)
| 5 | ALIMS Registration is not necessary
6 | 2,4,6 | Trolley for transport sterile material |  | 5
| 1 | Trolley for transport of sterile containers and baskets for 6 sterilization unit
| 2 | 2pushing handles
| 3 | Two doors on hinges, lockable
| 4 | Door seal around the whole door and frame
| 5 | Wheels diameter min. 160 mm that do not leave traces, 2 fixed and 2 craft item. brakes on 2 wheels
| 6 | External dimensions max. 980x710x1500 mm
| 7 | ALIMS Registration is not necessary
7 | 2,4,7 | System for tracking and tracing in the system of sterile instruments supply |  | 1
| 1 | A system that provides flow control of sterile instruments or instruments to be sterilized in more positions
| 2 | The system has the ability to make document at each stage of the sterile instruments supply
| 3 | The system is connected via computer to a hospital computer network and it can be supervised in accordance with individual access permissions
| 4 | The system is located at the following locations in the sterile supply chain:
| 5 | o return of instruments from operating theatres
| 6 | o washing and disinfection of instruments
| 7 | o washing and disinfection process control (data are generated automatically by the machine for each batch)
| 8 | o checking and packing of instruments
| 9 | o sterilization
| 10 | o sterilization process control (data are generated automatically by the machine for each batch)
| 11 | o delivery of sterile instruments
| 12 | o the place of use of sterile instruments (operating theatres)
| 13 | System that controls the flow of sterile instruments or instruments that are sterilized in 6 positions: the process of returning used instruments, cleaning and disinfection, checking and packaging, sterilizing, issuing and handing over to the operating room or other user.
| 14 | The system integrates all devices in the chain of sterilization: sterilizers, washer-disinfectors, heat sealing machines
| 15 | The system recognises instruments according to the identification located on the instrument and package (bar code, RFID or the corresponding)
| 16 | The ability to administratively update set data when there is a change in their content by completely ejecting sets or combining new sets.
| 17 | The system indicates a lack of instruments in the chain, as well as the insertion of instruments into the chain that are not found in the database
| 18 | Database with limited access. The modification of data only by authorized persons
| 19 | ALIMS Registration is not necessary
8 | 2,4,8 | General table |  | 2
| 1 | Table for instruments sorting,with shelf, fixed height
| 2 | Frame made from pipes of stainless steel AISI 304
| 3 | Panel and shelves made of stainless steel AISI304
| 4 | Dimensions min. 1800x700x900 mm (LxWxH)
| 5 | ALIMS Registration is not necessary
9 | 2,4,9 | Height adjustable packing table |  | 3
| 1 | Table for packing sterilization materials
| 2 | Electrically adjustable tabletop height by push buttons
| 3 | Possibility of presetting min 4 tabletop heights
| 4 | Tabletop made of AISI 304 stainless steel or better
| 5 | Columns for shelves made of stainless steel or aluminum
| 6 | One drawer on one side under the worktop
| 7 | 2 shelves above the table top, placed between vertical holders
| 8 | Power supply 220V-50Hz
| 9 | Dimensions min. 1800x700x770-1200 mm (WxWxH)
| 10 | ALIMS Registration is not neccessary
10 | 2,4,10 | Instrument washing sink |  | 1
| 1 | Instrument washing sink with two basins
| 2 | Made of stainless steel AISI 304
| 3 | Tap with one point, spiral extension and douche
| 4 | Barrier for liquid splashes
| 5 | Sink drainage control handle
| 6 | Dimension: max. 700x1800x850
| 7 | ALIMS Registration is not necessary
11 | 2,4,11 | Linen inspection table |  | 1
| 1 | Work table with all-clear illumintation part for easy inspection of linen.
| 2 | Illuminated part dim. 1000x800 mm, has min. 4 tubular lights and milky coloured glass. On/off switch on the side of the table.
| 3 | Tubular lights connected to 220V-50Hz power supply
| 2 | The working table for textile control should have neon or LED lighting, for an uninterrupted process of control
| 4 | The worktop made of HPL or better.
| 5 | Frame and shelf made of AISI304 stainless steel or better.
| 6 | Dimensions min. 2000x1200x900 mm (WxDxH)
| 7 | ALIMS Registration is not necessary
12 | 2,4,12 | Paper trolley |  | 2
| 1 | Made of stainless steel AISI 304 or better
| 2 | Min. 4 frames for paper holding with camps for holding paper in place
| 3 | Solid bottom shelf made of stainless steel AISI304 or better
| 4 | 4 castors, 2 of which have brakes
| 5 | Dimensions: 1250x600x900 ±200mm
| 6 | ALIMS Registration is not necessary

## 5.Auxiliary eq.
Bidder:
LOT 2 |  | STERILIZATION |  |  | Date:
Total DAP price:
Line item No. | ID |  | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no.in technical documentation
| 2.5 | Auxiliary equipment for sterilization |  |  |  | 0
1 | 2,5,1 | Sterilization box |  | 60 |  | 0
| 1 | Boxes are intended for storage and transport of non-sterile and sterile materials in CSSD and other wards.
| 2 | Box and lid is made of solid material
| 3 | Lid and the bottom of the box have perforations that allow steam penetration during sterilization process.
| 4 | Perforation should be protected by sterile filter that prevents penetration of potentially infectious air inside the box after sterilization, but allows steam penetration during sterilization process.
| 5 | Box has foldable handles
| 6 | Type: 1 STU
| 7 | Dimensions: approximately 520x280x260 mm
| 8 | Material: Aluminum alloy
| 9 | ALIMS registration not necessary
2 | 2,5,2 | Wire basket, 1 STU |  | 60 |  | 0
| 1 | Baskets are intended for storage and transportation of non-sterile and sterile material
| 2 | Baskets are made of wire mesh;raster prevents standard elements from falling out.
| 3 | Square shape
| 4 | Made of stainless steel AISI 304
| 5 | Type: 1 STU,
| 6 | Dimensions: approximately 600x300x300 mm (LxWxH)
3 | 2,5,3 | Wall brackets |  | 5
| 1 | For hanging varioous modular elements as well as sterilization baskets
| 2 | Made of stainless steel AISI 304
| 3 | With integrated hooks, min 2 hooks/bracket
| 4 | Tubing must be finished with plastic caps on both sides
| 5 | Carriers of 700 mm ± 5% length with interlocking hooks 350 mm ± 5%"
4 | 2,5,4 | Fixed solid shelf system |  | 2
| 1 | Made of stainless steel AISI 304 or better
| 2 | Equipped with 4 shelves
| 3 | The pillars are made of stainless steel tubing and equipped with four leveling feet
| 4 | Dimensions 1520x460x1830 mm ±5%
5 | 2,5,5 | Magnifying lamp |  | 2
| 1 | Stain free quality magnifiers
| 2 | Effortless adjustment without swinging movements
| 3 | Strong shadow free light
| 4 | Lighting with fluorescent lamp or LED, max 30W
| 5 | Distortion-free
| 6 | 4-diopter magnifier min. Ø120 mm
| 7 | Two arms of total length min 800mm
| 8 | Adjustment angle arms: 140°/130°, Rotation 360°
| 9 | El. power 230V/50Hz
6 | 2,5,6 | Spray gun |  | 1
| 1 | The device is intended for washing and cleaning, with the possibility of connection to hot water or compressed air (the same device has both possibilities)
| 2 | Adjustable water or air flow
| 3 | Min. 8 attachments for washing different elements
| 4 | Wall-moounted holder for attachments
| 5 | Connection hose – min. 2 m
| 6 | Spray gun made of stainless steel, aluminum, and similar
7 | 2,5,7 | Overhead shower |  | 1
| 1 | Intended for the first rough cleaning of instruments
| 2 | Activating the device by simply pressing the handle at the shower or pulling the shower to the sink opening
| 3 | Has to be connected to cold and hot water, mounted on the wall.
| 4 | The temperature setting is done by mixing warm and cold water with two taps
| 5 | Dimension: (LxWxH) 260x100x1250 ±5%
8 | 2,5,8 | Ultrasonic cleaning unit |  | 1
| 1 | The device is intended for ultrasonic cleaning of surgical and microsurgical instruments
| 2 | Capacity of the device 25-30 liters
| 3 | Adjustable cleaning temperature and time
| 4 | Acoustic alarm at the end of the cycle
| 5 | Transducer type –piezoelectric, frequency 30-40 kHz
| 6 | Power supply 220V/50Hz, single phase
| 7 | Setting the operating temperature max up to 80 ° C
ALIMS registration is not necessary for equipment 1-8