## Anesthesiology
Bidder: |  | Offered price:
LOT 5 | ANESTHESIOLOGY
ID | Equipment name | QTY | Technical Specification Offered(model) | Total Price per line item
5.1 | Anesthesia system with monitoring | 8
5.2 | DEFIBRILLATORS
5,2,1 | DEFIBRILLATOR AED with ECG monitoring and manual defibrillation option | 4
5,2,2 | DEFIBRILLATOR for HOSPITAL USE – ADVANCED | 2
5,2,3 | DEFIBRILLATOR for HOSPITAL USE – STANDARD | 10
5.3 | Blood and I.V. fluid Warmer | 10
5.4 | Monitoring system
5,4,1 | Advance Monitoring system (ICU 2) | 6
5,4,2 | Advance Monitoring system (ICU 3) | 6
5,4,3 | Advance Monitoring system - CCU /Coronary Care Unit | 7
5.5 | Ventilator with metabolic monitoring | 3
5.6 | Ventilator with standard capnography | 4
5.7 | ECG machines
5,7,1 | ECG machine 12-CHANNEL unit | 4
5,7,2 | Multifunctional ECG device (6/12-CHANNEL) | 5
5,7,3 | High resolution 12-lead ECG unit | 4
5.8 | Central station with WEB remote access to patients’ clinical information | 5
5.9 | ADDITIONAL WORK PLACE FOR ANALYSIS HOLTER RECORDER | 1
5.1 | Monitoring system with advance cardio function – (ICU 1, Post-CCU) | 22
Total Price per Lot

## 1.Anesthesia system
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 5.1 | Anesthesia system with monitoring |  | 8 |  | 0
| 1 | Mobile system with central breaking of wheels;
| 2 | Anesthesia system must have width maximum 105 cm and depth maximum 105 cm
| 3 | Must have minimum two drawers;
| 4 | System console must have work surface in the middle and upper shelf for additional devices and accessories;
| 5 | Besides main power supply must have backup internal battery supply;
| 6 | It has connections for pipeline gas supply with O2, N2O and air and on the back position for placement of back up cylinder. The cylinder should be delivered with machine;
| 7 | Must have safety solution that guarantee in in exceptional circumstances the ventilation can be performed using alternative oxygen supply including undisturbed vaporizer operation;
| 8 | Must have protection from hypoxy (hypoxyguard);
| 9 | Must have integrated suction system;
| 10 | Must have anesthestic gas scavenging system;
| 11 | Must have auxiliary O2 supply (for nasal or mask oxygenation).
| 12 | Must have active ventilation modes:
| 13 | Volume controlled, pressure controlled, pressure controlled with volume guarantee or equivalent, SIMV in variant with pressure and with volume control, pressure supported ventilation (PS);
| 14 | System has ability for lung recruitment either as automated procedure (e.g. mechanical breath delivery with pre-set inspiratory pressure and duration) or on other, alternative way (e.g.manually)
| 15 | It is desirable the system has possibility of delivery several consecutive mechanical breaths with different ventilation parameter settings or at least has possibility of preset ventilation parameters before starting procedure
| 16 | Minimal range of tidal volume adjustment: 20 – 1500 ml;
| 17 | It has ability of placement two vaporisers. Sevoflurane vaporiser should be delivered with machine;
| 18 | It has ability of flow measurement with proximally placed sensor (right in front of airway tube);
| 19 | It has inspiratory and expiratory conc. O2, CO2, N2O and agent via movable gas module, compatibile with acompanying patient monitor;
| 20 | System has its own color display on mobile arm, of size minimum 15”, with resolution of 1024x768 or higher, touchscreen. The display must have a rotary command for menu selection and settings;
| 21 | Display must show following measured data: three parametric curves at the same time (pressure, flow, CO2), respiratory mechanic curve (pressure-volume or flow-volume), airway pressure, breathing rate, tidal and minute volume,compliance, bargraf of flowmeters, digitally showed pipeline pressures, measured conc.of O2, CO2, N2O and anesthetic agent in inspirium and expirium, MAC value;
| 22 | It has continuous calculation with graphical/numerical display of flow range in which user can adjust flow settings according to target FiO2 and maintain oxygenation most economically on that way. This view must have also anesthetic agent consumption expediture expressed.
| 23 | Breathing system should be small, compact, designed for low flow, that uses reusable CO2 absorber canister;
| 24 | It has alarm system with adjustable limits;
| 25 | Acompanying haemodynamic monitor physically mounted on the anesthesia system;
| 26 | The monitor has main power and back up (battery) supply. The battery should allow minimum 1h of work without main power supply;
| 27 | Patient monitor has color LCD display of size minimum 15”, touchscreen;
| 28 | It must display 8 curves on screen;
| 29 | Active measuring parameters: ECG, saturation (SpO2), non invasive pressure, temperatura, invasive pressure x 2, measurement of sensitivity on surgical stimuli during general anesthesia (level of analgesia), status of CNS during anesthesia (anesthesia depth), NMT;
| 30 | It has trend of 24 h;
| 31 | It has adjustable alarm system according to type of patients;
| 32 | Networking ability (LAN, WLAN) and connection with central station;
| 33 | Minimum integrated 2 USB ports for periferal devices connection (keyboard, mouse, barcod reader, remote control etc);
| 34 | Ability of additional ( optional) upgrade on monitor: SvO2 module, C.O.module EEGmodule,

## 2.Defibrillators
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 5.2 | DEFIBRILLATORS |  |  |  | 0
1 | 5,2,1 | DEFIBRILLATOR AED with ECG monitoring and manual defibrillation option |  | 4 |  | 0
| 1 | Semiautomatic-automatic defibrillator with a biphasic defibrillation technology
| 2 | Defibrillator must have a manual mode so that a user chooses a moment for delivery of shock, regardless of the semi-automatic mode
| 3 | Device must have the automatic self test and the ability to visually or audibly warns if there is a specific problem
| 4 | Device must have standard IP54 protection against dust and moisture
| 5 | Analysis of the heart rhythm and the delivery of shocks have to be done through the self-adhesive electrodes where the patient cable for adult patients should not be less than 2m
| 6 | Detection VF / VT must be based on the analysis of ECG signals in semi-automatic mode
| 8 | Device must have the ability of safely discharge after charging if a shock is not delivered, there were a problem with electrodes, or the device has been turned off
| 9 | Device must have the ability to automatically identify the connected electrodes for children and adults and according to that to automaticaly selection adequate energy that should be delivered to the connected patient
| 10 | Defibrillator must have audio and visual instructions for operator in Serbian language
| 11 | Defibrillator must have integrated high resolution display LCD type backlit with minimum dimensions 100 x 35 mm
| 12 | The display must allow continuous monitoring of minimum 1-channel ECG, the number of delivered shocks, memory and battery capacity , type adhesive electrodes connected adults / children and text instructions for operator in Serbian language
| 13 | Defibrillator must have a memory card or internal memory for continuous recording min. 2 hours of continuous ECG or ambient sound
| 14 | Battery used by the defibrillator must be produced, certified and intended for use only in defibrillators
| 15 | Battery capacity minimum 180 shocks with maximum energy or minimum 3.5 hours of monitoring of the patient
| 16 | Device must have the option of giving a sound rhythm for properly performing of CPR
| 17 | Net weight with battery max. 1,6kg
| 18 | Device has to be delivered with the necessary accessories for the operation and instructions for use
2 | 5,2,2 | DEFIBRILLATOR for HOSPITAL USE – ADVANCED |  | 2 |  | 0
| 1 | Biphasic synchronous and asynchronous defibrillator that can be used as a patient monitor
| 2 | Integrated power supply, battery-powered defibrillator with a battery capacity of at least 2 h monitoring.
| 3 | Upgradeable capability with an extra battery with capacity of autonomy for a minimum 4h monitoring
| 4 | Power supply of 220-240V, 50Hz
| 5 | Defibrillation of patients through integrated pedals
| 6 | The pedals for defibrillation have to have integrated buttons to select energy, charge, shock delivery and start / stop function for integrated thermal printer
| 8 | Integrated color LCD monitor with diagonal minimum 10 "
| 9 | The monitor must have adjusting of brightness (brightness)
| 10 | Monitor must have the possibility of simultaneous display of 12 ECG leads
| 11 | Monitor must have the possibility of a trend display of graphic and tabular, 24 hours, for all parameters
| 12 | Monitor must have the ability to adjust the alarm limits for ECG, SpO2 and NIBP
| 13 | Device must have an integrated SpO2 option with the associated sensor
| 14 | Device must have an integrated NiBP option with its accessory
| 15 | The complete menu of user interface software must be in Serbian language
| 16 | Device must have a minimum of integrated following interfaces: 1 x USB, 1 x QRS trigger and 1 x RS 232
| 22 | Upgradeable ECG transfer option that allows the transfer of 12-lead ECG with a GSM or standard modem
| 23 | Upgradeable external pacemaker with FIX-DEMAND and OVERDRIVE modes
| 24 | Weight max. 6 kg with batteries and pedals
| 25 | Defibrillator must contain the complete standard accessories for work
3 | 5,2,3 | DEFIBRILLATOR for HOSPITAL USE – STANDARD |  | 10 |  | 0
| 1 | Hospital defibrillator for manual defibrillation with selectable energy levels of defibrillation
| 2 | Integrated power supply with battery in the defibrillator unit
| 3 | Battery capacity of min. 150 shocks with a maximum energy or 3.5 h of ECG monitoring
| 4 | Device must have an indicator of the charging-discharging of battery
| 5 | Biphasic defibrillation technology
| 6 | Device must have integrated pedals for defibrillation
| 8 | Device must have an integrated display readable in all conditions, minimum diagonal 5.9 " for display of ECG curves, heart rate (HR) and SpO2 values ​​
| 9 | Printing of report on the integrated 3-channel thermal printer with minimum speed of 25mm / s
| 10 | Thermo printer uses paper width minimum 70 mm width
| 11 | Device must have the ability to automatically print after completion of electrical shock and manually on users demand
| 12 | Device must have a minimum of integrated following interfaces: 1 x USB and 1 x RS 232
| 13 | Device memory for a minimum of 12 hours of ECG or 500 events
| 14 | A complete software user interface in Serbian language
| 15 | Device must have the ability to be upgraded with an optional external pacemaker modes Fix, Demand and Overdrive
| 16 | Device must have the ability to be upgraded with option for SpO2 measurement, which starts as soon as the measuring sensor is detected
| 17 | Device must have the ability to be upgraded with an optional semi automatic defibrillation, that uses sound and display instruction in order to guide users through the process of defibrillation
| 18 | Maximum gross weight of defibrillator including battery and pedals 5.5 kg
| 19 | Device must have minimum IP 53 protection standard
| 20 | Defibrillator should come with the following accessories for work:
| 21 | the pedals for defibrillation
| 22 | ECG patient cable
| 23 | 1 pack of gel for defibrillation
| 24 | the power cord
| 25 | the user manual.

## 3.Warmer
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 5.3 | Blood and I.V. fluid Warmer |  | 10 |  | 0
| 1 | Device for infusion solutions and blood heating using recirculating solution or water heating technology;
| 2 | It provides constant temperature of infusion solution or blood in whole system;
| 3 | Outlet temperature adjustment/settings range: 37 – 42 oC;
| 4 | It must heat whole system, form device till infusion/blood – set temperature is the same at end point /output part of the system (prevented the occurrence os sol called Line cool down effects);
| 5 | It has infusion solution or blood flow up to 5 lit/h.

## 4.Monitoring system
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
| 5.4 | Monitoring system |  |  |  | 0
1 | 5,4,1 | Advance Monitoring system (ICU 2) |  | 6 |  | 0
| 1 | Patient monitor with advanced parametar modules
| 2 | Integrated colour display 15 inch with resolution min. 1024 x 768 pixels
| 3 | Controls: touchscreen and rotary control button
| 4 | Number of traces : individual min. 8 or up to 14 with overlays and insets
| 5 | Automatic configuration according to parameter availability.
| 6 | Alarms with audiable and visual notification
| 8 | 72-hour parameter trends
| 9 | Network type (LAN, WLAN) and central station compatibility
| 10 | Integrated 2 USB ports
| 11 | Monitor configuration must include next algorithms and parameters :
| 12 | ECG including
| 13 | o Arrhythmia monitoring in four simultaneous leads , including atrial fibrillation detection and trending,
| 14 | o Continuous ST-segment monitoring of up to 12 leads ,
| 15 | o Continuous Qt/QTc monitoring , including measuremets, trending and alarming , to support in early identification of Torsades de pointes events.
| 16 | Respiration
| 17 | NIBP – non-invasive blood pressure
| 18 | SpO2 - Puls oximetry
| 19 | IBP x 2 -invasive blood pressure 2 channel
| 20 | Temperature
| 21 | Cardiac Output
| 22 | Additional Options and Features:
| 23 | SvO2 module
| 24 | EEG module
| 25 | Measuring the expiratory concentrations of CO2, through a portable module that is compatible with the ventilator device;
| 26 | Monitoring of Gas exchange and metabolics –continuous non-invasive measurement of oxygen consumption (VO2), carbon dioxide production (VCO2) , and calculation of respirtory quotient (RQ) and energy expenditure (EE)
| 27 | Patient spirometry
| 28 | PICCO module
| 29 | Integrated thermal recorder (min 3 channel)
| 30 | Remote controler
| 31 | Monitor Accessories
| 32 | ECG cable (5-leads) 1pc
| 33 | NIBP Cuffs (3 different size) 1pc
| 34 | SpO2 finger sensor –reausable 1pc
| 35 | Temperature probe 1pc
| 36 | IBP cables 1pc
| 37 | CO set 1pc
| 38 | Wall Mounting with bracket for monitor 1pc
2 | 5,4,2 | Advance Monitoring system (ICU 3) |  | 6 |  | 0
| 1 | Patient monitor with advanced parametar modules
| 2 | Integrated colour display 15 inch with resolution min. 1024 x 768 pixels
| 3 | Controls: touchscreen and rotary control button
| 4 | Number of traces : individual min. 8 or up to 14 with overlays and insets
| 5 | Automatic configuration according to parameter availability.
| 6 | Alarms with audiable and visual notification
| 8 | 72-hour parameter trends
| 9 | Network type (LAN, WLAN) and central station compatibility
| 10 | Integrated 2 USB ports
| 11 | Monitor configuration must include next algorithms and parameters :
| 12 | ECG including
| 13 | o Arrhythmia monitoring in four simultaneous leads , including atrial fibrillation detection and trending,
| 14 | o Continuous ST-segment monitoring of up to 12 leads ,
| 15 | Respiration
| 16 | NIBP – non-invasive blood pressure
| 17 | SpO2 - Puls oximetry
| 18 | IBP x 2 -invasive blood pressure 2 channel
| 19 | Temperature
| 20 | Additional Options and Features:
| 21 | Cardiac Output module
| 22 | SvO2 module
| 23 | EEG module
| 24 | Measuring the expiratory concentrations of CO2, through a portable module that is compatible with the ventilator device;
| 25 | Monitoring of Gas exchange and metabolics –continuous non-invasive measurement of oxygen consumption (VO2), carbon dioxide production (VCO2) , and calculation of respirtory quotient (RQ) and energy expenditure (EE)
| 26 | Patient spirometry
| 27 | PICCO module
| 28 | Integrated thermal recorder (min 3 channel)
| 29 | Remote controler
| 30 | Monitor Accessories
| 31 | ECG cable (5-leads) 1pc
| 32 | NIBP Cuffs (3 different size) 1pc
| 33 | SpO2 finger sensor –reausable 1pc
| 34 | Temperature probe 1pc
| 35 | IBP cables 1pc
| 36 | Wall Mounting with bracket for monitor 1pc
3 | 5,4,3 | Advance Monitoring system - CCU /Coronary Care Unit |  | 7 |  | 0
| 1 | Patient monitor with advanced parametar modules
| 2 | Integrated colour display 15 inch with resolution min. 1024 x 768 pixels
| 3 | Controls: touchscreen and rotary control button
| 4 | Number of traces : individual min. 8 or up to 14 with overlays and insets
| 5 | Automatic configuration according to parameter availability.
| 6 | Alarms with audiable and visual notification
| 7 | 72-hour parameter trends
| 8 | Network type (LAN, WLAN) and central station compatibility
| 9 | Integrated 2 USB ports
| 10 | Monitor configuration must include next algorithms and parameters :
| 11 | ECG including
| 12 | o Arrhythmia monitoring in four simultaneous leads , including atrial fibrillation detection and trending,
| 13 | o Continuous ST-segment monitoring of up to 12 leads ,
| 14 | o Continuous Qt/QTc monitoring , including measuremets, trending and alarming , to support in early identification of Torsades de pointes events.
| 15 | o 12 channel ECG gender-specific algorithm for interprtation
| 16 | o Algorithm for acute cardiac ischemia- triage of chest-pain patients with numeric score of probabilityof acute cardiac ischemia (%)
| 17 | Respiration
| 18 | NIBP – non-invasive blood pressure
| 19 | SpO2 - Puls oximetry
| 20 | IBP x 2 -invasive blood pressure 2 channel
| 21 | Temperature
| 22 | Additional Options and features:
| 23 | Cardiac Output module
| 24 | SvO2 module
| 25 | EEG module
| 26 | Measuring the expiratory concentrations of CO2, through a portable module that is compatible with the ventilator device;
| 27 | Monitoring of Gas exchange and metabolics –continuous non-invasive measurement of oxygen consumption (VO2), carbon dioxide production (VCO2) , and calculation of respirtory quotient (RQ) and energy expenditure (EE)
| 28 | Patient spirometry
| 29 | PICCO module
| 30 | Integrated thermal recorder (min 3 channel)
| 31 | Remote controler
| 32 | Monitor accessories
| 33 | ECG cable (max 6-leads cable for 12 channel ECG monitoring and interprtation)1 pc
| 34 | NIBP Cuffs (3 different size) 1 cpt
| 35 | SpO2 finger sensor –reausable 1 pc
| 36 | Temperature probe 1pc 1 pc
| 37 | IBP cables 1cpt 1 cpt
| 38 | Wall Mounting with bracket for monitor 1cpt 1 cpt

## 5.Ventilator
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 5.5 | Ventilator with metabolic monitoring |  | 3 |  | 0
| 1 | Compact device placed on its own mobile console with brakes on all wheels;
| 2 | Pipeline gas supply: O2 and Air;
| 3 | Possibilty of work with one gas supply source;
| 4 | Power supply: main and internal battery;
| 5 | Ventilator uses paramagnetic or ultrasound technology for O2 conc.measurement;
| 6 | It has internal electronic control of nebulizer operation;
| 7 | It has a port for optional auxiliary pressure measurement (tracheal);
| 8 | It has integrated or modular EtCO2 measurement;
| 9 | Ventilation modes: VCV, PCV, PRVC, Bi Level (or BiPAP or DuoPAP), APRV, SIMV-VC, SIMV-PC, SIMV-PRVC, VS, CPAP/PSV, NIV;
| 10 | It has adjustable Back-up ventilation mode;
| 11 | In sponataneous breathing modes (VS, CPAP, NIV) the ventilator will start to deliver pressure controlled mechanical breaths if the rate of patient spontaneous breathing is less then preset minimal rate;
| 12 | It has mode to help patient weaning (SBT or equivalent);
| 13 | It has adjustable compensations for tube resistance, leakage, trigger;
| 14 | It has flow and pressure triggering;
| 15 | Can automatically detect presence of patient, during (re)connection from standby;
| 16 | It has suction consisting of pre-oxygenation phase, temporary pausing alarms and post-oxygenation phase”;
| 17 | It has the routine for measuring patient functional residual capacity with numerical and graphical display of results;
| 18 | It has PEEP titration routine to achieve targeted functional residual capacity and faster patient recovery;
| 19 | It has monitoring of patient metabolic status (e.g. hypermetabolic changes due to infection or nutrition status information) expressed as calculated energy expediture in kCal during ventilation therapy and as respiratory coefficient;
| 20 | Besides ventilator flow sensors it has ability for alternative precise flow measurement with sensors placed right behind the airway tube;
| 21 | Minimal adjustment ranges :
| 22 | o Tidal volume: 20 - 2000 ml
| 23 | o Rate: 3 - 120 bpm;
| 24 | o P insp:1 - 95 cm H2O
| 25 | o P support: 0 - 50 cm H2O;
| 26 | o PEEP:1- 50 cm H2O;
| 27 | o Insp T:0.5 - 10 s ;
| 28 | o Bias flow: 2 -10 lit ;
| 29 | FiO2 :21 - 100%;
| 30 | Color LCD display of minimal size 15 ” (diagonal), touchscreen, with adjustable height and tilt;
| 31 | Ventilation monitoring on display consisting of: 4 breating curves showed at the same time, respiratory mechanic (pressure-volume or pressure-flow or flow-volume), pressure (peak,mean,plato), PEEPe, PEEPi, tidal and minute volumee (inspiratory and expiratory), compliance, resistance, FiO2, EtCO2,;
| 32 | It has procedure for measurement intrinsic (auto)PEEP and PEEPi volume;
| 33 | It measures/calculates parameters: P0.1, RSBI, vital capacity, NIF;
| 34 | It has ability to hold inspirium and to hold expirium;

## 6.Ventilator capn.
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 5.6 | Ventilator with standard capnography |  | 4 |  | 0
| 1 | Compact device placed on its own mobile console with brakes on all wheels;
| 2 | Pipeline gas supply: O2 and Air;
| 3 | Possibilty of work with one gas supply source;
| 4 | Power supply: main and internal battery;
| 5 | Ventilator uses paramagnetic or ultrasound technology for O2 conc.measurement;
| 6 | It has internal electronic control of nebulizer operation;
| 7 | It has a port for optional auxiliary pressure measurement (tracheal);
| 8 | It has integrated or modular EtCO2 measurement;
| 9 | Ventilation modes: VCV, PCV, PRVC, Bi Level (or BiPAP or DuoPAP), SIMV-VC, SIMV-PC, SIMV-PRVC, VS, CPAP/PSV;
| 10 | In sponataneous breathing modes (VS, CPAP, NIV) the ventilator will start to deliver pressure controlled mechanical breaths if the rate of patient spontaneous breathing is less then preset minimal rate;
| 11 | It has Back-up ventilation mode;
| 12 | It has mode to help patient weaning (SBT or equivalent);
| 13 | It has adjustable compensations for tube resistance, leakage, trigger;
| 14 | It has flow and pressure triggering;
| 15 | Can automatically detect presence of patient, during (re)connection from standby;
| 16 | It has suction consisting of pre-oxygenation phase, temporary pausing alarms and post-oxygenation phase;
| 17 | Besides ventilator flow sensors it has ability for alternative precise flow measurement with sensors placed right behind the airway tube;
| 18 | Minimal adjustment ranges :
| 19 | o Tidal volume: 20 - 2000 ml
| 20 | o Rate: 3 - 120 bpm;
| 21 | o P insp:1 - 95 cm H2O
| 22 | o P support: 0 - 50 cm H2O;
| 23 | o PEEP:1- 50 cm H2O
| 24 | o Insp T: 0,5-10 sec
| 25 | o Bias flow: 2 -10 lit ;
| 26 | FiO2 :21 - 100%;
| 27 | Color LCD display of minimal size 15 ” (diagonal), touchscreen, with adjustable height and tilt;
| 28 | Ventilation monitoring on display consisting of: 4 breating curves showed at the same time, respiratory mechanic (pressure-volume or pressure-flow or flow-volume), pressure, tidal and minute volume, compliance, FiO2, EtCO2,;
| 29 | It has procedure for measurement intrinsic (auto)PEEP and PEEPi volume;
| 30 | It measures/calculates parameters: P0.1, RSBI, vital capacity, NIF;
| 31 | It has ability to hold inspirium and to hold expirium;
| 32 | It has ability to take and memorize snapshot, including: 15 sec of all actual curves, active alarm messages, all measured parameters, all ventilation settings;
| 33 | It has 72 h trend of settings and measured parameters;
| 34 | It has alarm system with adjustable limits and auto limit ability.

## 7.ECG
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 5.7 | ECG machines |  |  |  | 0
1 | 5,7,1 | ECG machine 12-CHANNEL unit |  | 4 |  | 0
| 1 | Device must have automatic and manual mode
| 2 | 12 simultaneous channels in standard and cabrera modes
| 3 | Device must have an integrated 12-lead ECG display with minimum dimensions of 120x90 mm.
| 4 | Device must have an integrated 12-channel high resolution thermal printer to print ECG on thermal paper A4 format, minimum width 210 mm
| 5 | Device must have the ability to be directly connected on an external laser printer and to print 12 channels of ECG on A4 standard copy paper
| 7 | Print Speed: 5-10-25-50 mm / s adjustable in manual mode
| 8 | Automatic or manual selection of QRS amplitude signal: 5-10-20 mm / mV
| 9 | Device must have an integrated alphanumeric keyboard with a durable rubber keys. Foil and "touch screen" keyboard is not acceptable.
| 10 | Device must have an integrated power supply and battery with a battery capacity of at least 4 hours of standard usage which involves a minimum of 300 patients recorded in automatic mode with printed ECGs
| 11 | Device must have the QRS beeper for sound monitoring of QRS complexes
| 12 | Device must have the software to measure the ECG signal
| 13 | Device must have an integrated memory for at least of 30 ECG records
| 14 | Device must have a communication module for bi-directional communication with the associated PC software for storage and subsequent analysis of ECG recordings and direct networking option of ECG device to hospital information system.
| 15 | Communication via the integrated module must be enabled by XML standard.
| 16 | Device must have the ability to be upgraded with software option thrombolysis (gives a recommendation to users regarding thrombolytic therapy)
| 17 | Device must have the ability to be upgraded with software option of STRESS TEST
| 18 | Device must have the ability to work with a barcode reader
| 19 | Device must have the ability to be upgraded with an VGA connector in order to get connected to an external LCD display
| 20 | Device has to be delivered with the appropriate standard accessories for work.
2 | 5,7,2 | Multifunctional ECG device (6/12-CHANNEL) |  | 5 |  | 0
| 1 | The appliance must have automatic and manual mode
| 2 | 12 simultaneous sink in modes standard and cabrera
| 3 | The appliance must have an integrated 12-lead ECG display with minimum dimensions of 120x90 mm.
| 4 | The appliance must have an integrated 12-channel high resolution thermal printer to print ECG on thermal paper A4 format, minimum width 210 mm
| 5 | The appliance must have the ability to directly connect an external laser printer and press 6 and 12 channel ECG on A4 copy paper
| 6 | Print Speed: 5-10-25-50 mm /s adjustable in manual mode
| 7 | Automatic or manual selection of QRS amplitude signal: 5-10-20 mm/mV
| 8 | The appliance must have an integrated alphanumeric keyboard command with a durable rubber keys. Foil and "touch screen" keyboard is not acceptable.
| 9 | The appliance must have an integrated current battery power with a battery capacity of at least 4 hours of standard usage which involves a minimum of 300 patients performed in automatic mode with print ECG
| 10 | The appliance must have the QRS beeper for sound recording QRS complex
| 11 | The appliance must have an integrated memory for a minimum of 30 ECG records
| 12 | The appliance must have software for measurement and interpretation of ECG
| 13 | The appliance must have the thrombolysis software (recommended for users giving thrombolytic therapy)
| 19 | The appliance must be registered in ALIMS (Agency for Medicines and Medical Devices) Serbia
3 | 5,7,3 | High resolution 12-lead ECG unit |  | 4 |  | 0
| 1 | Device must have automatic and manual mode
| 2 | 12 simultaneous channels standard, Cabrera and Nehb modes
| 3 | Device must have an integrated 12-channel ECG display diagonal of 8.4 inches minimum
| 4 | Device must have an integrated 12-channel high resolution thermal printer to print ECG
| 5 | on thermal paper A4 format, minimum paper width 210 mm
| 6 | Device must have a minimum sampling rate of 4000 Hz
| 7 | Printing speed is adjustable in manual mode: 5-10-25-50 mm / s
| 8 | Automatic or manual selection of QRS amplitude signal: 5-10-20 mm / mV
| 9 | Device must have an integrated alphanumeric "hospital-grade" keyboard, that is easy to clean and resistant to water splashes and dust.
| 10 | Device must have an integrated power supply with a battery that has a capacity of at least 3 hours.
| 11 | Device must have the QRS beeper for sound monitoring of QRS complexes
| 12 | Device must have the software for the measurement and interpretation of ECG signals.
| 13 | Device must have the thrombolysis software (that gives recommendation to users regarding thrombolytic therapy)
| 14 | Device must have an integrated Print preview option
| 15 | Device must have the ability of minimum 5-minute, continuous ECG recording for all 12channels and subsequent display.
| 17 | Device must have an integrated memory for at least of 350 ECG records
| 18 | Device must have a communication module for bi-directional wireless and wired communication withthe associated PC software for storage and subsequent analysis of ECG recordings and for direct network integration of ECG machine with hospital information system.
| 20 | Communication via the integrated module must be enabled by the PDF, XML and DICOM standards.
| 21 | Device must have the ability of software upgrade with software option "Hyper Q" or equivalent
| 22 | software that uses the automatic analysis of the recorded high frequency QRS segment, detectschanges in myocardium depolarization and thus provides the doctor clear information on the existence of Acute Coronary Syndrome.
| 24 | Device must have an integrated WLAN 802.11g module
| 25 | Device must have a minimum of the following interfaces: 1xRS-232, 1x USB, 1xRJ -45.
| 26 | Device must have the ability to be upgraded with a barcode reader
| 27 | Device must have the ability to be upgraded with spirometry option
| 28 | Device is supplied with the appropriate standard accessories for work

## 8.Central monitoring
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 5.8 | Central station with WEB remote access to patients’ clinical information |  | 5 |  | 0
| 1 | Central station for ICU monitors
| 2 | Each system consist 2 medical grade displays min 19 inch, keyboard ,optical mouse and laser printer.
| 3 | Single viewer display format: min 8 waveforms per patient.Multy viewer display format: min. 4 waveforms per patient and 1-16 individual patient in standard display form
| 4 | Display information : Acess to patient demographic information , arrhythmias, ST-limit violations,and ECG samples in event review, parametar data in graphic and numeric trends, full disclosure data 24 h , parameter numerics and waveforms available in a live view presentation.
| 5 | Central station must provide monitor control: remote controle of admit/discharge settings, parameter priority , alarm and patient monitor settings.
| 6 | Alarm notification: audible, visual Patient alarms-4 configurable levels complies with IEC standard
| 7 | Active WEB remote access to patients’ clinical information in near real-time, enabling to closely observe a patient’s condition while clicicians are away from the bedside. It operates on a standard PC, Mac or tablet PC, enabling to quickly check the status of patients directly from PC without making a trip to the bedside .The multi-patient view capability supports viewing of up to 16 patients simultaneously.
| 8 | Additional Options and features: Full Disclosure and ST review records min. 6 days

## 9.Add for ex.holter
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 5.9 | ADDITIONAL WORK PLACE FOR ANALYSIS HOLTER RECORDER |  | 1 |  | 0
| 1 | Additional work place for analysis of the current EKG holter recorders SCHILLER AR-12plus and recorders for blood pressure measuring SCHILLER BR-102plus
| 2 | Additional work place needs the network connectivity with the existing system for the analysis of Holter SCHILLER Medilog DARWIN recorders to enable users that simultaneously two operators can use the same software analysis
| 3 | Additional work place must be equipped with the same software features and use identical database patients and carried out tests as well as the current system.
| 4 | With additional software infrastructure, workplace is supplied with the appropriate PC, 2 TFT color monitor and licensed Windows operating system.

## 10.Monitoring system advance
Bidder:
LOT5 |  | ANESTHESIOLOGY |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 | 5.1 | Monitoring system with advance cardio function – (ICU 1, Post-CCU) |  | 22 |  | 0
| 1 | Patient monitor with advanced parametar modules
| 2 | Integrated colour display 15 inch with resolution min. 1024 x 768 pixels
| 3 | Controls: touchscreen and rotary control button
| 4 | Number of traces : individual min. 8
| 5 | Automatic configuration according to parameter availability.
| 6 | Alarms with audiable and visual notification
| 7 | 72-hour parameter trends
| 8 | Network type (LAN, WLAN) and central station compatibility
| 9 | Integrated 2 USB ports
| 10 | Monitor configuration must include next algorithms and parameters :
| 11 | ECG including
| 12 | o Arrhythmia monitoring in four simultaneous leads , including atrial fibrillation detection and trending,
| 13 | o Continuous ST-segment monitoring of up to 12 leads ,
| 14 | o Continuous Qt/QTc monitoring , including measuremets, trending and alarming , to support in early identification of Torsades de pointes events.
| 15 | o 12 channel ECC gender-specific algorithm for interprtation
| 16 | o Algorithm for acute cardiac ischemia- triage of chest-pain patients with numeric score of probabilityof acute cardiac ischemia (%)
| 17 | Respiration
| 18 | NIBP – non-invasive blood pressure
| 19 | SpO2 - Puls oximetry
| 20 | IBP -invasive blood pressure (one channel)
| 21 | Temperature
| 22 | Additional Options and features:
| 23 | Cardiac Output module
| 24 | SvO2 module
| 25 | EEG module
| 26 | Measuring the expiratory concentrations of CO2, through a portable module that is compatible with the ventilator device;
| 27 | Monitoring of Gas exchange and metabolics –continuous non-invasive measurement of oxygen consumption (VO2), carbon dioxide production (VCO2) , and calculation of respirtory quotient (RQ) and energy expenditure (EE)
| 28 | Patient spirometry
| 29 | PICCO module
| 30 | Integrated thermal recorder (min 3 channel)
| 31 | Remote controler
| 32 | Monitor accessories
| 33 | · ECG cabel (max 6-leads cable for 12 channel ECG monitoring and interprtation) 1 pc
| 34 | · NIBP Cuffs (3 different size) 1 cpt