## Price Shedule
Bidder: |  |  |  | Offered price (currency):
LOT 1 | Angio operating theatre
ID | Technical Specification Requested | QTY | Technical Specification Offered (model) | Unit price | Total price per line item
1.1 | Digital biplane angiography system for cardiology diagnostic and interventional pocedures | 1 |  |  | 0
1.3.1 | Cardiovascular ultrasound machine | 1 |  |  | 0
1.3.2 | Ultrasound cardiovascular premium 4D system | 1 |  |  | 0
1.4 | Non-invasive hemodynamic monitoring system | 1 |  |  | 0
1.5 | ECG | 8 |  |  | 0
1.6 | Defibrillator with pacemaker option | 4 |  |  | 0
1.7 | Hospital Defibrillator | 3 |  |  | 0
|  |  |  | Total Price per Lot | 0

## Technical specifications
Bidder:
LOT 1 |  | Angio Operating Theatre |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | Unit Price DAP | Total Price per line item | Insert page no. in technical documentation
1.1 |  | Digital biplane angiography system for cardiology diagnostic and interventional pocedures |  | 1
| 1 | Biplane C-arm stand
| 2 | Monoplane floor and monoplane ceiling mounted C-arms, easy for operating for all kind of examinations and interventions, enabling cranio-caudal angulations, RAO/LAO and lateral projections in both planes (plane A and plane B)
| 3 | Variable focal spot-to-detector distance SID min. 95-120 cm for both planes (plane A and plane B)
| 4 | System functionality enables simple access to the patient during catheter interventions (by using femoral, brachial or radial access) positioning C-arm to left and right patient side, as well postioning C-arm in patient head-end position
| 5 | All movements of the system, including bi-plane C-arm, patient table and system for image acquisition, have control at the patient table side. Joysticks at command panel are above panel plane or equivalent solution at command panel
| 6 | C-arm memory positions can be programmable and can be activated with control panel at the patient table side
| 7 | Setting of following parameters in C-arm memory positions: cranio-caudal angulations, RAO/LAO, SID, patient table height
| 8 | Automatic C-arm/stand positioning depending on the reference image selected
| 9 | Software collision protection to prevent collision between C-arm and patient
| 10 | Patient table
Floor mounted patient table
| 12 | Telescoping column with motorized height adjustment, lowest tabletop position measured from floor level maximum 80 cm
| 13 | Free-floating flat tabletop in 4 directions, made of low absorption material, table length min. 260 cm
| 14 | Table rotation min. +/- 90°
| 15 | Maximum table load min. 320 kg
| 16 | Control panel or Touch screen , at the patient table side comprises: a) Motorized vertical table movement b) Blockage and release of the tabletop c) Motorized movement of the biplane C-arm for reaching of the desired position
| 17 | Touch screen control panel, at the patient table side comprises all operational functions, which includes following:
| 18 | image display adjustment
| 19 | image zoom in/out
| 20 | fluoroscopy level selection
| 21 | acquisition protocol selection
| 22 | field of view selection
| 23 | coronary analysis with stenosis quantification (QCA)
| 24 | Left ventricular analysis (LVA) in one plane and in both planes (plane A and plane B)
| 25 | image-quality optimized algorithm to improve the visibility of the deployed stent during cardiac interventions
| 26 | control of functions from hemodynamic system.
| 27 | control of functions from diagnostic workstation
| 28 | Standard accessories to the patient table includes following:
| 29 | Assembly armrest support
| 30 | Radiolucent foam mattress for patient table for comfortable position of the patient during interventions
| 31 | Table rails for equipment attach
| 32 | Infusion bootle holder
| 33 | Footswitch for fluoroscopy and acquisition for both planes (plane A and plane B)
| 34 | Two HV Generators
| 35 | Two microprocessor-controlled high-frequency X-ray generators with automatic dose rate control for fluoroscopy and acquisition
| 36 | For each of generators max. power min. 100 kW
| 37 | Generators enables protection of X-ray tube anode overheating
| 38 | Generators enables continuous display of actual anode heat (HU status), with display of free capacity left to reach maximum anode heat content OR an equivalent solution, only when there is a warning involved
| 39 | Maximum mA value in pulsed fluoro mode min. 200 mA with small X-ray focal spot for each X-ray tube
| 40 | Two X-ray tubes and collimators
| 41 | Focal spots, minimum 2 for each X-ray tube
| 42 | Size of smallest focal spot must be equal or smaller than 0,4 mm for each X-ray tube
| 43 | Maximum anode heat content min. 3 MHU for each tube
| 44 | /
| 45 | System must be equipped with manual rectangular blade for each X-ray tube
| 46 | System must be equipped with semi-transparent blade, which can be moved separately from the rectangular blade for each X-ray tube
| 47 | System must be equiped with min.three-level adaptive Cu prefiltration for each X-ray tube for X-ray beam quality improvment and dose reduction, automatic selection control based on the real-time patient dose absorption according to C-arm angulation, without interruption aquisition or changing existing organ program
| 48 | Virtual (radiation free collimation) on LIH
| 49 | System must have effective way to control skin dose. If the accumulated reference air kerma exceeds a configures threshold, a warning sound is given and/OR pop-up displays on the system, so operator must change existing C-arm position
| 50 | Both X-ray tubes must have ‘’grid-switch’’ or eqv. tehnology for dose reduction during pulsed fluoroscopy
| 51 | AEC - Automatic exposure control – min. 3 parameters automatically optimized in real time (kV, mA, ms) during different C-arm angulation, without need to change program or to interrupt exposure, which enables minimized dose and optimized resolution independent of C-arm angulation
| 52 | Two Flat panel detectors
| 53 | Active detector size min. 17x17cm, but not more than 30x30cm with rotation in landscape/portrait position for each detector, in case the detector is not square
| 54 | Pixel size maximum 200 µm for each detector
| 55 | DQE, min. 75% at 0 Lp/mm for each detector
| 56 | Input fileds min. 4 for each detector
| 57 | Digitization depth min. 14 bits for each detector
| 58 | Removable grid
| 59 | Acquisition matrix size 1024 x 1024 with highest speed of 30 fps
| 60 | Synchronized rotation of detector and collimator in all C-arm positions OR equivalent solution in order to get always upright live image, useful with radial approach
| 61 | Digital system
| 62 | Digital pulsed fluoroscopy, with minimum 3 selectable pulse rates: 7.5, 15 and 30 fps and with min. 3 additional pulse rates less than 7.5 fps
| 63 | LIH (Last Image Hold)
| 64 | Storage and display of dynamic fluoro sequences (Fluoro Loop or eqv.)
| 65 | Coronary analysis with stenosis quantification (QCA)
| 66 | Left ventricular analysis (LVA) in one plane and in both planes (plane A and plane B)
| 67 | Image-quality optimized algorithm to improve the visibility of the deployed stent during cardiac interventions
| 68 | ECG-triggered fluoroscopy/acquisition provides a still image of the catheter while compensating for cardiac movement
| 69 | Digital cardiac aquisition, with minimum 3 selectable pulse rates: 7.5, 15 and 30 fps and with pediatric option with 60 fps
| 70 | Low dose fluoroscopy and acquisition protocol
| 71 | Display of ECG waveform during aquisition on live image monitor
| 72 | System must have all available features for offered model for dose reduction with improved image quality at the same time (CARE&CLEAR, DOSEWISE&CLARITY IQ or eqv.)
| 73 | Acquisition console with two (2) min. 19” monitor in the control room for display of live image in both planes (plane A and plane B) and for patient data entry
| 74 | Image storage capacity min. 100.000 images, 1024 x 1024 pixels, 10 bits
| 75 | CD or DVD drive for automatic image storage in DICOM format (incl. DICOM viewer)
| 76 | Network DICOM 3.0 protocol:
| 77 | DICOM Storage
| 78 | DICOM Storage Commitment
| 79 | DICOM Query/Retrieve
| 80 | DICOM Modality Worklist
| 81 | DICOM Radiation Dose Structured Report
| 82 | Ceiling mounted suspension system with large LCD or TFT monitor min. 55 inch with freely selectable layout and with minimum 8 external signal inputs:
| 83 | for live image display (plane A and plane B)
| 84 | for reference image display (plane A and plane B)
| 85 | for display of images from other diagnostic systems (CT, MR, etc.) and with possibility of live (real-time) image display from other diagnostic systems (ultrasound, IVUS, OCT, EP systems)
| 86 | Intercom system for audio communication beetwen control and examination room
| 87 | Hemodynamic system
| 88 | System for complete hemodynamic recording with patient database for all kind of exhaminations with b/w laser printer, fully integrated with angiography system
| 89 | CD or DVD drive for archiving
| 90 | System must have following modules:
| 91 | 12 channel ECG (10 elektrodes)
| 92 | Invasive blood pressure (IBP), min. 4 channels
| 93 | SpO2
| 94 | CO (Cardiac Output)
| 95 | Non-Invasive blood pressure (NBP)
| 96 | Automated calculation of gradients, shunts, valve areas and resistances support all kinds of right heart and left heart diagnostic and interventional procedures
| 97 | One (1) monitor min. 19`` in control room and display provided on Large monitor in examination room
| 98 | Virtual pullback and pullback sequence
| 99 | Diagnostic workstation WITH ADVANCED 3D VISUALISATION
| 100 | Independent or integrated workstation with advanced tools for interventional procedures analyzes with imaging ability on large monitor in examination room
| 101 | One or more diagnostic monitors in control room of 19 inch (or more) and with 1,3 MP (or more)
| 102 | Processes volume datasets from angio system and various modalities such as CT, MR and provides MIP, MPR, VRT functions
| 103 | 3D rotational angiography reconstructs a high-contrast 3D image visualizing coronary vessels
| 104 | 3D rotational angiography with cone-beam reconstruction algorithms enables reconstruct of low-contrast 3D image, visualizing ventricles and vessels of the heart. (so called CT-like imaging)
| 105 | 3D/3D fusion spatially aligns two 3D volumes from angiography system or different modality CT/MR in such way that the anatomical structures overlay each other. With 2D/3D fuison only two fluoro projections are required to easily fuse 3D volumes from other imaging modalities for live image guidance.
| 106 | 3D Roadmap superimposes the 3D image color-coded on top of the live X-ray. Any change of the angulations of the C-arm or any change of the table position will cause the workstation to recalculate in real time the view on the 3D image to exactly match the view of the live 2D X-ray image. Without additional contrast agent injection, the physician can observe device movements simultaneously with the 3D overlay of the vessel contours in the fluoroscopy image.
| 107 | 3D application for electrophysiology guidance: A software module that supports electrophysiologists in the planning, execution and follow-up examination of atrial fibrillation treatments using automatic one-click segmentation, especially of the left atrium with display of pulmonary veins, but also of other cardiac chambers/vessels based on data acquired by 3D rotational angiography, CT or MR for planning of atrial fibrillation procedures in 3D. With help of image fusion function 3D-images acquired by CT or MR can be used for 3D Roadmap. The 3D-segmentation result and the ablation points can be used for superimposition via 3D Roadmap and to be transferred to electroanatomical mapping systems.
| 108 | CD/DVD device for image archiving in DICOM format
| 109 | Networking, connecting, archiving according DICOM 3.0 protocol: DICOM Storage, DICOM Storage Commitment, DICOM Query/Retrieve
| 110 | Additional equipment
| 111 | Radiation protection: Ceiling track mounted overhead suspended acrilyc shield 0.5 mm equivalent lead for effective protection to the upper body and x-ray protective curtains 0.5 mm equivalent lead mounted on table for lower body protection, 2 pcs. for left and right table side
| 112 | Ceiling mounted LED examination lamp with light intensity min. 70.000 Lux
| 113 | Two-piece wrap around protection consisting of vest and skirt for radiation protection of medical personnel, 0.5 mm equivalent lead – 4 pcs
| 114 | One piece double sided apron for radiation protection of medical personnel, 0.5 mm equivalent lead all-around protection – 4 pcs
| 115 | Thyroid radiation protection, 0.5 mm equivalent lead – 8 pcs
| 116 | Protection glasses, 0.5 mm equivalent lead frontal and lateral protection– 4 pcs
| 117 | Angio injector
| 118 | Other requirements for bidder
| 119 | To make location project based on space and room project where system will be installed.
| 120 | Installation of the offered system- “Turn key” project (preparation of facilities- technical room, examination room and control room. HV cable from HV substation to technical room has to be provided by Beneficiaries).
| 121 | Warranty min. 12 months from installation date
| 122 | 10 days on site training for staff (doctors and technicians) who will work on the system - 5 days after installation and 5 days during warranty time. Training has to be done by certified Application specialist.
1.3.1 |  | Cardiovascular ultrasound machine |  | 1
| 1 | GENERAL FEATURES
| 1.1 | Stationary ultrasound system on wheels with central brake, portable concept is not acceptable
| 1.2 | LCD monitor size min 23" ±1" Full HD on articulating arm
| 1.3 | Height adjustable and rotating control console
| 1.4 | Integrated QWERTY keyboard
| 1.5 | LCD touch - screen panel size min 12"
| 1.6 | System supports convex, linear, PA sector, matrix, intraoperative, 4D TEE an 4D electronic probes
| 1.7 | Minimum 4 active transducer ports
| 1.8 | At least 8 TGC sliders on operating console or LCD touch - screen panel for control of gain according to depth
| 1.9 | /
| 1.10 | Maximum display depth in B mode minimum 44 cm
| 2 | OPERATING MODES
| 2.1 | System supports following operating modes: B- mode, M-mode, Color Doppler, Power Doppler, HPRF PWD, Color M-mode, Tissue Doppler Imaging, CW Doppler
| 2.2 | Advanced technique for better visualization of low flow: B-flow, Advanced Dynamic Flow, Clarify or equivalent. Power Doppler is not acceptable
| 2.3 | Anatomical M mode
| 2.4 | Advanced technique for one button image optimization in B mode: TEQ, iScan, QuickScan, ATO or equivalent
| 2.5 | Automatic one button base line and velocity ranges in PW/CW Doppler optimization
| 2.6 | 4D acquisition and heart display in gray scale with color maps
| 2.7 | 4D acquisition and heart display in Color Doppler
| 2.8 | Simultaneous visualization in two planes in real time
| 2.9 | Automatic ejection fraction measurement
| 2.10 | Quantitative tool for systolic function assessment, based on speckle-tracking method, with display of segmented myocardial movement (left ventricle) as bull-eye graphic in 2D and 3D (automatic measurement in 2D)
| 2.11 | /
| 2.12 | Triplex mode with PW or CW Doppler
| 2.13 | Stress echo package
| 2.14 | Simultaneous display of last exam and current exam image. Last exam can be from the same ultrasound machine or other modalities, including CT
| 2.15 | Ultrasound system can be upgraded with fusion imaging which enables display of the same anatomical section on ultrasound live image as well as on last exam image performed on CT or MR systems. Image Fusion functionality should be available on 2D transthoracic matrix cardiac probe
| 3 | OTHER FUNCTIONALITIES
| 3.1 | Cine memory minimum 800 MB
| 3.2 | Prospective and retrospective clips recording
| 3.3 | Image storage in DICOM format on HDD, DVD/CD, USB Flash
| 3.4 | Image export in JPEG format and cine clips in AVI (MPEG-4) format
| 3.5 | /
| 3.6 | Device has minimum 3 USB ports of which at least one is on display or on operating panel
| 3.7 | System has following DICOM functionalities: Store, Print, Query/Retrieve, Verification, Modality Worklist Management, MPPS, Structured reporting,
| 4 | TRANSDUCERS
| 4.1 | 2D transthoracic matrix cardiology probe with field of view of minimum 120° and frequency range of 2,0 to 5,0 MHz +/- 1 MHz
| 4.2 | 4D transthoracic matrix cardiology probe frequency range of 2,0 to 5,0 MHz or wider
| 4.3 | 2D transthoracic pediatric cardiology probe with frequency range of 3,0 to 7,0 MHz or wider
| 4.4 | 2D transthoracic neonate cardiology probe with frequency range of 4,0 to 9,0 MHz or wider
| 4.5 | 2D transoesophageal pediatric probe
1.3.2 |  | Ultrasound cardiovascular premium 4D system |  | 1
| 1 | Digital color doppler 4D ultrasound system, not portable, with wheels. Front and rear handles for easy transport.
| 2 | Advanced beamforming technology, software based, C-Sound or similar with dynamic range minimum 290Db.
| 3 | Frame rate more than 800fps
| 4 | Four active probe ports, not including pencil ports.
| 5 | Min 22"±1" wide screen High-Definition (HD) OLED display with monitor on an articulating arm. Resolution 1920x1080 px
| 6 | Integrated Ultra HD LCD touch-screen panel, diagonal at least 10”.
| 7 | Floating operating panel, with integrated alphanumeric backlight keyboard, adjustable in three dimensions: Height, Rotation, Extension.
| 8 | Integrated HDD
| 9 | Maximum depth range in B-mode not less then 45 cm (+/- 15 cm) – probe specific
| 10 | Width range in 2D mode 10-120°
| 11 | Transducer Types:
| 11.1 | Convex array probes
| 11.2 | Micro convex array probes
| 11.3 | Linear array probes
| 11.4 | Sector phased array probes
| 11.5 | TEE probes
| 11.6 | Sector matrix phased array probes
| 11.7 | 4D transthoracic matrix phased array
| 11.8 | 4D TEE volume matrix array probes
| 12 | Operating mode
| 12.1 | B-Mode ( 2D tissue mode )
| 12.2 | M-Mode
| 12.3 | Anatomical M-mode
| 12.4 | /
| 12.5 | Color Flow Mode (CFM)
| 12.6 | 2D Virtual Apex or similar
| 12.7 | Color Angio Power Doppler Imaging (PDI)
| 12.8 | PW Doppler with High PRF
| 12.9 | TVI, Tissue Tracking
| 12.10 | CWD
| 12.11 | 4D Tissue
| 12.12 | 4D Color flow
| 13 | Digital imaging technique that provides real-time visualization of vascular hemodynamics by directly visualizing blood reflectors and presenting this information in a grayscale display as B-flow or similar
| 14 | An advanced image processing technique to remove speckle in real-time examining. As UD clarity and UD speckle reduce imaging or similar.
| 15 | Multiple-angle compound imaging in real time which reduce angular dependence of border or edge as compared to no-compound imaging as Compound Imaging, Real time sono CT, SieClear or similar.
| 16 | Automatic Optimization of B-mode image and spectral image with one key touch as TEQ, Iscan, AO or similar
| 17 | Automatic Spectrum Optimization – provides a single push, automatic, real-time optimization of PW or CW spectrum scale and baseline display. ASO or similar
| 18 | CW Doppler Triplex mode (simultaneous display of 2D, Color Flow and CW Doppler image in real time )
| 19 | Automated Function Imaging which allows assessment at a glance by combining longitudinal views into one comprehensive bulls-eyeview
| 20 | /
| 21 | Software which, depending on the given parameters, allows a volume rate of min. 800fps on 4D TEE probes in single beat acquisition, without reducing image quality Vmax or similar.
| 22 | Simultaneous visualization of min two plane acquisition.
| 23 | Simultaneous display of three independent random slices through the 4D volumen (tissue and color) as FlexiSlice, Qlab or similar
| 24 | Automated measurement of LV volume and EF from volumetric data as 4D Auto LVQ or similar.
| 25 | Simultaneus display of combined of short-axis and long-axis standard slices extracted from the volume data (tissue and/or color)
| 26 | Transducers:
| 26.1 | 2D Sector pediatric cardiac probe. Bandwidth from 3 to 8.0 MHz (+/-1 MHz) or wider. Depth of field: minimum 16cm +/- 1 cm. Sector width: minimum 110 °( +/- 20°)
| 26.2 | 2D Sector pediatric cardiac probe. Bandwidth from 4.0 to 12.0 MHz (+/-1 MHz) or wider. Depth of field: minimum 12 cm+/- 1 cm. Sector width: 100 ° ( +/- 20°)
| 26.3 | 4D Volume transthoracic matrix array cardiac transducer for three-dimensional imaging ( matrix technology provides probe elements in several rows in matrix order ). Bandwidth from 1.5 MHz to 5 MHz (+/- 0,5 MHz). Sector width: minimum 90°
| 26.4 | 4D Volume Transesophageal (TEE) matrix array probe transducer for real-time three-dimensional imaging (matrix technology provides probe elements in severel rows in matrix order). Bandwidth from 3 MHz to 8 MHz (+/- 0,5 MHz) or wider
| 26.5 | 2D convex probe, Bandwidth from 1 to 5 MHz (+/- 1 MHz) or wider. Width: min 70 °( +/- 10°)
| 27 | Possible to upgrade system with operating modes and probes
| 27.1 | /
| 27.2 | /
| 27.3 | The semi-automatic surface detecting algorithm helps clinical users get fast, reproducible and accurate 4D quantification with 4D TTE probes
| 27.4 | Automated EF measurement tool based on 2D-speckle tracking algorithm.
| 27.5 | ECG cable
| 27.6 | Thermal B&W printer
1.4 |  | Non-invasive hemodynamic monitoring system |  | 1
| 1 | Non-invasive hemodynamic monitoring system with tilt table testing, integrated non-invasive continuous "beat to beat" blood pressure measuring and non-invasive measurement of Cardiac Output.
| 2 | complete hemodynamic assessment of the patient integrated in one device
| 3 | autonomic function assessement
| 4 | monitoring system must enable the following measurement methods:
| 5 | • impedance cardiography ICG
| 6 | • continuous "beat to beat" non-invasive blood pressure measurement (finger pletism)
| 7 | • oscillometric blood pressure measurement
| 8 | • 6-channel ECG
| 9 | monitoring system should measure, calculate and display minimum following parameters:
| 10 | • RRI (RR interval)
| 11 | • HR (heart rate)
| 12 | • Blood pressure (systolic, diastolic and middle)
| 13 | • PP (pulse pressure)
| 14 | • Zo (impedance of the thorax derived from the ICG signal)
| 15 | • LVET (left ventricular ejection time)
| 16 | • PEP (pre-ejection period)
| 17 | • SV (stroke volume)
| 18 | • SI (stroke index)
| 19 | • CO (cardiac output)
| 20 | • CI (cardiac index)
| 21 | • TPR (total peripheral resistance)
| 22 | • TPRI (total peripheral resistance index)
| 23 | • TFC (thoracic fluid content)
| 24 | • IC (index of contractility)
| 25 | • ACI (acceleration index)
| 26 | • STR (systolic time ratio)
| 27 | • ER (ejection rate)
| 28 | • LVWI (left ventricular work index)
| 29 | • MSER (mean systolic ejection rate)
| 30 | • REP (rapid ejection rate)
| 31 | • HIrev (revised heather index)
| 32 | • Power spectrum analysis of Heart Rate and Blood Pressure Variability (HRV anf BP)
| 33 | • BRRS (baroreceptor reflex sensitivity)
| 34 | the application diagnostic software should enable the following analysis, review and printing of data:
| 35 | • "beat to beat" review of the trend of hemodynamic parameters
| 36 | • heart rate and blood pressure variability/spectral analysis, display of results and graphics in 2D and 3D form.
| 37 | • baroreceptor sensitivity/sympathetic and parasympathetic activity
| 38 | • complete graphical and numerical "on-line" simultaneous monitoring of all active parameters
| 39 | • the possibility of free marking in the report of certain moments of the test
| 40 | • the possibility of free definition of the pressure measurement interval
| 41 | • automatic calculation of all values ​​during the test
| 42 | continuous non-invasive blood pressure measurement / min. measuring range 50-250 mm/Hg
| 43 | oscillometric blood pressure measurement / min. measuring range 50-250 mm/Hg
| 44 | ECG / sampling rate min.1000 Hz
| 45 | non-invasive hemodynamic monitoring system must be delivered with all-in-one PC system ( min 21", touch screen), licensed WINDOW sofware and appropriate trolley for complete system.
| 46 | system must be delivered with the fully compatible tilt table and all necessary accessories
| 47 | system must have complete software-hardware support for tilt table testing
| 48 | complete set of accessories and 120 sets of eletrodes should be supplied with the device.
1.5 |  | ECG |  | 8
| 1 | 12-channel, high resolution ECG device
| 2 | Integrated mains and battery supply with battery capacity of minimum 2 hours
| 3 | Sampling frequency, min. 30 KHz
| 4 | Frequency range min. 0-250 Hz
| 5 | Input impedance min. 50 MΩ
| 6 | CMRR > 90dB
| 7 | Defibrillation protection min. 5000 VDC
| 8 | ECG amplifier resolution max 1µV and min.24 bit
| 9 | Pacemaker detection frequency min 20 kHz
| 10 | Device must have an integrated color LCD, size not less than 8 inches with the minimum resolution 1024 x 768.
| 11 | Device must have lead attachment quality check.
| 12 | Device must have lead reversal detection
| 13 | /
| 14 | Integrated memory for a minimum of 350 ECG recordings and 50 resting rhythms.
| 15 | Integrated USB interface for export of ECG recordings directly to USB stick
| 16 | Device must support minimum standard and cabrera recording functions
| 17 | Device must support minimum Standard 12-leads, Standard C4r, Pediatric, Right precordial, Left posterior and Nehb lead configurations
| 18 | Recording function must support minimum following QT correction formulas: Bazett, Fredericia, Fremingham, Hadges
| 19 | Device must have a computerized measurement program and interpretation of ECG recordings dedicated for adult and children
| 20 | Must have ECG interpretation based on the "Seattle criteria"
| 21 | Device must have ECG resting rhythm function with full disclousure, review, store and print functions.
| 22 | ECG resting rhythm must be recorded
| 23 | Integrated 6/12-channel high resolution A4 thermal printer for thermal paper min width of 210 mm.
| 24 | Printer must be compatible with z-folded thermal paper
| 25 | Speed min. 12,5-25-50 mm/s.
| 26 | Sensitivity min: 5-10-20 mm/mV.
| 27 | Device must have an sealed alphanemeric keyboard with durable rubber/silicon buttons and direct function keys.
| 28 | Device must have integrated interfaces: min. 2xUSB, LAN and WIFI modul
| 29 | Device must support bidirectional communication with HIS systems.
| 30 | /
| 31 | The complete device software must be in Serbian or English
| 32 | Weight of the device, maximum 5kg
| 33 | Standard accessories should be supplied with the device.
1.6 |  | Defibrillator with pacemaker option |  | 4
| 1 | Biphasic, synchronous and asynchronous manual defibrillator for hospital and outside environment
| 2 | Selectable energy levels for defibrillation
| 3 | Device must have integrated paddles for defibrillation with buttons for charge and shock.
| 4 | Integrated mains and battery supply
| 5 | Battery capacity of min. 150 shocks with a maximum energy or 3.5 h of ECG monitoring
| 6 | Device must have an indicator for charging and discharging of battery
| 7 | Device must have integrated min 6" display readable in all conditions.
| 8 | Device must have integrated min. 3-channel thermal printer with automatic and manual operating mode
| 9 | Thermo printer uses paper min. 70 mm width
| 10 | Device must have integrated interfaces: min. 1xUSB and 1xRS 232
| 11 | Integrated device memory for a minimum of 12 hours of ECG or 300 events
| 12 | TThe complete device software must be in Serbian or English
| 13 | Device must have integrated external transthoracic pacemaker option with Fix, Demand and Overdrive operating mode
| 14 | Device must have pads connector integrated on front panel of the device
| 15 | Weight of defibrillator including battery and pedals max. 6 kg+/- 1kg
| 16 | Device must have minimum IP 53 protection standard, optional IP 34, IP4X
| 17 | Defibrillator should come with the following accessories: defibrillation paddles, defibrillation paddles adapter for children, 1 x pacing/defibrillation pads, 1 x min.4-lead ECG patient cable, 1 x defibrillation gel, 1 x mains cable, 1 x user manual, 1 x thermic paper
1.7 |  | Hospital Defibrillator |  | 3
| 1 | Biphasic, synchronous and asynchronous, manual defibrillator for hospital and outside environment
| 2 | Selectable energy levels for defibrillation
| 3 | Device must have integrated paddles for defibrillation with buttons for charge and shock.
| 4 | Integrated mains and battery supply
| 5 | Battery capacity of min. 150 shocks with a maximum energy or 3.5 h of ECG monitoring
| 6 | Device must have an indicator for charging and discharging of battery
| 7 | Device must have integrated min 6" display readable in all conditions.
| 8 | Device must have integrated min. 3-channel thermal printer with automatic and manual operating mode
| 9 | Thermo printer uses paper min. 70 mm width
| 10 | Device must have integrated interfaces: min. 1xUSB and 1xRS 232
| 11 | Integrated device memory for a minimum of 12 hours of ECG or 300 events
| 12 | TThe complete device software must be in Serbian or English
| 13 | Device must have the ability to be upgraded with an optional external transthoracicpacemaker with Fix, Demand and Overdrive operating mode
| 14 | Weight of defibrillator including battery and pedals max. 6 kg +/-1 kg
| 15 | Device must have minimum IP 53 protection standard, optional IP 34, IP4X
| 16 | Defibrillator should come with the following accessories: defibrillation paddles, defibrillation paddles adapter for children, 1 x min.4-lead ECG patient cable, 1 x defibrillation gel, 1 x mains cable, 1 x user manual, 1 x thermic paper