## Price schedule
Bidder: |  |  |  | Offered price (currency):
LOT 3 | INTENSIVE CARE
ID | Technical Specification Requested | QTY | Technical Specification Offered (model) | Unit price | Total price per line item
3.1 | Reanimation table | 2 |  |  | 0
3.2 | Hybrid incubator | 6 |  |  | 0
3.3 | Stationary Incubator | 7 |  |  | 0
3.4 | Infant Incubator with phototherapy | 6 |  |  | 0
3.5 | Infant Warmer Bed | 6 |  |  | 0
3.6 | Vital sign patient monitors and central monitors | 48 |  |  | 0
3.7.1 | Intensive care mechanical ventilator for pediatric application | 7 |  |  | 0
3.7.2 | Universal mechanical ventilator for adult, paediatric and neonatal application | 3 |  |  | 0
3.8 | Intensive care mechanical ventilator for neonatal application | 5 |  |  | 0
3.9.1 | Surgical suction unit | 20 |  |  | 0
3.9.2 | Aspirator for intensive care unit | 15 |  |  | 0
3.9.3 | Aspirator for thoracic drainage | 5 |  |  | 0
3.10.1 | Hospital bed for intensive care with mattress | 4 |  |  | 0
3.10.2 | Pediatric IC bed | 2 |  |  | 0
3.11 | Monitor for measurement hemodynamic parameters | 2 |  |  | 0
3.12 | Infusion pump | 40 |  |  | 0
3.13 | Syringe pump | 120 |  |  | 0
3.14 | Pulse oximeter | 30 |  |  | 0
3.15 | Hyper-hypothermia system | 1 |  |  | 0
3.16 | Nitric oxide delivery system | 1 |  |  | 0
3.17 | Cerebral/Somatic Oximeter monitor | 2 |  |  | 0
3.18 | Phototherapy lamp | 2 |  |  | 0
3.19 | Mixer low-flow | 6 |  |  | 0
3.20 | Mixer high -flow | 2 |  |  | 0
3.21 | Mobile video intubation endoscope/bronchoscope | 1 |  |  | 0
3.22 | Fiberoptic laryngosope | 2 |  |  | 0
3.23 | Fiberoptic laryngosope for difficult intubations | 1 |  |  | 0
3.24 | Portable tablet Ultrasound system for anesthesia | 1 |  |  | 0
3.25 | Nerve Stimulator for Peripheral Regional Anesthesia | 1 |  |  | 0
3.26 | Combined monitor for respiratory status | 2 |  |  | 0
3.27 | Device for measuring and monitoring of the depth of anesthesia and sedation | 4 |  |  | 0
3.28 | Heart Lung machine | 1 |  |  | 0
3.29 | Autotransfusion device | 1 |  |  | 0
3.30 | Calibrated Monitor for measurement hemodynamic parameters | 1 |  |  | 0
3.31 | Cardiology Ultrasound Device | 1 |  |  | 0
3.32 | MONITORS FOR MONITORING PATIENT VITAL FUNCTIONS | 1 |  |  | 0
3.33 | Patient monitor with transport module | 2 |  |  | 0
3.34 | ACT analyzer | 1 |  |  | 0
3.35 | In line Blood Parameter Monitoring System | 1 |  |  | 0
3.36 | Transcutaneous bilirubinometer | 1 |  |  | 0
3.37 | Anesthesia device | 5 |  |  | 0
|  |  |  | Total Price per Lot | 0

## Technical specifications
Bidder:
LOT 3 |  | INTENSIVE CARE |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | Unit Price | Total Price | Insert page no. in technical documentation
3.1 |  | Reanimation table |  | 2
| 1 | The warm bed must have a resuscitation module
| 2 | Integrated inspection lamp> 0.11 lumens / cm2 (100 ft. Candles)
| 3 | alarms:
| 4 | heater alarms
| 5 | Check patient
| 6 | Baby temperature
| 7 | high temperature
| 8 | probe
| 9 | System error
| 10 | Power failure
| 11 | Setting 34-38 ° C (> 37.0 ° override)
| 12 | Heater 0-100%, 10% increments
| 13 | Temperature display 18-43 ° C
| 14 | precision +/- 0.2 ° C (31 ° C to 37 ° C)
| 15 | Breath frequency 18-60 BPM (+/- 10% of setting)
| 16 | I: E ratio Fixed at 1: 2 Nominal
| 17 | Pressure (PEEP) 0-18 H20 (0-1.77 kPa)
| 18 | Gas 5 LPM Max
| 19 | Precision blender 21-100% O2 +/- 3% 02
| 20 | Suction 0-150 mmHg (0-20 kPa)
| 21 | height 180-200cm
| 22 | width 70-80 cm
| 23 | Depth 110-115 cm
| 24 | Weight 90-120 kg
3.2 |  | Hybrid incubator |  | 6
| 1 | Thermoregulation in closed care
| 2 | Thermoregulation in open care with radiant warmer
| 3 | Quick transition between open care and closed care
| 4 | Trolley with 4 castors with locking brake
| 5 | Integrated work light with intensity levels
| 6 | Possibility for mounting arms for monitor, phototherapy, suction unit etc.
| 7 | Mattres that is heated and x-ray translucent
| 8 | Device has mechanical opening and closing hood (opening and closing possible when device is switched off)
| 9 | hand ports, that can be easy opened, for access to the baby - 2 ports on each side and one at the front side
| 10 | Access panels for interventions, that can be opened on each side and front access panel
| 11 | tubing grommets
| 12 | Electrical height adjustment operable from both sides by foot pedal - mattres adjustment hight range 700mm - 1100mm
| 13 | Integrated electrical tilt adjustment
| 14 | Integrated scale
| 15 | Operating volume 40 dB
| 16 | Flow speed over mattress 10 cm/s
| 17 | Device has three heat sources - convective heating, radiant warmer and heated mattress, that are synchronized
| 18 | Display
| 19 | min 10” colour touchscreen display
| 20 | Graphical and tabular display of trends 2 hours to 7 days for temperatures, heating, noise level/ light level measurement, weight, events
| 21 | Timer Stopwatch and APGAR timer with alarms after 1 min, 5 min, 10 min.
| 22 | Thermoregulation
| 23 | Warm up time less than 20min
| 24 | Operation in air controlled mode and skin controlled mode
| 25 | Air temerature control and monitoring
| 26 | Air temperature measurement range from 13 to 45 °C
| 27 | Air temperature setting range from 20 to 39 °C
| 28 | Extended Air temperature setting range (after confirmation), < 28 °C and > 37°C
| 29 | Skin Temperature Control and monitoring
| 30 | Skin temperature measurement range from 13 to 43 °C
| 31 | Central Skin temperature setting range from 34 to 38 °C
| 32 | Extended Central Skin temperature setting range (after confirmation) >37 °C
| 33 | Relative humidity control and monitoring
| 34 | Servo-controlled humidification system integrated
| 35 | Contamination free humidification system through boiling and evaporation of water
| 36 | Closed humidification system with external water reservoir
| 37 | Humidification measurement range from 30 to 99 % RH
| 38 | Humidification setting range from Off, 30 to 99 % RH in steps of 1 %, AUTO
| 39 | Radiant warmer
| 40 | heating elements
| 41 | automatic adaption of radiant warmer when mattress is tilted
| 42 | Setting range Off, 10 % to 100 %
| 43 | Pre-warm mode
| 44 | Heated mattress
| 45 | X-ray translucent
| 46 | Heated mattress temperature setting range from Off, 35 to 39 °C, AUTO
| 47 | Oxygen regulation (Optional)
| 48 | Servo-controlled Oxygen enrichment with connection to central O2 supply or to O2-cylinder
| 49 | Measurement range: 18 to 65 Vol% (normal) and 66 to 99 Vol% (extended)
| 50 | O2 rise time from 21 to 65 Vol% less than 10 minutes
| 51 | Alarm management
| 52 | Ramping alarm volume
| 53 | Suppression and delaying of alarms automatically
| 54 | Alarm volume setting range 50 – 70 dB(A)
| 55 | Data Exchange and Interfaces
| 56 | Serial port COM
| 57 | USB port
| 58 | Nurse call
| 59 | USB port
3.3 |  | Stationary Incubator |  | 7
| 1 | Access to:
| 1.1 | – front and rear access panel
| 1.2 | – 6 access ports
| 1.3 | – 3 left and 3 right tubing grommets -front
| 1.4 | – 2 left and 2 right tubing grommets – rear"
| 2 | Double temperature monitoring
| 3 | 2-4-8-12-24 h trend: air temperature, skin temperature, humidity, O2 concentration, hetarer power
| 4 | Seven days trends gain / weight loss
| 5 | Temperature range when measuring the air inside the incubator 20 to 39 C, with a confirmation when you want to over 37 C
| 6 | flow rate of air flow through the mattress <10 cm / sec
| 7 | Temperature range when the temperature of the baby's skin 34 to 38 C, with a confirmation when you want to over 37 C
| 8 | integrated servo control with a moisture range from 30 to 95% with an accuracy of up to 6%, reservoir capacity 1,600 ml
| 9 | An integrated servo control oxygen accuracy ± 2% of full scale
| 10 | Calibration O2 may be carried out ambient air or 100% oxygen
| 11 | The electroluminescent display for the moisture content, temperature, and oxygen weight
| 12 | The integrated built-in scale to the range of 7 kg with a resolution of 1g, the accuracy of 2g (for weight up to 2 kg) and 5g (for weight over 2kg)
| 13 | Adjust the range of 21-65% oxygen, CO2 level less than 0,5 %
| 14 | Internal noise level less than 47 dB.
| 15 | The existence of an alarm:
| 15.1 | - High / low temperature skin baby
| 15.2 | - Air temperature
| 15.3 | - Error probe for measuring air temperature
| 15.4 | - Error probe for measuring temperature bebe
| 15.5 | - Humidity
| 15.6 | - High / low oxygen concentration
| 15.7 | - The error small airflow
| 15.8 | - The required calibration of the oxygen sensor
| 15.9 | - oxygen sensor fault
| 16 | The existence of the alarm system:
| 16.1 | - engine fault
| 16.2 | - fault sensor module
| 16.3 | - errors heater
| 16.4 | - disconnection sensor
| 16.5 | - jammed keys
| 17 | The possibility of installing the tube holder, the carrier and shelving infusion to monitor
| 18 | The slope of the mattress ± 12%, the tolerance of the inclination of ± 1%
| 19 | Adjustable height 135cm to 150 cm
3.4 |  | Infant Incubator with phototherapy |  | 6
| 1 | Microprocessor control system
| 2 | Main Unit at high adjustable stand with four locking casters
| 3 | Mttress high position is adjusted: 800-1200mm or larger
| 4 | Mattress tilting range: 13° or more
| 5 | Foot switches provided on both sides of the Hi-Low stand
| 6 | TFT LCD “touch screen“ display size: min 8 inches
| 7 | Automatic temterature control: Incubator air temperature setting range: 25-37C Incubator air temperature display range: 20-42°C Skin temperature setting range: 34-37.5°C Skin temperature display range: 30-42°C
| 8 | Heater power adjustable in minimum 10 levels
| 9 | Warm-up time to 36°C: 60min or less
| 10 | Automatic oxygen control: Oxygen concentration setting range: 22-65% Oxygen concentration display range: 19-105%
| 11 | Automatic humidity control: Relative humidity setting range: 40 – 95% Relative humidity display range: 15 – 99% Maximal relative humidity: 90% or more
| 12 | Pulse oximeter is integrated: SpO2 display range: 10-100% Pulse rate display range: 30-240bpm
| 13 | Noise level: maximum 45dB
| 14 | Double lock system of each admittance panel
| 15 | An X-ray cassette can be inserted and removed without opening an admittance panel so that the incubator air temperature will not drop
| 16 | Alarms: high temperature, power failure, set humidity, skin temperature probe, oxygen concentracion, SpO2 limit, pulse rate limit and system failure
| 17 | Dual wheel caster for more stability of the unit with four locking casters
| 18 | AC power source: AC230V ±10%; 50/60Hz
| 19 | Power consumption: maximum 600VA
| 20 | Degree of protection: Type BF applied part
| 21 | Incubator have to be delivered with all accessories needed, including: - Skin temperature probe reusable, 1pc - Infant SpO2 probe reusable, 1 pc - Dust cover, 1 pc - Operation manual, 1 pc
| 22 | External Phototherapy Unit at flexibile arm with attachment for incubator stand Examination LED light built-in to monitor the infant's condition Radiation level: 30 - 40 uW/cm2, at distance of 30cm Light illumination can be adjusted: 500 - 2000Lux, at distance of 30cm AC power source: AC220-240V 50/60Hz Power consumption: max 30VA
3.5 |  | Infant Warmer Bed |  | 6
| 1 | Main Unit at the stand with four locking casters
| 2 | Mattress high adjustable: 850-1000mm or larger
| 3 | Automatic temterature control: Skin temperature setting range: 34-38°C Skin temperature display range: 30-42°C
| 4 | APGAR and CPR timer are available
| 5 | Heater power adjustable in minimum 10 levels
| 6 | Warming head: rotation minimum +/-85° or more
| 7 | Mattress longitudinal tilting range: 13° or more
| 8 | Built in LED lighting lamp
| 9 | X-ray cassette tray
| 10 | Alarms: power failure, skin temperature probe and system error
| 11 | Degree of protection: Type BF applied part
| 12 | AC power source: AC230V+/-10%; 50/60Hz
| 13 | Power consumption: max 700VA
| 14 | Drawer for accessories
| 15 | Warming Unit have to be delivered with all accessories needed, including: - Skin temperature probe reusable, 1pc - Mattress sheet, 1 pc - Dust cover, 1 pc - Operation manual, 1 pc
| 16 | Resuscitation external module: Pressure gauge indication: -20 to 80cmH2O Pressure limit: 20 to 60cmH2O Oxygen concentration control: 21-100% Flowmeter adjusment: 0-15l/min Accessories: Patient circuit with T-flow safety valve (40 pcs), Test bag, AIR and O2 connecting hoses, Infant face masks size 1 (20 pcs) and size 2 (20 pcs)
| 17 | External suction unit, venturi (ejector) type: Continuous vacuum regulator (analog gauge) Vacuum level: 0-250 mbar Horizontal rotation of the vacuum gauge from -45° to +45° for a better visibility Complete mounting system: rail clamp, AIR hose and DIN connection Accessories: 150 ml safety jar, 1 L polysulfone collection jar with rail mount, antibacterial filters (100 pieces), Suction tip for drainage (10 pieces)
3.6 |  | Vital sign patient monitors and central monitors |  | 48
| 1 | TFT LCD color display, min. 10", touchscreen
| 2 | Simultaneous display of min. 5 waveforms and 10 numeric parameters
| 3 | Connections for monitoring of ECG, Respiration, SpO2, NIBP, Temperature
| 4 | Connection for IBP monitoring
| 5 | NIBP measurements triggered by changes in pulse wave propagation time
| 6 | Pacemaker signal detection
| 7 | Automatic ST segment level detection, alarming and storing
| 8 | Automatic arrhythmia detection, alarming and storing: min. 10s per episode, total min. 120h
| 9 | Graphic and numerical parameters trend, min. 120h
| 10 | ECG full disclosure min. 120h
| 11 | Visual and sound alarm, 3 priority levels, adjustable limits, storage for min. 120h
| 12 | Continuous noninvasive monitoring of estimated hemodynamic parameters CO and SV
| 13 | ECG sync sound with pitch variable according to current SpO2 and IBP value
| 14 | Choice for enlarged numerical display mode, readable from long distance
| 15 | Power supply AC 220V/50Hz with battery backup min. 3h
| 16 | Upgradeable with thermal recorder, min. 3 channels
| 17 | Defibrillation proof patient connectors
| 18 | LAN connection for central monitoring system and interbed monitoring
| 19 | Upgradeable with wireless telemetry
| 20 | Onscreen operation guide with illustrations
| 21 | Maximal weight: 4kg
| 22 | Each monitor to be delivered with the following accessory:
| 22.1 | ECG cable with 3 leads, 1 pc
| 22.2 | Reusable waterproof SpO2 probe, clip type, 1 pc
| 22.3 | Reusable waterproof SpO2 probe, neonatal, 1 pc
| 22.4 | Reusable NIBP cuff, large, 1 pc
| 22.5 | Reusable NIBP cuff, medium, 1 pc
| 22.6 | Reusable NIBP cuff, small, 1 pc
| 22.7 | Reusable NIBP cuff, extra small, 1 pc
| 22.8 | Disposable NIBP cuff, neonatal, 10 pcs
| 22.9 | Disposable ECG electrodes, neonatal, 120 pcs
| 22.1 | IBP cable, 1 pc
| 22.11 | Accessory for continous noninvasive monitoring of CO and SV, 1 set
| 22.12 | Holder for monitor, fixation to medical rail or wall, 1 pc
| 22.13 | Battery, 1 pc
Minimal technical specifications of central monitors:
| 23 | PC based central station for monitoring of min. 16 simultaneous patients
| 24 | TFT LCD color display, min 22", upgradeable with secondary display
| 25 | Waveforms and numerical parameters monitoring and storing for all beds
| 26 | Upgradeable for up to min. 48 monitors
| 27 | Capable to display min 12 waveforms for selected monitor, simultaneously
| 28 | Numeric and graphical trend min. 120h for each patient
| 29 | Memory for storing min. 750 arrhythmia episodes and min 7200 ST episodes
| 30 | Memory for storing min. 120h of full disclosure, min. 8 waveforms per bed
| 31 | Displaying of 12-ch ECG for selected patient, if available at bedside monitor
| 32 | Displaying of patient graphic and visual alarms, three priority levels
| 33 | Alarm limits adjustment
| 34 | Memory for storing min. 1000 alarm episodes per bed
| 35 | Upgradeable for 12 lead ECG interpretation
| 36 | Eash patient monitor electrically insulated by dedicated insulation units (transformers)
| 37 | Patient reports printed on laser printer A4
| 38 | Unconditional power supply (UPS)
Network infrastructure (LAN) will be provided by hospital
3.7.1 |  | Intensive care mechanical ventilator for pediatric application |  | 7
| 1 | Intensive care mechanical ventilator for pediatric and adult patients, bodyweight 3kg to 130kg.
| 2 | Upgradeable with neonatal option, bodyweight 0.2 - 30kg or better
| 3 | Ventilator based on dedicated trolley with castors and brakes, with arm for patient circuit and two quick-lock holders for the arm (left and right side of the trolley).
| 4 | Ventilator must have carrying handle, with easy one-button removal from the trolley.
| 5 | Power supply AC 220V/50Hz, DC 12-24V from wehicle, with integrated Li-Ion batery backup min. 3 h.
| 6 | Upgradeable with additional hot-swappable battery, min. 3h (total 6h).
| 7 | Integrated air turbine, quiet and efficient, with HEPA filter, flow min. 240l/min.
| 8 | DIN connection for O2 high pressure central gas supply.
| 9 | Low pressure connection for O2 concentrator.
| 10 | TFT LCD color touchscreen, min. 12", with tilt adjustment.
| 11 | Ventilation modes needed: - Volume controlled CMV and pressure controlled CMV - Volume controlled SIMV and pressure controlled SIMV - Pressure support - BiPAP (or equivalent) and APRV - NIV with guaranteed minimal frequency - High-Flow O2 therapy - Upgradeable with neonatal nCPAP therapy
| 12 | Ventilator must have closed loop ventilation mode with MV guaranteed, adapting ventilation parameters according to current patient respiratory mechanics (R, C and activity). This adaptive ventilation should work automatically: since intubation, through weaning until extubation. Intended for both intubated and nonintubated patients, for obstructive, restrictive or healthy lungs.
| 13 | Upgradeable with fully closed loop ventilation mode: automatic adaptation of minute volume, PEEP and FiO2 according to current values of EtCO2 and SpO2.
| 14 | Ventilation parameters minimal settings range: - Pinsp: 5 - 60 cmH2O over PEEP - Psupport: 0 - 60 cmH2O over PEEP - PEEP: 0 - 35 cmH2O - I:E: 1:9 do 4:1 - f: 1 - 80 bpm - FiO2: 21 - 100% - Vt: 2 - 2000 ml
| 15 | Flow trigger with automatic adjustment to current leak.
| 16 | Automatic tube resistance compensation.
| 17 | Reversible apnea backup function (off when apnea is over).
| 17 | Capability to drive and control pneumatic nebuliser.
| 18 | Equipped with electronic nebulizer, connected via USB port.
| 19 | Needed functions: preoxygenation, suction, manual breath, inspiration hold.
| 20 | Visual and audible alarms with priorities and adjustments of alarm limits.
| 21 | Alarm lamp on top of ventilator, visible from all directions (360°).
| 22 | Monitoring trends min. 72h.
| 23 | Capability to display min. four simultaneous waveforms, loops and trends.
| 24 | Ventilator must have graphic animation for visualization of lung dynamics (Vt, R, C) and patient activity, as well as graphic animation for visualization of patient - ventilator dependency (oxygenation, CO2 elimination and spontaneous breathing) with adjustable weaning criteria.
| 25 | Monitoring parameters: - Tidal volume, spontaneous MV and total MV - Spontaneous frequency, total frequency, TI, TE - Minute volume of leak - Ppeak, Pmean, AutoPEEP - Airway resistance, lung compliance, RC and RSB - Oxygen concentration
| 26 | Proximal flow sensor.
Event log for alarms and ventilation settings.
| 27 | Program tool for automated lung assessment and recruitment maneuver. Recording of complete quasi static compliance curve (inflation and deflation), for evaluation of hysteresis, with adjustable maneuver parameters and pause for recruitment. Cursor for measurements and finding optimal pressures (PEEP, Pplato). Saving history of loops and reference loops.
| 28 | Volumetric capnography, graph FCO2 vs. V, measurement of etCO2, Vdeadspace and VTalveolar.
| 29 | Communication: USB port for screen snapshots and saving log files, upgradeable with RS232 for remote alarms, nurse call and telemetry.
| 30 | Warm humidifier with manual and automatic temperature / humidity controll and watter filling.
| 31 | To be delivered with all accessories needed, including: - Patient breathing sets for adult/pediatric, with humidifier chamber, 15 pcs - Flow sensor adult/pediatric, 15 pcs - NIV masks, three different sizes, 3 pcs - High Flow O2 therapy cannulas, 10 pcs - Accessory for electronic nebulizer, 1 set - Accessory for volumetric capnography, 1 set
| 32 | Maximum ventilator weight without trolley: 10 kg.
3.7.2 |  | Universal mechanical ventilator for adult, paediatric and neonatal application |  | 3
| 1 | Patient category: adult, pediatric and neonate
| 2 | Device is on mobile stand, with brakes on castors
| 3 | Central supply with medical gas (Compressed air and Oxygen)
| 4 | Electrical supply:
| 4.1 | - 220V AC
| 4.2 | - internal battery with 30 min. capacity
| 5 | Display: size minimum 17“, colour, touchscreen;
| 6 | Screen configuratuion with 6 different views
| 7 | Ventilator uses permanent method of O2 concentration measurement, without sensor replacing – ultrasound or paramagnetic probe
| 8 | Device is using standard patient breathing hoses
| 9 | Ventilation modes:
| 9.1 | Volume controlled modes:
| 1 | - assisted (VC-AC)
| 2 | - synchronized (VC-SIMV)
| 3 | - ventilation with guaranteed minimal minute volume
| 9.2 | Pressure controlled modes:
| 1 | - assisted (PC-AC)
| 2 | - synchronized (PC-SIMV)
| 3 | - synchronized with possibility for breathing on two levels (PC-BIPAP, PC-BiLevel or relevant)
| 9.3 | Spontaneous modes (with continuous positive airway pressure):
| 1 | - CPAP/PS
| 2 | - CPAP/VS
| 3 | - CPAP/ variable pressure support
| 4 | - Variable Pressure Support ventilation
| 5 | - Pressure Support proportional to patients effort – PAV or PPS or NAVA or adequate
| 6 | Integrated O2 therapy
| 7 | Device has possibility to deliver set tidal volume with minimal possible pressure – PRVC, Autoflow, ASV or relevant
| 8 | Non invasive ventilation – NIV in all ventilation modes and all categories, with automatic leakage compensation of 180 L/min
| 9 | Protocol for patient weaning from mechanical ventilation - PPS / SBT/ Intelli Vent/ PAV +/ NAVA or adequate
| 9.4 | Ventilator settings:
| 1 | Tidal volume (VT): 20 - 3000 ml
| 2 | Breathing frequency (RR): 1 - 150 bpm
| 3 | Inspiratory pressure: 1 -95 cm H2O;
| 4 | PEEP: 0 - 50 mbar
| 5 | Pressure assist/support: 0 -95 cm H2O
| 6 | O2 concentration – FiO2: 21 - 100%
| 7 | Inspiratory time (Tinsp): 0,1 – 10s
| 8 | Triggering: flow trigger, with auto trigger adjust corresponding to leakage;
| 9 | Trigger sensitivity: 0,2 – 15 L/min
| 10 | Automatic tube compensation – АTC
| 11 | Leakage compensation;
| 12 | Measurement of expiratory CO2 concentration (mainstream);
| 13 | Volumetric CO2-Monitoring : VCO2, VTCO2
| 14 | Measurement and calculating:
| 14.1 | - Compliance
| 14.2 | - Resistance
| 14.3 | - Intrinsic PEEP or AutoPEEP
| 9.5 | View on display:
| 1 | - 4 respiratory curves simultaneously (pressure, volume, flow)
| 2 | - pressure-volume loop
| 9.6 | Parameters:
| 1 | - airway pressure
| 2 | - tidal volume
| 3 | - minute volume
| 4 | - FiО2
| 5 | - ЕtCО2
| 6 | - compliance (C)
| 7 | - resistance (R)
| 8 | Special maneuvres for inspiration and expiration hold
| 9 | O2 Suction manouvre, with auto disconnection an reconnection detection, and oxygen enrichment
| 10 | Calculating parameters for weaning from mechanical ventilation:
| 10.1 | - P0.1
| 10.2 | - RSBI
| 10.3 | - NIF
| 11 | Graphical representation of lungs, and change of airway resistance and lung compliance
| 12 | Trends: minimum 7 days
| 13 | With device should be delivered:
| 14 | - Arm for breathing hoses
| 15 | - nebulizer
| 16 | - breathing hoses for adults (1.5m)
| 17 | - instructions for use
| 18 | - active humidifier for neonate and hoses
3.8 |  | Intensive care mechanical ventilator for neonatal application |  | 5
| 1 | Intensive care mechanical ventilator for adult, pediatric and neonatal patients, bodyweight > 0.2kg.
| 2 | Ventilator based on dedicated trolley with castors and brakes, with arm for patient circuit.
| 3 | Power supply AC 220V/50Hz, with integrated battery backup for minimum 60 min.
| 4 | Standard connection for Air and O2 high pressure central gas supply.
| 5 | TFT LCD color touchscreen, min. 15", with tilt and pan adjustment.
| 6 | Ventilation modes needed: - Volume controlled CMV and pressure controlled CMV - Volume controlled SIMV and pressure controlled SIMV - Pressure support - BiPAP (or equivalent) and APRV - NIV with guaranteed minimal frequency - High-Flow O2 therapy - Neonatal nCPAP therapy
| 7 | Ventilation parameters minimal settings range: Pinsp: 5 - 90 mbar Psupport: 0 - 90 mbar PEEP: 1 - 50 mbar I:E: 1:6 do 4:1 f: 1 - 150 bpm FiO2: 21 - 100% Vt: 2 - 2000 ml
| 8 | Reversible apnea backup function (off when apnea is over).
| 9 | Flow trigger and pressure trigger, adjustable.
| 10 | Tube Resistance Compensation and Leak Compensation.
| 11 | Inspiration hold and Expiration hold functions.
| 12 | Preoxygenation and Manual breathe functions.
| 13 | Capability to drive and control pneumatic nebulizer.
| 14 | Equipped with electronic nebulizer, controlled by ventilator.
| 15 | Upgradeable with integrated ET cuff control system: cuff pressure setting, monitoring and alarming via ventilator display.
| 16 | Visual and audible alarms with priorities and adjustments of alarm limits.
| 17 | Alarm lamp on top of ventilator, visible from all directions (360°).
| 18 | Capability to display min. 8 simultaneous waveforms, loops and trends.
| 19 | Monitoring trends min. 96h.
| 20 | Event log for alarms and ventilation settings.
| 21 | Monitoring parameters: - Tidal volume, spontaneous MV and total MV - Spontaneous frequency, total frequency, TI, TE - Minute volume of leak - Ppeak, Pmean, Pplato, AutoPEEP, P01 - Insp. and exp. airway resistance, lung compliance, PTP, RC and RSB - Oxygen concentration
| 22 | Ventilator must have closed loop ventilation mode with MV guaranteed, with automatic adaptation of ventilation parameters according to current patient respiratory mechanics (R, C and activity). With automatic protective limits, preventing barotraumas, autopeep and deadspace ventilation. This adaptive ventilation should work automatically: starting with intubation, through weaning, until extubation. Intended for both intubated and nonintubated patients, for obstructive, restrictive or healthy lungs.
| 23 | Program tool for automated lung assessment and recruitment maneuver. Recording of complete quasi static compliance curves (inflation and deflation), for evaluation of hysteresis, with adjustable maneuver parameters and pause for recruitment. Cursor for measurements and finding optimal pressures (PEEP, Pplato). Saving history of loops and reference loops.
| 24 | Ventilator must have graphic animation for visualization of lung dynamics (Vt, R, C) and patient activity, as well as graphic animation for visualization of patient - ventilator dependency (oxygenation, CO2 elimination and spontaneous breathing) with adjustable weaning criteria.
| 25 | Upgradeable with fully closed loop ventilation mode: automatic adaptation of minute volume, PEEP and FiO2 according to current values of EtCO2, SpO2 and measured hemodynamic stability.
| 26 | Upgradeable for helliox application and gas supply.
| 27 | Proximal flow sensor.
| 28 | Auxiliary pressure input for esophageal (transpulmonary) pressure measurement.
| 29 | Upgradeable with paramagnetic O2 sensor.
| 30 | Output connection for external video screen.
| 31 | Equipped with volumetric capnography: graph FCO2 and PCO2 vs. V, measurement of etCO2, Vdeadspace, VTalveolar and display on ventilator's screen.
| 32 | Upgradeable with plethismography: plethismography curve, measurement of pulse rate, SpO2, HLI and display on ventilator's screen.
| 33 | Upgradeable with additional hot-swappable batteries.
| 34 | Communication: upgradeable with 2x RS232 for digital communication, connection to warm humidifier, remote alarms and telemetry.
| 35 | Function of screenshots saving to a memory card.
| 36 | Warm humidifier with manual and automatic temperature / humidity controll and watter filling.
| 37 | To be delivered with all accessories needed, including: - Patient breathing sets for neonatals, with humidifier chamber, 15 pcs - Flow sensor neonatal, 15 pcs - NIV masks, three different sizes, 3 pcs - "High Flow" O2 therapy cannulas, 10 pcs - Accessory for nCPAP therapy, 20 pcs - Accessory for electronic nebulizer, 1 set - Accessory for volumetric capnography, 1 set
3.9.1 |  | Surgical suction unit |  | 20
| 1 | Flow rate: not less than 90 liters/minute
| 2 | Vacuum -0.9 bar
| 3 | Noise level ≤ 50dB
| 4 | Foot control
| 5 | Mobile frame with antistatic castors
| 6 | Power supply 230V – 50Hz
| 7 | Device weight maximum: 20kg
| 8 | Standard accessories: - two polycarbonate, autoclavable bowls of 2 liter equipped with the anti-overflow valve, - set of silicone hoses, conical appendage
3.9.2 |  | Aspirator for intensive care unit |  | 15
| 1 | Electrical suction pump on trolley
| 2 | Suction capacity 30 L/min
| 3 | Max vacuum -90 kPa (-675 mm Hg)
| 4 | Silent vacuum pump, maximum noise up to 37 dB
| 5 | Membrane vacuum regulator
| 6 | Vacuum gauge in kPa and mmHg
| 7 | Safety set with polysulfone bottle 0.25 L, large lid and overflow protection
| 8 | Touch-sensitive on/off button
| 9 | Standby indicator
| 10 | Trolley with four antistatic castors and brakes, on/off switch integrated on trolley
| 11 | Weight of suction pump up to 10 kg
| 12 | Equipotential conductor integrated on suction pump housing
| 13 | Double-reusable 2L set with polysulfone jar, change-over valve and overflow protection
3.9.3 |  | Aspirator for thoracic drainage |  | 5
| 1 | Compact digital, mobile,electric aspirator for thoracic drainage
| 2 | Full digital monitoring of the patient treatment
| 3 | Intended to be used for aspiration and removal of surgical fluids, tissue, gases, bodily fluids or infectious materials
| 4 | Electronic measuring and monitoring system with optical and acoustic status indications
| 5 | Dry system, which means no fluids are necessary for operation
| 6 | Battery operated device or from main power supply
| 7 | Monitors the air leak and only applies the suction required to maintain the negative pressure prescribed and set by the managing physician
| 8 | Numerical and graphic data for pressure, liquids and air leak are shown on digital display
| 9 | Pressure range from -0,1 to -10 kPa (-100 mbar, -75 mmHg).
| 10 | Flow 5 l/min
| 11 | Aspirator weight 1 kg
| 12 | System use single or double tubing for connection between patient and aspirator
3.10.1 |  | Hospital bed for intensive care with mattress |  | 4
| 1 | External dimension with split plastic siderails and standard bumpers: 220 x 95±5cm
| 2 | Bed extension: min 30 cm
| 3 | Bed Height (Hi-Lo adjustment) : 40 - 75 cm
| 4 | Height of split plastic siderails: 40 cm
| 5 | SWL -Safe Working Load: 250 kg
| 6 | Max. Patient Weight : 185 kg
| 7 | TR/RTR angle: 16°/ 16°
| 8 | Standard : EN ISO 9001:2000; EN ISO 13485:2003;EN 60601-1; EN 60601-1-2; EN 60601-1-4; EN 60601-2-52; EN ISO 14971;
| 9 | IP protection : IP X4
| 10 | The bed complies with all requirements stated in: MDD 93/42/EEC; 90/384/EEC
| 14 | BED ENDs:
| 14.1 | Plastic bed ends
| 14.2 | Removable: with safety lock
| 14.3 | IV poles lifting pole holders: 4x in the corner of bed
| 15 | SIDERAILS
| 15.1 | Split plastic siderails
| 15.2 | Holders for urine bag integrated in split plastic siderails
| 16 | MATTRESS PLATFORM
| 16.1 | 4 section mattress platform
| 16.2 | Maximum Backrest Angle: 60°
| 16.3 | Back rest adjustable: electric
| 16.4 | Maximum Thigh Rest Angle: 30°
| 16.5 | Thigh rest adjustable: electric
| 16.6 | Calf rest adjustable: Mechanic
| 16.7 | Mattress platform : All sections are covered with blow molding plastic parts (incl. fixed parts)
| 17 | CONTROLS
| 17.1 | Activation with GO button on all controlls
| 17.2 | Cardiac chair: preprogrammed one button on supervisor
| 17.3 | Examination position: preprogrammed one button on supervisor panel
| 17.4 | Supervisor (ACP) panel on flexible cable
| 17.5 | Functions lockouts through supervisor (ACP) panel
| 18 | CONSTRUCTION
| 18.1 | Column construction: 2 columns
| 18.2 | Manual CPR backrest handle
| 18.3 | Bed frame colour: 9002
| 18.4 | Pair of accesories rails
| 19 | UNDERCARIAGE , BRAKES and CASTORS:
| 19.1 | centrally operated castors
| 19.2 | 4 castors 150 mm
| 19.3 | 2 brake levers, foot side operated
| 19.4 | Number of actuators: 4 linear actuators, backrest with CPR mechanical releasing
| 19.5 | Battery back up
| 20 | Safe use:
| 20.1 | Risk category of presure ulcers (*with an appropriate turning and re-positioning regime): II
| 23 | Cover:
| 23.1 | Vapour permeability
| 23.2 | Fluid resistant
| 23.3 | Inflammability: CRIB 7
3.10.2 |  | Pediatric IC bed |  | 2
| 1 | Electric ICU paediatric bed with sleep deck in four sections
| 2 | Height adjustable
| 3 | Trendelenburg and reverse-Trendelenburg movements
| 4 | Robust steel frame with a thick epoxy coated finish for high scratch and corrosion resistance
| 5 | Electrical back section movement. Manual CPR lever on both sides of the bed with dampened motion
| 6 | Fixed pelvic section.
| 7 | Electrical upper leg section and manual lower leg section movement via a locking latch mechanism.
| 8 | Electrical sleep deck height variation on robust telescopic columns. Minimum height approx. 65cm up to 85cm from ground
| 9 | Equipped as standard with transparent side rails in Plexiglas, manually adjustable and with a semi-automatic lock mechanism. Side rail height from sleep deck min 35cm
| 10 | Base frame with easy to clean plastic cover, mounted on 4 castors with simultaneous and directional lock
| 11 | Fitted with a cable wind-up support and a back-up battery to operate the bed in the absence of mains power
| 12 | Trendelenburg: max 15 degrees.
| 13 | Reverse Trendelenburg: max 15 degrees.
| 14 | Commands to activate BACK, LEG, HI-LO, TILT. Command lock key. Pre-configured synchronized commands including AUTOCONTOUR, EXAM, FOWLER, SEAT, FAINT positions. Emergency one-touch commands CPR and SHOCK. Visual display Handset position at head end of bed
| 15 | Mattress
| 16 | Pull-out bar at the head end
| 17 | IV pole with 4 hooks
| 18 | Vertical oxygen bottle holder
3.11 |  | Monitor for measurement hemodynamic parameters |  | 2
| 1 | Derived cardiac output from arterial pressure waveform based on »Pulse Power« analysis
| 2 | Validated techology with vasoactive drugs usage
| 3 | Touch screen, minimum size 12 inch
| 4 | Minimally invasively continuous measurement of hemodynamic patients with the existing arterial line and non-invasive continuos measurement via cuff controler and finger cufs (at least 2)
| 5 | Display of hemodinamic parameters on display (graph and numerical) : CO, SVR, SV, MAP, HR, BIS
| 6 | Absolute or indexed values
| 7 | Stroke volume variation (SVV) and pulse pressure variotion (PVV) numerical and trend values of chosen indicator
| 8 | Numeric value for heart rate variation (HRV)
| 9 | Hemodynamic baseline for pre-induction blood pressure, stroke volume and cardiac output values, long term and short term trend data for parameters
| 10 | Event response – response of hemodynamic parameters on inotrops or fluid challenge
| 11 | Event markers
| 12 | Patient data stored for 6 months
| 13 | Data can be tranfered to computer via USB
| 14 | Program for viewing patient data on computer
| 15 | Monitor uses disbosable cards for unlocking monitoring functions for non-invasive or minimal invasive hemodynamic measurement
| 16 | Possibility to add BIS module for monitoring depth of anaesthesia
3.12 |  | Infusion pump |  | 40
| 1 | Volumetric infusion pump for Codan standard and czto IV sets and other validated sets made of PVC, PVC-free, No Dehp.
| 2 | Aplications: Intensive care, anesthesia, standard care, neonatology, IV administration of fluids, citostatics, oncology, parenteral nutrition, blood transfusion
| 3 | Intuitive and symple to use
| 4 | High accuracy ±5% depending on the used IV system,
| 5 | Flow rate 0,1-1200ml/h in steps of 0,1ml/h, modification without infusion interruption. Automatic rate calculation on the basis of volume and time.
| 6 | Adjustable infusion volume 0,1-9999ml (0,1 - 999,9 ml in steps of 0,1ml from 1 - 9999 in steps of 1 ml),
| 7 | Adjustable infusion time 1 min 99h 59 min. In steps of 1 min.
| 8 | Dosage units: g, mg, ug, ng, IU, mmol, hour, min., kg, day
| 9 | Provision to limit the flow rate according to the medication name with optional DERS system.
| 10 | Purging of the IV system at the rate up to 1200ml/h
| 11 | Bolus manual and automatic
| 12 | Bolus programming: Bolus flow rate 0,1-1200ml/h, Bolus volume 0,1-999ml
| 13 | Drop detector (device can also operate without drop detector)
| 14 | Air bubbles detection: individual bubbles configurable 50-1000µl, air accumulation 100-2000µl
| 15 | KVO 0.1-10.0ml/h configurable; deactivation depending on medication
| 16 | Adjustable sensitivity of the occlusion alarm in 10 steps
| 17 | Automatic pressure decrease after occlusion
| 18 | Graphic display of pressure in IV line (bar graph)
| 19 | Data display in graphic LCD display. Permanent display of rate, volume and infusion time
| 20 | Displaying additional information: menu for programming specific functions and alarms, display of additional parameters and settings (pressure limit, batery status, bolus dose calculation etc.)
| 21 | Operation in the following modes: Flow rate, Volume + Time
| 22 | Provision to operate in the IV stand or on its own
| 23 | Connectable to the docking station, connection to the PDMS system, RS-232 interface
| 24 | Transport function
| 25 | Key lock
| 26 | Memory for more than 1000 events
| 27 | Alarms: Visual alarm in red color, discernable from great distance + dispalyed message on the alarm type, accustic alarm with adjustable volume
| 28 | Pump adjustment and service through computer
| 29 | Mains and battery operated. Battery operation minimum 6 hour at the rate of 25ml/h.
| 30 | Mains adapter 220V 50Hz and charger integrated into the pump housing.
| 31 | Dimensions 190x 160x130mm and smaller
| 32 | Weight up to 2 kg including the battery and charger
| 33 | CE certification
| 34 | Compliance with: EN 60601-1-1, EN 60601-1-4, EN 60601-2-24, ISO 13485, ISO 9001
3.13 |  | Syringe pump |  | 120
| 1 | Syringe infusion pump for standard use, anaesthesia and intensive care
| 2 | Intuitive and very simple to use
| 3 | High accuracy ±2% depending on the syringe type
| 4 | Flow rate 0,1-1100ml/h in steps of 0,1ml/h, modification without infusion interruption. Automatic rate calculation on the basis of volume and time.
| 5 | Adjustable infusion volume 0,1- 999,9 in steps of 0,1 and 1 ml,
| 6 | Adjustable infusion time 1 min - 99h 59 min. In steps of 1 min.
| 7 | Provision to limit the flow rate according to the medication name using an optional DERS system
| 8 | Purging and filling the IV line at the rate up to 1100ml/h
| 9 | Bolus manual and automatic
| 10 | Bolus programming: Bolus flow rate 0,1-1100ml/h, Bolus volume 0,1-50ml depending on the size of the syringe.
| 11 | Able to use syringes of: 10, 20 and 50/60ml
| 12 | Automatic syringe detection on display
| 13 | KVO 0.1-5,0 ml/h configurable; deactivation depending on medication
| 14 | Adjustable sensitivity of the occlusion alarm from 200 to 1200 mbar in 11 steps. Permanent monitoring of the line pressure and permanent display on the screen
| 15 | Automatic pressure decrease after occlusion
| 16 | Data display in graphic LCD display. Permanent display of rate, volume and infusion time
| 17 | Displaying additional information: menu for programming specific functions and alarms, display of additional parameters and settings (pressure limit, batery status, bolus dose calculation etc.)
| 18 | Operation in the following modes: Flow rate, Volume + Time,
| 19 | Optional Medication database with hard and soft dosage limits, provision for dosage calculation. Dose error reduction system - DERS
| 20 | Connectable to the docking station, connection to the PDMS system, RS-232 interface
| 21 | Provision to operate in the IV stand, rail system or be placed on an even surface
| 22 | Key lock
| 23 | Memory for more than 1000 events
| 24 | Pre-alarms. Warnings: infusion near end, syringe almost empty, battery running out. Alarms: Visual alarm in red discernable from great distance + alarm type message on the display, accustic alarm with adjustable volume. Alarms for the end of infusion, occlusion of IV line, badly positioned syringe, empty syringe, wrong rate, wrong concentration, alarm for insufficient dosage when dosing induction in TIVA mode.
| 25 | Pump adjustment and service through computer
| 26 | Mains and battery operated. Battery operation min 9 hours at the flow rate of 5ml/h. Provision to display the remaining time for the battery operation
| 27 | Mains adapter 22V 50Hz and charger integrated into the pump housing.
| 28 | Dimensions 320x 100x 180mm and smaller together with the battery charger
| 29 | Weight up to 2,2kg including the battery and charger
| 30 | CE certification
| 31 | Compliance with: EN 60601-1-1, EN 60601-1-4, EN 60601-2-24, ISO 13485, ISO 9001
3.14 |  | Pulse oximeter |  | 30
| 1 | digital signal processing technology for accurate, reliable readings even during low perfusion and other forms of signal interference
| 2 | continuous SpO2 and pulse rate monitoring, trending data and SatSeconds alarm management
| 3 | compact, bedside design with improved functionality, built in handle and an intuitive, easy-to-use and easy-to-read, color user interface
| 4 | Variable pitch beep tone which enables clinicians to hear pointby- point changes in SpO2
| 5 | Five-hour battery life
| 6 | 96-hour trend memory captured every four seconds
| 7 | Patient trend data can be stored on a PC for archive and analysis
| 8 | On-screen help messages to assist the user in the use of the monitor.
| 9 | Measurement Range SpO2: 1% to 100% Pulse rate: 20 to 250 beats per minute (bpm) Pulse amplitude: 0.03% to 20%
| 10 | Accuracy Saturation (% SpO2 ± 1 SD) Adult: 70% to 100% ± 2 digits Neonate: 70% to 100% ± 3 digits Low perfusion: 70% to 100% ± 2 digits Pulse rate: 20 to 250 bpm ± 3 digits Low perfusion: 20 to 250 bpm ± 3 digits
| 11 | Visual indicators: Pulse search, audible alarms silenced or off, interference indicator, battery charging, and SatSeconds alarm management clock, Pleth
3.15 |  | Hyper-hypothermia system |  | 1
| 1 | Three Blanket connections
| 2 | High Flow rate for rapid temperature change in water
| 3 | Gradient technology to help prevent overshoot
| 4 | Data Export Software
| 5 | Water Temperature Range 4°C – 42°C
| 6 | Patient Temperature Set Point Range 30°C – 40°C
| 7 | Reservoir capacity (7.5 liters)
| 8 | Gradient Technology minimizes fluctuations in water temperature and maintains a stable patient temperature
| 9 | Lite Blanket, Esophageal/Rectal Probe & Mittens and Booties package (10 pcs)
3.16 |  | Nitric oxide delivery system |  | 1
| 1 | neonatal, paediatric or adult application
| 2 | direct connection by ventilator interface
| 3 | optional flow sensor for support of external ventilators of different brands
| 4 | convenient design and comprehensive power supply for intrahospital transport (mains, 12V, battery)
| 5 | dual cylinder management: operation with one or two cylinders with automatic flush and cylinder switching
| 6 | wide-ranging cylinder compatibility (50 – 1000 ppm)
| 7 | fail-safe dosing by secondary dosing system
| 8 | wide dosing range by clocked dosing (13 µl/min – 17 l/min)
| 9 | extended concentration range (0.1 – 99.9 ppm)
| 10 | interface connection to PDMS
| 11 | automatic adjustment for therapeutic maneuvers
| 12 | short NO standing times in the breathing gas and consequently minimization of risks caused by conversion products
3.17 |  | Cerebral/Somatic Oximeter monitor |  | 2
| 1 | Cerebral/Somatic Oximeter monitor a noninvasive oximeter for simultaneously monitor changes in regional blood oxygen saturation in the brain and skeletal muscle (somatic) tissue of the body. It may be used for cerebral oximetry, somatic oximetry or both simultaneously.
| 2 | technology is applicable and approved for use in adult, pediatric and neonatal patients in any clinical setting where the brain and body are at risk of reduced-flow or no-flow ischemic states.
| 3 | four data channels, simultaneously view on monitor
| 4 | manual setting of alarm limits for all channels
| 5 | Range of rSO2 15-95
| 6 | Alarm Limit Range High: 20 – 95; Low: 15 – 90
| 7 | Trend Memory 24 hours at 2 samples per minute
| 8 | 28 cases saved in memory
| 9 | Diagnostics Automatic self-test
| 10 | Input Voltage 100 – 240VAC
| 11 | Frequency 50/60Hz
| 12 | Backup Battery 12VDC; approximately 20 minutes
| 13 | Digital Output RS-232 communications
| 14 | USB Port USB 2.0 Flash Memory
3.18 |  | Phototherapy lamp |  | 2
| 1 | LED phototherapy light system for the treatment of neonatal jaundice
| 2 | The phototherapy light can be placed on an incubator hood, mounted with the spring arm to warming therapy devices or ceiling supplies or used with a trolley
| 3 | Mode of operation Continous
| 4 | Operating conditions:
| 5 | Temperature 18 °C (64.4 °F) to 40 °C (104 °F) Ambient pressure 1,100 hPa to 700 hPa Relative humidity 10 % to 95 % relative humidity, noncondensing
| 6 | Noise level during normal operation ≤20 dB(A)
| 7 | Irradiance
| 8 | 30 cm distance >85,5
| 9 | 40 cm distance >50,1
| 10 | 50 cm distance >33,4
| 11 | Spectral irradiance distribution – Peak between 460–490nm
| 12 | Effective surface area at 40 cm: 30 cm x 50 cm
3.19 |  | Mixer low-flow |  | 6
| 1 | enables a continuous and individually adjusted mixture of medical air and medical oxygen to be administered to infants, children and adults
| 2 | ensures that each patient receives the O2 therapy that is precisely tailored to their needs
| 3 | Oxygen concentrations can easily and reliably be set between the range of 21–100%
| 4 | Low Flow with low flow capacity
| 5 | Main flow outlet 3 – 30 l/min
| 6 | Secondary flow outlet 0 – 27 l/min
| 7 | Bleed flow ≤ 3 l/min at 3.4 bar (49.3 psi)
| 8 | Maximum combined flow (all outlets) ≥ 30 l/min
| 9 | Emergency flow (in the event of the air or oxygen supply failing) >15 l/min
3.2 |  | Mixer high -flow |  | 2
| 1 | enables a continuous and individually adjusted mixture of medical air and medical oxygen to be administered to infants, children and adults
| 2 | ensures that each patient receives the O2 therapy that is precisely tailored to their needs
| 3 | Oxygen concentrations can easily and reliably be set between the range of 21–100%
| 4 | High Flow with high flow capacity
| 5 | Main flow outlet 15 – 120 l/min
| 6 | Secondary flow outlet with bleed 0 – 105 l/min
| 7 | Bleed flow ≤ 13 l/min at 3.4 bar (49.3 psi)
| 8 | Maximum combined flow (all outlets) ≥ 120 l/min
| 9 | Emergency flow (in the event of the air or oxygen supply failing)> 85 l/min
3.21 |  | Mobile video intubation endoscope/bronchoscope |  | 1
| 1 | Flexible endoscope with integrated User interface/monitor, designed for intubation of adults and children in normal conditions or in case of difficult intubation, to assist in intubation with double lumen tubes and broncho blocker positioning as well as for broncho lavage and aspiration.
| 2 | The device must be portable which allows the use in the Operating Rooms, patient rooms, Intensive Care Unit, infirmaries, clinics and ambulance vehicles.
| 3 | The device is equipped with a rechargeable battery with charger for continuous operation for a minimum of 150 minutes.
| 4 | The monitor should be firmly placed on the upper end of the operating part of the endoscope with the possibility of separation when changing the flexible operating part or in case of the device cleaning and disinfection.
| 5 | The monitor can be connected to more different flexible operating parts.
| 6 | The ability to move the monitor, fixed to the operating part, in all directions - up / down and left /right.
| 7 | TFT LCD color touchscreen, with a minimum resolution of 800x480 pxl.
| 8 | The minimum diagonal of the visible image of 10 cm (4 inches).
| 9 | Flexible insertion tube made without an optical cable.
| 10 | The device shall have two flexible operating parts: I. operating part with outer diameter 4.2 mm with working channel diameter minimum 1.2 mm. II. operating part with outer diameter 2.8 mm without working channel.
| 11 | Insertion section working length minimum 600 mm.
| 12 | The possibility of angulation of the bending section in a minimum range upwards 180°/downwards 130 °.
| 13 | The image sensor (camera) integrated into the lower end of the insertion tube of the endoscope.
| 14 | The depth of clear view in the minimum range of 3-50 mm.
| 15 | The possibility of recording a minimum of 9,000 individual images, at a minimum resolution of 384 x 384 pxl.
| 16 | The suitcase also should have necessary accessories for a cleaning, disinfection and a leak tester for checking the correctness of the endoscope.
3.22 |  | Fiberoptic laryngosope |  | 2
| 1 | Handle with replaceable LED bulb which provides white illumination with a LUX output of 123,000
| 2 | Handle manufactured in accordance with ISO 7376 standard
| 3 | Fully autoclaveable metal handle, covered by a 5 year guarantee or 4,000 autoclaves
| 4 | A polished metal collar wich reflects and directs light up towards the fibre core of the blade for maximum illumination
| 5 | Secure screw cap on handle for easy access for batteries
| 6 | LED bulb provides 5000 operating hours and a temperature of 5300 (k)
| 7 | Carved surface of the handle for easier and safer keeping
| 8 | Laryngosope blade made of stainless steel with atraumatic distal end
| 9 | Blades fully autoclaveable
| 10 | The fiber optic core in the blade need to have a mechanism for easy separation and resetting - with one click / move without the use of additional accessories, for easy cleaning or replacement.
| 11 | Possibility of changing the fiber optic core independently of the blade
| 12 | Fiber optic fiber marked in accordance with the size of the spatula for proper installation
| 13 | Laryngoscope should be supplied with: •One handle with LED light - size medium •Three Macintosh reusable spatules of the following sizes: 1. Children 100.2 x 21 mm (size 2) 2. Infant 92 x 21.8 mm (size 1), 3. Neonatus 79 x 15.5 mm (size 0) •Spare LED bulb with a LUX output of 123,000 •A rechargeable battery, Lithium Ion technology •Charger with all the necessary equipment for charging the battery with the colour changing LED indicator which show the health of the battery and the charging cycle. Typically less than 2 hours for a full charge and deliver consistent and reliable power of battery. •Storage box for handle and blades
3.23 |  | Fiberoptic laryngosope for difficult intubations |  | 1
| 1 | Handle with replaceable LED bulb which provides illumination with a LUX output of 123,000
| 2 | Handle manufactured in accordance with ISO 7376 standard
| 3 | Fully autoclaveable metal handle, covered by a 5 year guarantee or 4,000 autoclaves
| 4 | A polished metal collar wich reflects and directs light up towards the fibre core of the blade for maximum illumination.
| 5 | Secure screw cap on handle for easy access for batteries nad easy access
| 6 | LED bulb provides 5000 operating hours and a temperature of 5300 (k)
| 7 | Carved surface of the handle for easier and safer keeping
| 8 | Laryngosope blade for difficult intubations made of antimagnetic stainless steel with atraumatic distal end, and tilting tip for elevation of the epiglottis
| 9 | Blades for difficult intubations fully autoclaveable covered by a 5 year guarantee or 4,000 autoclaves
| 10 | Lever positioned at the rear of the blade for control the tilting tip
| 11 | The non-replaceable glass fibre bundle fully enclosed in the blade which allows easy cleaning
| 12 | Laryngoscope should be supplied with:
| 13 | •One handle with LED light - size medium •Two Macintosh reusable spatules for difficult intubation of the following sizes: 1. Child 115 x 24 mm (size 2) •Spare LED bulb with a LUX output of 123,000 •A rechargeable battery, Lithium Ion technology •Charger with all the necessary equipment for charging the battery with the colour changing LED indicator which show the health of the battery and the charging cycle. Typically less than 2 hours for a full charge and deliver consistent and reliable power of battery. •Storage box for handle and blades
3.24 |  | Portable tablet Ultrasound system for anesthesia |  | 1
| 1 | Portable tablet ultrasound system
| 2 | Console weight max. 6 kg+/-10%
| 3 | Lithium-Ion battery pack
| 4 | Advanced imaging platform with digital beamformer
| 5 | High resolution (min. 1024x768) color LCD min. 12,1"
| 6 | Multi-touch user-interface with gesture recognition
| 7 | Internal solid-state drive (SSD)
| 8 | Data encryption
| 9 | System boot up max. 30 sec
| 10 | /
| 11 | User-define presets
| 12 | Maximum depth of field min. 25 cm (probe dependent)
| 13 | High definition speckle reduction imaging as SRI-HD or corresponding
| 14 | Compounding technique as CrossXBeam or corresponding
| 15 | Automatic tissue optimization
| 16 | TGC
| 17 | Split window
| 18 | Pinch zoom
| 19 | CINE memory
| 20 | Connectors on console: 1 x USB, 1 x SD
| 21 | Docking cart with features including: - height adjustable - handle for maneuverability - large high clearance swivel wheels with locking front wheels - monitor tilt - 2 removable transducer holders
| 22 | Transducer types:
| 22.1 | Linear Array
| 22.2 | /
| 22.3 | /
| 22.4 | Convex Array
| 23 | Operating modes:
| 23.1 | B-Mode
| 23.2 | Color Flow Mode (CFM)
| 23.3 | Power Doppler Imaging (PDI)
| 24 | Modes and options which can be added later:
| 24.1 | M-Mode
| 24.2 | Needle Recognition or corresponding
| 24.3 | /
| 25 | Transducers:
| 25.1 | High-frequency linear array transducer - Bandwidth 8 - 15 MHz (+/3 MHz)
3.25 |  | Nerve Stimulator for Peripheral Regional Anesthesia |  | 1
| 1 | The following information is permanently displayed: - Stimulus amplitude in mA (large digits) - Stimulus duration in ms - Stimulus frequency in Hz - Load impedance in kOhm - Charge delivered in nC (if activated)
| 2 | The enlarged full graphic LC-display
| 3 | Battery: 9V (alkaline)
| 4 | Stimulation frequency: 1 Hz / 2 Hz ±1% / SENSe (3 Hz)
| 5 | Stimulus duration: 0.05 ms – 0.10 ms – 0.30 ms – 0.50 ms – 1.00 ms ±1%
| 6 | SENSe (0.10 ms - 0.10 ms - 0.15 ms to 1.0 ms)
| 7 | Allowable load impedance: 0 kΩ–12 kΩ
| 8 | Current measuring accuracy: ± 0.02 mA
| 9 | Impedance measuring range: 1 kΩ – 90 kΩ for target stimulation current > 0.5 mA
| 10 | Impedance measuring accuracy: ±10% / ±20% for target stimulation current > 1mA / <= 1mA
| 11 | Sound pressure level: 51 dB / 54 dB / 63 dB for stimulation / warning / error
| 12 | Weight: 250g
3.26 |  | Combined monitor for respiratory status |  | 2
| 1 | Compact, transportable monitor for monitoring of respiratory status, for all type of patients (adults, infant, neonates)
| 2 | Monitor must be equipped with color display, minimum 15 cm diagonal
| 3 | AC power and battery power (min. 2 hours of battery operation)
| 4 | Monitor must have an ability to measure, monitor and display EtCO2 based on following principles and characteristics: -EtCO2 sampling possible for intubated and non-intubated patients via plastic sample lines, with or w/o O2 connection, single use, with filters for moisture elimination -Flow rate necessary for sampling: 50 ml/min -Small sampling cell: 15 μl -Sample analysis using CO2 specific wavelengths of IR light -Algorithm unaffected by presence of other gases (O2, NO2, inhaled anesthetics…) -EtCO2, FiCO2 range 0-150mmHg, resolution 1mmHg -CO2 unit : mmHg or kPa or Vol% -CO2 accuracy : 0-38 mmHg ± 2 mmHg, 39-150 mmHg ± 5% + 0,08% for every mmHg above 38 mmHg -EtCO2 results display – capnography – numeric (tabular) and graphic -Calibration interval : no need for individual kalibration (null) for every patient, initially after 1200 operating hours, then after 4000 operating hours or once a year
| 5 | Monitor must be able to detect, monitor i display of following parameters and index values of respiratory status: -IPI (integrated pulmonary index): index number of ventilation adequacy integrating all meausred parameters (numeric and graphic), 1-10 range -Displaying number of apneas during past 60 minutes -Displaying number of desaturation events during past 60 minutes
| 6 | Required alarms : EtCO2 high and low, no breath, RR high and low, SpO2 high and low, Pulse rate high and low, IPI low, alarm of apnea and desaturation event, adjustable alarm management system for managing SpO2 alarms and reduction of false alarms
| 7 | Integrated USB port
| 8 | Trend storage minimum 48 hours at 30 s resolution
| 9 | Monitor accessories:
| 10 | SpO2 connection cable
| 11 | SpO2 sensor, reusable
3.27 |  | Device for measuring and monitoring of the depth of anesthesia and sedation |  | 4
| 1 | Monitor for noninvasive measuring and monitoring of the depth of anesthesia and sedation
| 2 | Technology based on processed EEG signal
| 3 | Acquisition of one-sided 2-channels EEG signal via adhesive sensor with minimum 4 electrodes
| 4 | Displaying of measured reference with single index number, on a 0-100 scale
| 5 | Compact monitor, up to 2 kg
| 6 | Optional mounting on IV stand
| 7 | Built-in rechargeable battery for min. 30 minutes of work without AC
3.28 |  | Heart Lung machine |  | 1
| 1 | System design is fully modular, to be configured according to the specific clinical needs of the perfusionist, allowing to accomodate a large variety of clinical practices and applications.
| 4 | The pediatric configuration can hold up to one large and two integrated small double roller pumps or alternatively three integrated small pumps, allowing the lowest possible priming circuit design for the benefit of the smallest patients.
| 6 | Cable guides enable a lean configuration and design with cables protected and clearly organized
| 7 | Tubing clamp module for the heads of roller pumps.
| 8 | System panels with 3 to 6 slots accommodate the control and monitoring modules. All display and control modules can be replaced during operation if a fault occurs.
| 9 | Panel contains the display and control modules for all of the monitoring, control and measuring devices
| 10 | C-shaped holder for mast-mounted pumps
| 11 | The temperature monitor allows the simultaneous measurement and display of up to four tempera-tures
| 12 | Measuring and displaying the pressure in the extracorporeal circuit. The display range extends from -200 mmHg to +800 mmHg. The values can be displayed in either mmHg or kPa.
| 13 | Controlling perfusion with constant adjustable pressure (set value)through automatic variation of the pump speed (control mode).
| 14 | Cable for transducer
| 15 | Holder for one transducer
| 16 | Holder for 2 transducers
| 17 | The level sensor is mounted on disposable, adhesive mounting strips which are attached to the reservoir or oxygenator. The sensor can be used with most hardshell oxygenators and reservoirs.
| 18 | In monitoring mode, the sensor is attached to the minimum filling level. If the current blood level drops below the set minimum level, a visual and acoustic alarm is triggered and the pump stops
| 19 | A bubble detector that can be used for 3/8“ tubing. Alternatively, 1/2“ and 1/4“ sensors are available.
| 20 | Bubble Detector Sensor - 1/4" - (6.35 mm)
| 21 | The 3-joint mast holder with a fast clamp connector allows the bubble sensor to be positioned at the tubing system.
| 22 | The operator can choose between two operating modes that can be selected in the menu of the control and display module. The cardioplegia sensor module has its own connectors for a bubble sensor and a pressure transducer
| 23 | The centrifugal pump can be mounted quickly on to a mast and integrated into the available system. The system is delivered complete with drive unit, control panel, connection cable, sensor module for flow measurement, flow sensor and emergency system.
| 24 | The ERC closes the arterial line in a fraction of a second if an alarm is triggered (e.g. retrograde flow, level or bubble alarm), keeping the danger of air delivery to a minimum
| 25 | The clamp closes automatically when the stop link function to the arterial pump is activated, if the latter has been stopped by monitoring functions in case of an alarm or is stopped manually.
| 26 | The mast system extension can be mounted on the left or on the right side of the console and folded prior to transport.
| 27 | Cable holders are used for the secure and correct routing of cables and tubing on the mast system. The holders can be installed anywhere on the mast system. Sets of 6.
| 28 | Cable holder for masts with Ø25mm - 6 clips
| 29 | The hair helps the operator maintain a healthy posture. The ergonomic swivel design is mounted on 5 self-locking castors. The back and seat height are adjustable. It is suitable for the operating theatre and easy to clean
| 30 | A second version of the console lamp uses LEDs.
| 31 | Cover for fast clamp connectors - (set of 6 pieces)
| 32 | Has one water tank and two water circuits. It is operated and monitored from its own control panel
| 33 | Heating-cooling blankets connected to the heater-cooler provide an additional support for regulating the patient‘s blood temperature.
| 34 | Tubing set for Heating-Cooling blankets
| 35 | 1 set of adapters for heating/cooling blankets (2 pieces)
3.29 |  | Autotransfusion device |  | 1
| 1 | Full-color touch screen LCD, TFT (min8") with a large view area
| 2 | Internal memory
| 3 | USB ports for simultaneous download of multiple cases in MS® Excel format
| 4 | Integrated ink printer
| 5 | RS232 ports for external PC connection
| 6 | Dual RBC detector technology
| 7 | Integrated Hct sensor (inlet and outlet Hct)
| 8 | Waste line color indicator
| 9 | The four wheels coupled with the front and rear handles
| 10 | IV pole
| 11 | Automatic start at operator determined target volume in the reservoir
| 12 | Continuous processing of bowls in sequence
| 13 | Automatic stop when reservoir is empty
| 14 | Operating protocols: POPT, PSTD, POST-OP
| 15 | preoperative sequestration protocols (PPP,PRP1, PRP2)
| 16 | HCT 60-65%
| 17 | Wash time approx. 5min
| 18 | Heparin, Protein, Albumin, Potassium removal rate >95%
| 19 | modular vacuum pump with possibility to be used as stand alone
| 20 | 50 to -300 mmHg (Intra-operative mode -High vacuum capacity)
| 21 | minus (-) 10 to minus (-)100 mmHg (Post-operative mode Gentle aspiration for wound drainage connection)
| 22 | Availability of disposables in sizes: 55ml, 125ml, 175ml, 225ml
| 23 | Availability of Blood Collection Reservoir with 4 liter capacity and internal defoaming polyurethane layer (30 ppi), Intermediate woven, non-woven filter (40 μm)
| 24 | Availability of Sequestration Set intended for pre-operative sequestration of the blood
3.30 |  | Calibrated Monitor for measurement hemodynamic parameters |  | 1
| 1 | Derived cardiac output from arterial pressure waveform
| 2 | Non-Invasive, Minimally Invasive, Calibrated
| 3 | 3 in 1 platform for monitoring in any clinical situation
| 4 | Validated techology with vasoactive drugs usage
| 5 | Touch screen, minimum size 12 inch
| 6 | Minimally invasively continuous measurement of hemodynamic patients with the existing arterial line and non-invasive continuos measurement via cuff controler and finger cufs (at least 2)
| 7 | Calibrate using Lithium technology or another absolute cardiac output measurement value
| 8 | Reduced infection risk with less invasive catheters
| 9 | Display of hemodinamic parameters on display (graph and numerical) : CO, SVR, SV, MAP, HR, BIS
| 10 | Absolute or indexed values
| 11 | Stroke volume variation (SVV) and pulse pressure variotion (PVV) numerical and trend values of chosen indicator
| 12 | Numeric value for heart rate variation (HRV)
| 13 | Hemodynamic baseline for pre-induction blood pressure, stroke volume and cardiac output values, long term and short term trend data for parameters
| 14 | Event response – response of hemodynamic parameters on inotrops or fluid challenge
| 15 | Event markers
| 16 | Patient data stored for 6 months
| 17 | Data can be tranfered to computer via USB
| 18 | Program for viewing patient data on computer
| 19 | Monitor uses disbosable cards for unlocking monitoring functions for non-invasive or minimal invasive hemodynamic measurement
| 20 | Possibility to add BIS module for monitoring depth of anaesthesia
3.31 |  | Cardiology Ultrasound Device |  | 1
| 1 | Basic Characteristics:
| 1.1 | Integrated colour LCD or TFT high resolution display diagonal at least 21” on artificial arm with possibility of moving in all directions
| 1.2 | Operating panel with backlit buttons
| 1.3 | QWERTY keyboard
| 1.4 | Operating panel can move up/down, and rotate horizontally
| 1.5 | Buttons on operating panel (not on keyboard) can be programmed according to customer needs
| 1.6 | LCD touch - screen panel diagonal at least 8“ built into operating panel for faster access to device functions.
| 1.7 | Highly movable device with 4 rotating casters weighing less than 150 kg
| 1.8 | At least 3 active probe connectors (not counting connector for pencil probes) , electronic probe switching
| 1.9 | Every offered transducer must have choice of minimum 3 fundamental operating frequencies
| 1.1 | Device has following scan modes: B mode, M mode, color M mode, Color Doppler, PW Doppler with HPRF, POWER Doppler, Tissue Doppler Imaging CW Doppler
| 1.11 | Maximum scanning depth in B mode at least 30 cm
| 1.12 | Simultaneous display in real time of B mode image and Color Doppler image, in display layout side by side and one below another
| 1.13 | Dynamic range minimum 200 dB
| 1.14 | /
| 1.15 | /
| 2 | Advanced functions
| 2.1 | "Tissue Harmonic Imaging" based on newest techniques like: pulse subtraction, pulse phase inversion, wide band pulse inversion coded harmonics or similar
| 2.2 | Spatial image compounding technology for real time noise reduction Real time sonoCT, CROSSX Beam, SieClear, Aplipure, etc.
| 2.3 | Some of technologies for automatic B mode image optimization with one click of a button: TEQ, iScan, QuickScan, ATO or similar
| 2.4 | Automatic PW Doppler base line and velocity scale optimization with one click of a button.
| 2.6 | /
| 2.7 | Autotrace in PW Doppler mode with measurement
| 2.8 | Trapezoid imaging
| 2.9 | Automatic measurement of intima-media
| 2.1 | Possibility to zoom image before and after freezing
| 2.11 | Possibility of choosing which part of the image is zoomed („Spot Zoom“ or similar)
| 2.12 | Possibility to zoom the image without loss of diagnostic imaging quality and increase of granulation („HD Zoom“or similar)
| 2.13 | System has technique for enhanced needle display during biopsies (state the name of technique)
| 2.14 | Package of standard and cardiac measurements and calculations
| 2.15 | /
| 3 | Archiving and networking
| 3.1 | Cine memory
| 3.2 | Prospective and retrospective clip recording
| 3.3 | Archiving of images in DICOM format on HDD, DVD/CD, USB Flash memory
| 3.4 | Export of images in JPEG format and Cine clips in AVI (MPEG-4) format or equivalent PC formats.
| 3.5 | Network connection, minimum speed 1000 Mbps
| 3.6 | Device has at least 3 USB ports
| 3.7 | /
| 3.8 | Users can create personalized reports on system, these reports can be exported to PDF or equivalent PC format or printed on office printer
| 4 | Transducers and accessories
| 4.1 | Sector probe with operating frequencies in range 2,0-4,0 MHz or wider, with field of view of at least 90⁰
| 4.2 | Sector probe for pediatric cardiac exams
| 4.3 | Sector probe for neonate cardiac exams
| 4.4 | Black and white thermal printer
| 4.5 | Possibility to connect TEE probes
3.32 |  | MONITORS FOR MONITORING PATIENT VITAL FUNCTIONS |  | 1
| 1 | MONITORS FOR MONITORING PATIENT VITAL FUNCTIONS
| 1.2 | A stationary monitor for patients of all ages with a suitable wall/rail suspension bracket in the OR. The monitor shall have a color display with panel size min. 15 in, high-resolution or more, with touch screen controls.
| 1.3 | Display: 6 curves or more (selectable) and possibility to upgrade it to min. 8
| 1.4 | Adjustable audio / visual alarm system for all parameters
| 1.5 | Trend storage min. 48 hours
| 1.6 | Possibility of networking LAN and connection to central monitoring system, as well a possibility WLA N to upgrade
| 1.7 | Integrated USB port for the connection of keyboard, mouse or other peripherals
| 1.8 | Monitor with passive cooling (cooling of the monitor without fan)
| 1.9 | Internal power supply with battery life min. 2,5h
| 1.1 | Ambiental light sensor for automatically adjusting the screen intensity based on the environmental light
| 1.11 | Synchronizing port for connection to defibrillator
| 1.12 | Clock has to be present on the screen all the time
| 1.13 | Possibility to view other monitor data from network in real time
| 1.14 | Monitor has the ability to connect at least 3 single-parameter modules besides those with multiples.
| 2 | The monitor shall have algorithms and modules for the following measurements:
| 2.1 | ECG /HR
| 2.2 | ST segment analysis of all 12 ECG leads
| 2.3 | Simultaneous analysis of min. 2 leads for the detection of arrhythmia
| 2.4 | Pacemaker detection
| 2.5 | Respiration –thoracic impedance
| 2.6 | NIBP - non-invasive blood pressure
| 2.7 | SpO2 - saturation(possibility to upgrade to dual SpO2)
| 2.8 | IBP - invasive blood pressure min. 4 channel
| 2.9 | * T - temperature min. 1 channel
| 3 | All the monitors shall have the option of additional (optional) upgrade:
| 3.1 | Measurement and analysis of CO2, O2 and anesthetic gases
| 3.2 | Measurement of continuous cardiac output and hemodynamic parameters ("PICCO")
| 3.3 | Measurement of SvO2 - mixed venous saturation
| 3.4 | Measurement of EEG - multichannel electroencephalograph
| 3.5 | Upgrade of thermal printer with three or more channels
| 3.6 | Possibility of freezing curves and measuring values at the selected points.
| 3.7 | Connection with a barcode reader allowed
| 3.8 | Monitoring of the diagnostic 12-channel ECG with interpretation on the monitor with 6 lead cable
| 3.9 | Non-invasive, continuous and real-time hemoglobin monitoring
| 3.1 | Monitor to have software that can follow resuscitation of the septic patient using the standard protocol
| 4 | Accessories with every monitor:
| 4.1 | ECG cable, 3-wire (1 pc.)
| 4.2 | Neonatal SP NIBP cuff (1 pc.)
| 4.3 | Neonatal Sensor for SpO2, reusable (1 pc.)
| 4.4 | Neonatal NIBP Air Hose(1 pcs.)
| 4.5 | Preattached leadwire electrode (1 pcs.)
| 5 | MODULE WHICH CAN BE USED AS A TRANSPORT MONITOR WHILE STILL MONITORING ALL PARAMETERS-1 pcs
| 5.1 | Color display min. 6 in. or more (touch screen)
| 5.2 | Waveforms can be displayed in landscape and portrait view
| 5.3 | Display: min. 3 waveforms or more (selectable) and possibility to upgrade to min. 5 waveforms
| 5.4 | Adjustable audio / visual alarm system for all parameters
| 5.5 | Battery capacity min. 300 min, for intra-hospital transport
| 5.6 | Module has to be water and dust resistant under min. IP32 protection
| 5.7 | For ease of transport, the weight of the multiparametric transport mode is less than 1,5 kg
| 5.8 | When detaching the multiparametric module, it will ensure the storage of at least 4 hours of trends and patient demographics.
| 5.9 | During disconnection of the multiparametric screen module, it is not necessary to detach the patient's electrodes, sensors and ECG wires
| 6 | OVERVIEW DISPLAY FOR MONITORING VITAL SIGNS FROM MONITOR IN OPERATION ROOM, MINIMUM 32 INCH.
3.33 |  | Patient monitor with transport module |  | 2
| 1 | MONITORS FOR MONITORING PATIENT VITAL FUNCTIONS
| 1.1 | STATIONARY MONITOR
| 1.2 | A stationary monitor for patients of all ages with a suitable wall/rail suspension bracket in the Cardio Surgery department. The monitor shall have a color display with panel size min. 15 in, high-resolution or more, with touch screen controls.
| 1.3 | Display: 6 curves or more (selectable) and possibility to upgrade it to min. 8
| 1.4 | Adjustable audio / visual alarm system for all parameters
| 1.5 | Trend storage min. 48 hours
| 1.6 | Possibility of networking LAN and connection to central monitoring system, as well a possibility WLA N to upgrade
| 1.7 | Integrated USB port for the connection of keyboard, mouse or other peripherals
| 1.8 | Monitor with passive cooling (cooling of the monitor without fan)
| 1.9 | Internal power supply with battery life min. 2,5h
| 1.1 | Ambiental light sensor for automatically adjusting the screen intensity based on the environmental light
| 1.11 | Synchronizing port for connection to defibrillator
| 1.12 | Clock has to be present on the screen all the time
| 1.13 | Possibility to view other monitor data from network in real time
| 1.14 | Monitor has the ability to connect at least 3 single-parameter modules besides those with multiples.
| 2 | The monitor shall have algorithms and modules for the following measurements:
| 2.1 | ECG /HR
| 2.2 | ST segment analysis of all 12 ECG leads
| 2.3 | Simultaneous analysis of min. 2 leads for the detection of arrhythmia
| 2.4 | Pacemaker detection
| 2.5 | Respiration –thoracic impedance
| 2.6 | NIBP - non-invasive blood pressure
| 2.7 | SpO2 - saturation(possibility to upgrade to dual SpO2)
| 2.8 | IBP - invasive blood pressure min. 2 channel
| 2.9 | * T - temperature min. 1 channel
| 2.1 | * Measurement and analysis of CO2, C.O., C.C.O., PICCO- only for two monitors
| 2.11 | * Additional IBP module
| 3 | All the monitors shall have the option of additional (optional) upgrade:
| 3.1 | Measurement and analysis of CO2, O2 and anesthetic gases
| 3.2 | Measurement of continuous cardiac output and hemodynamic parameters ("PICCO")
| 3.3 | Measurement of SvO2 - mixed venous saturation
| 3.4 | Measurement of EEG - multichannel electroencephalograph
| 3.5 | Monitoring of the diagnostic 12-channel ECG with interpretation on the monitor with 6 lead cable
| 3.6 | Non-invasive, continuous and real-time hemoglobin monitoring
| 3.7 | Monitor to have software that can follow resuscitation of the septic patient using the standard protocol
| 4 | Accessories with every monitor:
| 4.1 | ECG cable, 3-wire (3 pc.)
| 4.2 | Neonatal SP NIBP cuff (3 pc.)
| 4.3 | Neonatal Sensor for SpO2, reusable (3 pc.)
| 4.4 | Neonatal NIBP Air Hose(3 pcs.)
| 4.5 | PICCO temperature probe (2 pcs.)
| 4.6 | SP use Infant Airway Adapt (2 pcs.)
| 4.7 | CBL CO temp probe Set (2pcs.)
| 4.8 | CBL CO Cable (2 pcs.)
| 5 | MODULE WHICH CAN BE USED AS A TRANSPORT MONITOR WHILE STILL MONITORING ALL PARAMETERS- 1 pcs.
| 5.1 | Color display min. 6 in. or more (touch screen)
| 5.2 | Waveforms can be displayed in landscape and portrait view
| 5.3 | Display: min. 3 waveforms or more (selectable) and possibility to upgrade to min. 5 waveforms
| 5.4 | Adjustable audio / visual alarm system for all parameters
| 5.5 | Battery capacity min. 300 min, for intra-hospital transport
| 5.6 | Module has to be water and dust resistant under min. IP32 protection
| 5.7 | For ease of transport, the weight of the multiparametric transport mode is less than 1,5 kg
| 5.8 | When detaching the multiparametric module, it will ensure the storage of at least 4 hours of trends and patient demographics.
| 5.9 | During disconnection of the multiparametric screen module, it is not necessary to detach the patient's electrodes, sensors and ECG wires
3.34 |  | ACT analyzer |  | 1
| 1 | Handheld device for measurement High range ACT, Low range ACT,PT/INR and aPTT C20
| 2 | tests require just a single drop of blood to provide prompt results using a sophisticated combination of electromagnet, photodetector and software algorithms.
| 3 | card based tests
| 4 | Color LCD Touch Screen
| 5 | Built-in bar code reader
| 6 | Built-in WiFi and USB
| 7 | Electronic QC
| 8 | Internal battery
| 9 | Data storage on instrument
| 10 | Docking station available for instrument charging and communication link
| 11 | Weight up to 1 kg
3.35 |  | In line Blood Parameter Monitoring System |  | 1
| 1 | Measures or calculates pH, pCO , pO, K+, temperature, SO , hematocrit, hemoglobin, base excess, bicarbonate, oxygen consumption and oxygen
| 2 | Fast two-point gas calibration
| 3 | Mountable onto monitor pole clamp
| 4 | sensor technology measures pH, pCO , pO , and K+.
| 5 | installation of the sensor into the shunt line using luer connections
| 6 | Large, color liquid crystal display
| 7 | Self-diagnostic system
| 8 | visual and audible indicators
| 9 | blood parameter values in either numeric, graphic, or tabular formats
3.36 |  | Transcutaneous bilirubinometer |  | 1
| 1 | Screens infants as young as 24 weeks gestational age
| 2 | Reusable probe just wipe with alcohol and reuse.
| 3 | Light source: Pulse xenon arc lamp; lamp life 150000 measurements
| 4 | Battery Internal number of measurements (when fully charged) up to 250
| 5 | Touch screen display
| 6 | Easy to navigate through menu settings;
| 6.1 | No need for recalibration after every baby
| 7 | Measurement range
| 7.1 | choose to see results in mg/dL or μmol/L;
| 7.2 | measurement range:
| 7.3 | 0.0 mg/dL to 20.0 mg/dL (0 μmol/L to 340 μmol/L)
| 8 | adjust contrast and sensitivity of display
| 9 | Internal memory
3.37 |  | Anesthesia device |  | 5
| 1 | Dedicated for anstesia for neonates, children and adults, without changing components (sensors)
| 2 | Power supply 220VAC and battery (operating capacity for 90 minutes)
| 3 | Color, touch-screen display min. 15“
| 4 | Central gas supply for O2, N2O and Air
| 5 | Backup O2 cylinder (10 lit. capacity) can be mounted on the back of the device
| 6 | Integrated suction unit on (aspirator)
| 7 | Anesthetic gas scavenging system
| 8 | All parts that comes in contact with patient gas can be easily disassembled and sterilized
| 9 | Anesthesia device accessories: reusable patient breathing circuit, connection hoses, breathing bag hose, breathing bag and all necessary accessories needed for normal functioning of device
| 10 | Device should have minimum one drawer for accessories Also device must be portable (mobile) on castors with brake. Device should have writing table. With device should be delivered all belonging components and connections, breathing hoses, breathing bag hoses, connection hoses for central gas supply
| 11 | Device hasoption to predict and display concentration of anestetic agent
| 12 | Breathing system and gas delivery:
| 13 | Breathing circuit system
| 14 | Device has posibility for low-flow and minimal-flow anesthesia
| 15 | Manual and mechanical ventilation
| 16 | Possible emergency manual ventilation in case of power supply failure (external and battery supply)
| 17 | O2 flush command
| 18 | Automatic N2O and anesthetic gas shut-off in case of O2 failure
| 19 | Electronic regulation of O2 concentration in O2 and N2O mixture of 25%
| 20 | Possibility for mechanical ventilation with ambiental air even in case of medical gas supply failure (central gas supply and cylinders) - (electrically driven and electronically controlled ventilator
| 21 | Adjustable valve for pressure control in manual mode of ventilation
| 22 | Simple transition between mechanical and manual ventilation
| 23 | Reusable CO2 absorber, minimal volume 1,5lit
| 24 | Sevoflurane vapor
| 25 | Posibility for instalation of two vapors, with block for simultaneous work of both vapors
| 26 | Posibility for filling vapors during the operation of device
| 27 | Ventilation modes
| 28 | Manual with CPAP
| 29 | Volume controlled ventoilation
| 30 | Presssure controlled ventoilation
| 31 | Pressure support breathing
| 32 | Apnea ventilation
| 33 | Simple transition between ventilation modes
| 34 | Tidal volume: 20ml to minimal 2000ml
| 35 | Breathing frequency: from 3 to minimal 100/min
| 36 | PEEP: 2 to minimal 35 cm Н2О
| 37 | Pinsp (inspiratory pressure): 80 mbar
| 38 | Pressure support: 70 mbar
| 39 | I:Е minimal range: 1:5 до 5:1
| 40 | Adjustable flow trigger 0,3 - 15 l/min
| 41 | Inspiratory flow 180 l/min
| 42 | Monitoring
| 43 | Pressure
| 44 | Volume, inspiratory and expiratory
| 45 | Minute Volume, inspiratory and expiratory
| 46 | Integrated gas analyzer with display for inspiratory and expiratory concentrations of: CO2, N2O, anesthetic gas, MAC values
| 47 | Device can show simultaneously ventilation and gas values numerically, graphically and trends
| 48 | Digital display of medical gas pressures (О2, Air и N2O) for central gas supply and cylinders
| 49 | Logbok, with anesthetic gas consumption
| 50 | Simultaneous display of two loops
| 51 | Graphical or tabelar view of trends or minitrends simultaneously with ventilation curves
| 52 | Calibration and alarms
| 53 | Device has automated selftest wich includes leakage test, compliance test and all sensors calibration
| 54 | Programable self-test, that can be started on programed time
| 55 | Alarms for: gas supply low pressure, power supply failure, battery charge
| 56 | Audio-visual alarms with multi level priority: minute ventilation, leakage,high airway pressure , gas concentration, alarm for sensor failure.
| 57 | Hemodynamic monitor
| 58 | Monitor for all patient categories – neonatal, pediatric and adult
| 59 | Touchscreen high resolution color display. Minimal size 15“
| 60 | Monitor has possibility to navigate around the screen, select and confirm function. Monitoring of the following functions: - pause sound alarm - start NIBP measurements - enter trends - freezing curves on screen - printing on the integrated printer - enter the main menu
| 61 | Monitor can show on display 13 or more curves
| 62 | Maximum weight not above 7 kg (with all modules)
| 63 | Audio-visual alarm system, with centraly positioned alarm lamps
| 64 | Integrated 3-channel recorder with recordin types: ‒ Continuous real-time recording ‒ 8/20 seconds real-time recording ‒ Trend graph recording ‒ Trend table recording ‒ NIBP review recording ‒ Arrhythmia review recording ‒ Alarm review recording ‒ Frozen waveform recording ‒ Drug calculation titration recording
| 65 | Internal battery with capacity of 4 hours of operating
| 66 | Monitor has modules for measurement of following parameters: - ECG - NIBP - SpO2 – saturation - temperature - 3 IBP - BIS - BIS EEG waveform, BIS trend and parameters: BIS, SQI, SR, SEF, TP
| 67 | With monitor following accessories must be delivered: - 3 lead ECG kit and ECG set for neonate - SpO2 reusable sensor for adults and SpO2 for neonate - 3 NIBP cuffs (different sizes) and NIBP cuffs for neonate - Reusable temp sensor and skin sensors for neonate - kit for measurement of 1 IBP