## Price Schedule
Bidder: |  |  |  | Offered price (currency):
LOT 8 | Sterillization
ID | Technical Specification Requested | QTY | Technical Specification Offered (model) | Unit price | Total price per line item
8.1 | Shelf rack, open, 5 shelves, 120x60x200 cm | 1
8.2 | Working table with 1 sink, lower shelf and tap, 200x65x90 cm | 1
8.3 | Set of wall shelves, 200x30x60 cm | 1
8.4 | Universal spray gun for washing and drying | 1
8.5 | Working table with lower shelf, 120x65x90 cm | 1
8.6 | Set of wall shelves, 120x30x60 cm | 1
8.7 | Ultrasonic bath, tank volume 30 liters | 1
8.8 | Working table with 2 sinks, lower shelf and tap, 200x65x90 cm | 1
8.9 | Set of wall shelves, 200x30x60 cm | 1
8.10 | Universal spray gun for washing and drying | 1
8.11 | Working table with lower shelf, 120x65x90 cm | 1
8.12 | Auxiliary table for inner trolley of washing machine | 1
8.13 | Unit for washing of trolleys | 1
8.14 | Trolley for washing machine, 5 levels | 1
8.15 | Machine for washing, disinfection and drying of instruments, chamber volume 230 liters, 5 levels, 10 DIN baskets | 2
8.15A | Set of basic accessories for charging of washing machine, 10 DIN | 1
8.15B | Set of accessories for washing of tubular elements for washing machine, 10 DIN | 1
8.16 | Vertical sliding window, pass through, Aluminum/ AISI 304, 110x65 cm | 1
8.17 | Transport trolley for washing machine, 5 levels | 1
8.18 | Working table with lower shelf, 160x65x90 cm | 1
8.19 | Shelf rack, open, 5 shelves, 120x60x200 cm | 2
8.20 | Working table for packaging, with shelves, 200x80x90 cm | 2
8.21 | Cutting device for sterilisation rolls, 120 cm | 2
8.22 | Sealing device for sterilisation rolls | 2
8.23 | Working chair, adjustable height | 4
8.24 | Transport trolley for medical material | 1
8.25 | Working table for control and sorting of textile, with shelves, 200x125x90 cm | 1
8.26 | Mobile storing and carrier device for sterilisation rolls | 1
8.27 | Working table for sorting and packaging, with shelves, 140x80x90 cm | 1
8.28 | Shelf rack, open, 5 shelves, 120x60x200 cm | 2
8.29 | Trolley for steam steriliser, 8 STU | 2
8.30 | Inner trolley for steam steriliser, 8 STU | 2
8.31 | Automatic steam steriliser, 8 STU, 2 doors pass-through, right-hand service access | 1
8.32 | Automatic steam-formaldehyde steriliser, 8 STU, 2 doors pass-through, left-hand service access | 1
8.33 | H2O2 Plasma Steriliser, 100 liter | 1
8.34 | Trolley for steam steriliser, 8 STU | 1
8.35 | Shelf rack, open, 5 shelves, 120x60x200 cm | 4
8.36 | Vertical sliding window, pass through, Aluminum/ AISI 304, 110x65 cm | 1
8.37 | Water softener | 1
8.38 | Reverse Osmosis (RO) unit | 1
8.39 | Transport trolley, closed, 6 STU | 4
8.40 | Transport trolley, open, 6 STU | 4
8.41 | Sterilisation container, 1 STU, 58x28x26 cm | 10
8.42 | Sterilisation container, 1/2 STU, 58x28x13.5 cm | 10
8.43 | Sterilisation container, 1/2 STU, 28.5x28x26 cm | 10
8.44 | Wire basket for sterilisation, 1 STU, 58x28x26 cm | 50
8.45 | Wire basket for sterilisation, 1/2 STU, 58x28x13.5 cm | 50
8.46 | Bedpan washer | 1
8.47 | Working table, 80x60x90 cm | 1
8.48 | Working table with 1 sink, 100x60x90 cm | 1
8.49 | Working table, 120x60x90 cm | 1
8.50 | Cabinet for drying of hoses | 1
8.51 | Shelf rack, open, 5 shelves, 120x50x200 cm | 1
8.52 | Shelf rack, closed - cabinet type, 5 shelves, 90x50x200 cm | 1
8.53 | Unit for decontamination of working surfaces by means of fogging chemicals without silver ions | 1
8.54 | Unit for rapid neutralisation/ desorption of chemicals used for decontamination of working surfaces | 1
8.55 | Reconstruction works at Central Sterilisation Department - 1st floor - Preparation/ finishing works | 1
8.56 | Reconstruction works at Central Sterilisation Department - 1st floor - Dismount and demolition works | 1
8.57 | Reconstruction works at Central Sterilisation Department - 1st floor - Masonry, insulation, plaster, tiling and flooring works | 1
8.58 | Reconstruction works at Central Sterilisation Department - 1st floor - Painting works, sanitary engineering, sewerage works, electrical installations and mechanical works | 1
8.59 | Reconstruction works at Area for washing and disinfection of incubators - Basement - Preparation/ finishing works | 1
8.60 | Reconstruction works at Area for washing and disinfection of incubators - Basement - Dismount and demolition works | 1
8.61 | Reconstruction works at Area for washing and disinfection of incubators - Basement - Masonry, insulation, plaster, tiling and flooring works | 1
8.62 | Reconstruction works at Area for washing and disinfection of incubators - Basement - Painting works, sanitary engineering, sewerage works, electrical installations and mechanical works | 1
WHITE DISHES WASHING
8.63 | Tray holder trolley | 1
8.64 | Worktable | 1
8.65 | Waste bin with pedal | 1
8.66 | Double-bowl sink | 1
8.67 | Hand operated mixer tap with spray arm | 1
8.68 | Double plastic drain siphon | 1
8.69 | Dishwasher | 1
8.70 | Storage cabinet with sliding doors | 1
8.71 | Wash basin | 1
8.72 | Waste bin with pedal | 1
8.73 | Two level platform trolley | 1
8.74 | Refrigerated table | 1
8.75 | Worktable with single-bowl sink | 1
8.76 | Hand operated mixer tap | 1
8.77 | Plastic drain siphon | 1
8.78 | Refrigerated table | 1
8.79 | Worktable | 1
8.80 | Gravity slicer | 1
8.81 | Vegetable cutter | 1
8.82 | Digital desktop scales with display | 1
8.83 | Planetary mixer | 1
8.84 | Crepe machine | 1
8.85 | Hot Dog machine | 1
8.86 | Cast iron electrical grills | 1
8.87 | Tabletop Bain-Marie | 2
8.88 | Lunch counter with infrared lamps | 1
8.89 | El.fryer, tabletop | 1
8.90 | Neutral worktop | 1
8.91 | El.cookers with 4 plates and oven | 1
8.92 | El. Fry top | 1
8.93 | El. steam- convection oven | 1
8.94 | Base for convection oven | 1
8.95 | El. Pizza oven | 1
8.96 | Base for pizza oven | 1
FOOD DELIVERY LINE
8.97 | Warm showcase | 1
8.98 | Refrigerated showcase | 1
8.99 | Cash desk unit | 1
8.100 | Tubular tray slide | 1
8.101 | Table for trays | 1
STORAGE
8.102 | Beverage Fridge | 1
8.103 | Horizontal freezer | 1
DISHES WASHING
8.104 | Two level platform trolley | 1
8.105 | Prewash table with single-bowl sink | 1
8.106 | Hand operated mixer tap with spray arm | 1
8.107 | Single plastic drain siphon | 1
8.108 | Waste bin with pedal | 1
8.109 | Hood-type dishwasher | 1
8.110 | Unloading table from dishwasher | 1
8.111 | Double-bowl sink | 1
8.112 | Hand operated mixer tap | 1
8.113 | Double plastic drain siphon | 1
8.114 | Worktable | 1
8.115 | Dishwasher | 1
8.116 | Table-Machine stand | 1
8.117 | Worktable with single-bowl sink | 1
8.118 | Hand operated mixer tap | 1
8.119 | Plastic drain siphon | 1
8.120 | Waste bin with pedal | 1
8.121 | Storage cabinet with sliding doors | 4
VEGETABLES AND MEAT PREPARATION
8.122 | One level platform trolley | 1
8.123 | Worktable | 2
8.124 | Waste bin with pedal | 2
8.125 | Double-bowl sink | 1
8.126 | Hand operated mixer tap | 1
8.127 | Double plastic drain siphon | 1
8.128 | Wash basin | 2
8.129 | Worktable | 1
8.130 | Sink trolley for vegetables | 1
8.131 | Vegetable peeler | 1
8.132 | UV Knife Sterilizer | 2
8.133 | Worktop cupboard with sliding doors | 1
8.134 | Worktable | 1
8.135 | Worktop cupboard with sliding doors | 2
8.136 | Refrigerated table | 2
8.137 | Worktable with single-bowl sink | 1
8.138 | Hand operated mixer tap | 1
8.139 | Plastic drain siphon | 1
8.140 | Chopping block | 1
8.141 | Worktable | 1
8.142 | Planetary mixer | 1
8.143 | Vegetable cutter | 1
8.144 | Cutter | 1
8.145 | Meat mincer | 1
8.146 | Digital desktop scales with display | 2
8.147 | Magnetic knife rack, L=45 cm | 2
THERMAL BLOCK AND BAKERY / PASTRY PREPARATION
8.148 | Worktop cupboard with sliding doors | 1
8.149 | Worktable | 1
8.150 | Ingredient bin | 1
8.151 | Wash basin | 2
8.152 | Waste bin with pedal | 2
8.153 | Planetary mixer | 1
8.154 | Transportation trolley | 1
8.155 | Transportation trolley | 2
8.156 | El.steam-convection oven | 1
8.157 | Open stand with rach supports | 1
8.158 | Automatic water softener , 8 l. | 1
8.159 | Three level oven | 1
8.160 | Worktable with single-bowl sink | 1
8.161 | Hand operated mixer tap | 1
8.162 | Plastic drain siphon | 1
8.163 | Worktable | 2
8.164 | Worktable | 1
8.165 | El. cookers with 4 plates and oven | 2
8.166 | Steam boiling pan 100 l | 4
8.167 | Neutral worktop | 1
8.168 | El. bratt pan 80 l. | 1
ISSUE FOOD
8.169 | Worktop cupboard with sliding doors | 3
STORAGE OF BREAD
8.170 | Bread slicer | 1
8.171 | Worktable | 1
WAREHOUSE
8.172 | Digital refrigerator 1400 l | 4
FOOD CABINET
8.173 | Peristaltic enteral feeding pump | 5
8.174 | Peristaltic enteral feeding pump | 5
8.175 | Body composition analyzer | 1
WASHING AND DRYING
8.176 | Washer extractor | 1
8.177 | Washer extractor | 2
8.178 | Tumble dryer | 1
8.179 | Tumble dryer | 2
|  |  |  | Total Price per Lot | 0

## Technical specifications
Bidder:
LOT 8 |  | Sterillization |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | Unit Price DAP | Total Price per line item | Insert page no. in technical documentation
Sterilization
8.1 |  | Shelf rack, open, 5 shelves, 120x60x200 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | non-perforated shelves, 5 pcs
| 3 | shelf dimension 120x60 cm (+/- 5 cm)
| 4 | shelf rack height 180 - 200 cm
8.2 |  | Working table with 1 sink, lower shelf and tap, 200x65x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | 1 sink, dimension 50x40x20 cm
| 3 | combined tap for cold/ hot water
| 4 | lower shelf at 30 cm height from floor
| 5 | dimensions 200x65x90 cm (+/- 5 cm)
8.3 |  | Set of wall shelves, 200x30x60 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | non-perforated shelves, 2 pcs
| 3 | height adjustable shelves
| 4 | shelf dimension 200x30 cm (+/- 5 cm)
8.4 |  | Universal spray gun for washing and drying |  | 1
| 1 | water and compressed air connection
| 2 | thermally insulated handle of the spray gun
| 3 | hose, min. 1 m length
| 4 | nozzles, 8 pcs
| 5 | wall mount set
8.5 |  | Working table with lower shelf, 120x65x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | lower shelf at 40 cm height from floor
| 3 | dimensions 120x65x90 cm (+/- 5 cm)
8.6 |  | Set of wall shelves, 120x30x60 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | non-perforated shelves, 2 pcs
| 3 | height adjustable shelves
| 4 | shelf dimension 120x30 cm (+/- 5 cm)
8.7 |  | Ultrasonic bath, tank volume 30 liters |  | 1
| 1 | tank made of stainless steel AISI 304
| 2 | tank volume 30 l
| 3 | depth of tank min. 18 cm
| 4 | ultrasound frequncy 35-40 kHz
| 5 | power connection 220 V, 50 Hz
| 6 | total power max. 2 kW
| 7 | weight of unit max. 15 kg
| 8 | dimensions 60x35x35 cm (+/- 5 cm)
8.8 |  | Working table with 2 sinks, lower shelf and tap, 200x65x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | 2 sinks, dimension 50x40x20 cm
| 3 | combined tap for cold/ hot water
| 4 | lower shelf at 30 cm height from floor
| 5 | dimensions 200x65x90 cm (+/- 5 cm)
8.9 |  | Set of wall shelves, 200x30x60 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | non-perforated shelves, 2 pcs
| 3 | height adjustable shelves
| 4 | shelf dimension 200x30 cm (+/- 5 cm)
8.10 |  | Universal spray gun for washing and drying |  | 1
| 1 | water and compressed air connection
| 2 | thermally insulated handle of the spray gun
| 3 | hose, min. 1 m length
| 4 | nozzles, 8 pcs
| 5 | wall mount set
8.11 |  | Working table with lower shelf, 120x65x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | lower shelf at 40 cm height from floor
| 3 | dimensions 120x65x90 cm (+/- 5 cm)
8.12 |  | Auxiliary table for inner trolley of washing machine |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | carrier of inner trolley of washing machine
| 3 | height 70-90 cm
| 4 | load min. 50 kg
| 5 | dimensions appropriate for inner trolley
8.13 |  | Unit for washing of trolleys |  | 1
| 1 | unit for washing of trolleys and other equipment
| 2 | tank for detergents
| 3 | dosing system for detergents
| 4 | hose, min. 10 m length
| 5 | adjustable nozzle
8.14 |  | Transport trolley for washing machine, 5 levels |  | 1
| 1 | trolley for storage and transport of inner trolley of washing machine
| 2 | positioning elements for inner trolley
| 3 | material: stainless steel, AISI 304 and/ or other appropriate materials
| 4 | 4 wheels
| 5 | dimensions appropriate for inner trolley of washing machine
8.15 |  | Machine for washing, disinfection and drying of instruments, chamber volume 230 liters, 5 levels, 10 DIN baskets |  | 2
| 1 | to be used for washing, disinfection and drying of surgical instruments, hollow elements, bottles, boxes, anaesthesia accessories etc.
| 2 | rotational spray arms at top and bottom of chamber
| 3 | inner trolley equipped with rotational spray arms, which are not fixed to the inner trolley
| 4 | independant temperature and validation test port
| 5 | PLC controlled with microprocessor, battery for back-up, with analogue and digital connections
| 6 | control system controls all operative functions and has audio and visual alarms, min. 6 programs which are started with 6 separate push buttons on the control panel
| 7 | Control panel above chamber with LCD screen and control buttons and/ or touch screen
| 8 | control panel displays all informtions about the cycle and alarms
| 9 | choosing of programs by buttons or by touch-screen
| 10 | possibility to store 20 programs
| 11 | chamber capacity: 5 levels, i.e. 10 DIN baskets (480x250x50 mm)
| 12 | 2 doors, pass-through model, with hinges
| 13 | chamber volume: 220 liters (+/- 5 %)
| 14 | chamber material: AISI 316L or better
| 15 | chamber doors: doors open by rotation on bottom edge, so the doors can be used as auxiliary table for inner trolley
| 16 | doors made mainly of glass
| 17 | light inside the chamber
| 18 | heating of water for washing by electrical heaters
| 19 | 2 dosing pumps for detergents
| 20 | system for control of remaining detergent and flow of detergents on both pumps. Containers for chemicals are placed inside the washing machine, below the device chamber in a dedicated cabinet with its door
| 21 | thermal disinfection at temperatures above 90*C including AO 600 value calculated (90*C, 1 min), max. water temperature 92*C,
| 22 | equipped with system for cooling of waste water, max. allowed temperature of waste water 70*C
| 23 | drying by hot air generated in washing machine by electrical heaters
| 24 | circulation of air by drying fan
| 25 | HEPA filter controlled by differential pressure measurement
| 26 | equipped with communication module for connection to external modules or for connection to system for traceability of instruments
| 27 | integrated printer of process parameters at front panel
| 28 | connections for cold water, hot water, RO water with R 1/2" and/ or R 3/4"
| 29 | sewerage connection diameter 50 mm below washing machine
| 30 | electrical connection 3x380-400 V, 50 Hz
| 31 | connection for removal of hot air above washing machine
| 32 | heaters for water: max. 10 kW
| 33 | heaters for drying air: max. 3 kW
| 34 | max. working temperature: 92*C
| 35 | max. flow of recirculation pump for water: min. 450 l/min
| 36 | max. flow of waste water pump: min. 50 l/min
| 37 | max. flow of dosing pumps for detergents: min. 100 ml/min
| 38 | dimensions: width max. 700 mm, depth max. 700 mm, height max. 2.100 mm, weight max. 200 kg
| 39 | Standards: EN 15883 (Declaration of Conformity), MDD 93/42/EC, CE marked
| 40 | Washing machine registered as medical device according to local regulations with local authorities
8.15A |  | Set of basic accessories for charging of washing machine, 10 DIN |  | 1
| 1 | to be used for placing medical elements, i.e. for charging of the washing machine, consisting of:
| 2 | instruments rack with 5 shelves for washing on 5 levels, capacity 10 DIN baskets (480x250x50 mm) - 2 pcs
| 3 | rack for washing, disinfection and drying of anaesthesia accessories with capacity of min. 12 anaesthesia hoses with length up to 1.500 mm and inner diameter of 8-30 mm. has at least 8 pieces of spray holders for bellows and / or breathing bags - 1 pc
| 4 | rack for washing of shoes for the OR. Capacity of min. 20 shoes - 1 pc
| 5 | rack for STU containers. Capacity of min. 3 STU boxes (600x300x300 mm) - 1 pc
| 6 | baskets for instruments, DIN, 480x250x50-70 mm - 20 pcs
| 7 | baskets for instruments, 1/2 DIN, 240x250x50-70 mm - 20 pcs
| 8 | fine mesh baskets for small elements, 100x100x60 mm - 2 pcs
8.15B |  | Set of accessories for washing of tubular elements for washing machine, 10 DIN |  | 1
| 1 | rack with appropriate elements for washing of tubular elements (MIS). Capacity of min. 5 elements for LUER connection, min. 10 rigid instruments with holder, min. 10 MIS instruments - 1 pc
| 2 | adjusted for use with washing machine with 5 levels of washing
8.16 |  | Vertical sliding window, pass through, Aluminum/ AISI 304, 110x65 cm |  | 1
| 1 | vertical sliding window
| 2 | implementation into wall
| 3 | aluminum/ stainless steel frame, glass window
| 4 | manual or automatic
| 5 | dimensions 110x65 cm (+/- 10 cm)
8.17 |  | Transport trolley for washing machine, 5 levels |  | 1
| 1 | trolley for storage and transport of inner trolley of washing machine
| 2 | positioning elements for inner trolley
| 3 | material: stainless steel, AISI 304 and/ or other appropriate materials
| 4 | 4 wheels
| 5 | dimensions appropriate for inner trolley of washing machine
8.18 |  | Working table with lower shelf, 160x65x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | lower shelf at 40 cm height from floor
| 3 | dimensions 160x65x90 cm (+/- 5 cm)
8.19 |  | Shelf rack, open, 5 shelves, 120x60x200 cm |  | 2
| 1 | material: stainless steel, AISI 304
| 2 | non-perforated shelves, 5 pcs
| 3 | shelf dimension 120x60 cm (+/- 5 cm)
| 4 | shelf rack height 180 - 200 cm
8.20 |  | Working table for packaging, with shelves, 200x80x90 cm |  | 2
| 1 | material: stainless steel, AISI 304
| 2 | shelves above table, width min. 30 cm, 2 pcs
| 3 | LED light below lower shelf
| 4 | 3 power connections 220 V, 50 Hz
| 5 | castor element with drawers below working surface/ table
| 6 | dimensions 200x80x90 cm (+/- 5 cm)
8.21 |  | Cutting device for sterilisation rolls, 120 cm |  | 2
| 1 | material: stainless steel, AISI 304
| 2 | for storing and cutting of sterilisation rolls
| 3 | 4 carrying and positioning bars for rolls
| 4 | dimensions 120x35x20 cm (+/- 5 cm)
8.22 |  | Sealing device for sterilisation rolls |  | 2
| 1 | microprocessor controlled automatic sealing device to be used for sealing of sterilisation reels
| 2 | possibility to validate sealing machine
| 3 | control of critical parameters according to ISO 11607-2
| 4 | equipped with Ethernet and USB connections
| 5 | possibility to connect device to external printer
| 6 | user interface by means of a colour touch-screen with min. 4" diagonal
| 7 | possibility to record process parameters with 3 programs for quick access
| 8 | automatic stand-by, when not in use
| 9 | slow-start function for reels with bigger dimensions
| 10 | safety system for overheating
| 11 | working temperature: 100-200*C
| 12 | sealing speed: 8-10 m/min
| 13 | sealing force: 50-100 N
| 14 | sealing width: min. 15 mm
| 15 | sealing length: no limitation
| 16 | stand-by time: 1-45 min
| 17 | electrical connection 220 V, 50 Hz
| 18 | dimensions max. 600x300x200 mm
| 19 | Standards: ISO 13485, LV 2006/95/EC, EMC 2004/108/EC
8.23 |  | Working chair, adjustable height |  | 4
| 1 | height adjustable
| 2 | 5 wheels
| 3 | sitting height 60-80 cm
8.24 |  | Transport trolley for medical material |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | 2 shelves
| 3 | 4 wheels, wheel diameter 10-15 cm
| 4 | dimensions 110 (+/-10 cm) x 60 (+/- 10 cm) x 80 cm (+/- 5 cm)
8.25 |  | Working table for control and sorting of textile, with shelves, 200x125x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | integrated partially transparent plastic part in working surface with lamp for illumination installed below working surface
| 3 | lamp 220 V, 50 Hz installed below transparent part of the working surface
| 4 | dimensions 200x125x90 cm (+/- 5 cm)
8.26 |  | Mobile storing and carrier device for sterilisation rolls |  | 2
| 1 | material: stainless steel, AISI 304
| 2 | for storing of sterilisation rolls
| 3 | 4 carrying and positioning bars adjustable for 4 different diameters of rolls
| 4 | mobile, with 4 wheels, where of at least 2 wheels with brakes
| 5 | dimensions 140 (+/-10 cm) x 70 (+/- 10 cm) x 100 cm (+/- 10 cm)
8.27 |  | Working table for sorting and packaging, with shelves, 140x80x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | shelves above table, width min. 30 cm, 2 pcs
| 3 | LED light below lower shelf
| 4 | 3 power connections 220 V, 50 Hz
| 5 | castor element with min. 1 drawer below working surface/ table
| 6 | dimensions 140x80x90 cm (+/- 5 cm)
8.28 |  | Shelf rack, open, 5 shelves, 120x60x200 cm |  | 2
| 1 | material: stainless steel, AISI 304
| 2 | non-perforated shelves, 5 pcs
| 3 | shelf dimension 120x60 cm (+/- 5 cm)
| 4 | shelf rack height 180 - 200 cm
8.29 |  | Trolley for steam steriliser, 8 STU |  | 2
| 1 | trolley for charging/ discharging of steriliser
| 2 | appropriate for steam steriliser 8 STU
| 3 | 4 wheels, 2 with brakes
| 4 | docking mechanism for fixing position to steriliser
| 5 | material stainless steel; handles, wheels made of other appropriate materials
| 6 | adjusted to inner chassis for steriliser with 8 STU
| 7 | length max. 1.700 mm
8.30 |  | Inner trolley for steam steriliser, 8 STU |  | 2
| 1 | used in combination with outer trolley for charging/ discharging of steriliser
| 2 | inner trolley is placed on outer trolley
| 3 | material 316 L
| 4 | trolley equipped with 2 shelves, each has a load capacity of min. 75 kg
| 5 | height adjustable shelves
| 6 | inner trolley capacity of 8 STU (600x300x300 mm)
| 7 | appropriate for steam steriliser with 8 STU
8.31 |  | Automatic steam steriliser, 8 STU, 2 doors pass-through, right-hand service access |  | 1
| 1 | A sterilizer for general-purpose steam sterilization of instruments, textiles and other medical goods at health care facilities. The temperature range is 121-134°C (or wider, including this range). Sterilizer performs sterilization by means of dry-saturated steam. Pre-vacuum, sterilization and post-vacuum drying are main cycle phases. A sterilizer for general-purpose steam sterilization of instruments, textiles and other medical goods at health care facilities. The temperature range is 121-134°C (or wider, including this range). Sterilizer performs sterilization by means of dry-saturated steam. Pre-vacuum, sterilization and post-vacuum drying are main cycle phases.
| 2 | The device is free-standing unit with, or without side panels, positioned between 2 walls. The unit is serviced from the front and right side of the sterilizer. Two sterilizers standing one beside another can use the same service area. The device is free-standing unit with, or without side panels, positioned between 2 walls. The unit is serviced from the front and right side of the sterilizer. Two sterilizers standing one beside another can use the same service area.
| 3 | The unit has two doors (pass-through model), which is vertically sliding (up-down). The doors movement is automatic, achieved by means of pneumatic cylinder. The door plate surface is in the same level with the front sterilizer panel (no depth distance between the door plate and front panel), which enables easy cleaning (hygienic design). The door interlock prevents the door opening during the cycle. The door stops in case of any obstacle during closing (prevents hands squeezing).
| 4 | Sterilization chamber has a rectangular shape with rounded corners (not round chamber) for optimal chamber space usage. Chamber is polished. Internal surface roughness: max. 0,5 µm. Nominal internal chamber volume: 600-650 lit Chamber capacity: 8 STU (STU = standard sterilization unit of nominal size 600x300x300 mm)
| 5 | The sterilizer is equipped with integrated steam generator with electrical heaters. It is positioned bellow or above the chamber. Integrated steam generator installed power: 50-55 kW.
| 6 | The unit has integrated multi-stage (double-stage or better) water-ring vacuum pump for pre and post-vacuum cycle phases.
| 7 | The process is completely controlled by micro-processor control system (PLC). It controls all sterilizer’s functions and monitors its operation. It visually and audibly alerts in case of any malfunction, or process irregularity. The sterilizer has its own independent monitoring and documentation system. This system evaluates the sterilizer operation independently from main Control system and it has its own sensors for temperature and pressure.
| 8 | The user interface (operating panel) is in the form of color touch-screen, positioned above, or beside the sterilization chamber, on the front panel of the machine (display with soft buttons is not acceptable). Touch screen diagonal is minimum 10”. The same touch screen (type, size) is positioned both on loading and unloading side of the sterilizer.
| 9 | The user identification (for accessing sterilizers functions) is done by typing in personal password over the user interface and by using some “contact-less” identification system, such as RFID personal tag, bar-code etc.
| 10 | The external communication is achieved over integrated RS232, or Ethernet port. The unit can be connected to the central information system for collecting process data in a common data base. The sterilizer is equipped with hardware and software which enables connecting the sterilizer to “cloud”- based data base. The users are able to access the data base by using their smart phones or computers and to check certain machines functions and parameters (depending on the access level).
| 11 | Pressure and temperature sensors control the sterilization process. The chamber pressure gauge is mounted on the front sterilizer’s panel.
| 12 | Programs: Sterilizer has following sterilization programs: - Sterilization at 121°C with pre and post vacuum (B-cycle) for wrapped goods – type-tested - Sterilization at 134°C with pre and post vacuum (B-cycle) for wrapped goods – type-tested - Rapid sterilization at 134°C for non-porous, non-wrapped goods (N-cycle) - Pre-heating cycle. The cycle used to warm up the sterilizer after long non-operational period (like the first morning cycle). The time of automatic pre-heating cycle start can be set. - Test cycles: “Bowie and Dick”, “leak test”
| 13 | The unit is equipped with a printer for recording process parameters in alpha-numerical and graph (chart) format. The printer is positioned on the front of the machine. The printer is 4” or bigger. The sterilizer is equipped with USB connection for recording process cycles in PDF file format. The USB memory stick can easily be disconnected from the sterilizer and taken to any PC with printer for printing.
| 14 | Operating Performance: - Softened (cold) water consumption: max. 250 lit per cycle - Highly purified water consumption (RO quality), only if integrated steam generator is activated: max. 20 lit per cycle - Clean steam consumption, only if steam supply from hospital’s central system is activated: max. 20 kg per cycle - Electrical power consumption, if integrated steam generator is activated: max. 15 kWh per cycle - Process time: max. 35 minutes per cycle All performance points above refers (applies) to the steam sterilization cycle which is defined in EN 285 standard: 8 STU containers are placed in the chamber (full chamber), there is 15 kg of metal load in each container, cycle sterilization temperature is 134°C, vacuum level 85 mbar.
| 15 | Utilities/power connections: - Softened (cold) water and purified water (RO quality) - Drainage (to be connected in the floor drain) - Compressed air (6-8 bar provided by user) - Electrical power: 3x400V, 50 Hz
| 16 | Chamber/door surface: stainless steel AISI 316L, or AISI 316Ti Side/front/top panels: stainless materials Frame construction: stainless steel, or other corrosion protected material Process pipes and valves: stainless steel
| 17 | Installation Width (left to right) without service door: max. 100 cm Installation Height (bottom to top): max. 220 cm Installation Depth (front to back): max. 170 cm Internal chamber height: min. 69 cm Chamber loading height (bottom internal chamber surface): 700-800 mm
| 18 | Standards: - Sterilizer is MDD 93/42/EC compliant and has proper CE mark. - Sterilizer chamber and steam generator are manufactured in accordance to PED 97/23/EC directive for pressure vessels and have proper CE mark. In addition, both pressure vessels are manufactured in accordance to local regulations in Serbia. - Sterilizer is EN 285 compliant. - Sterilizer is registered in Serbian Agency for Medical Goods (ALIMS)
8.32 |  | Automatic steam-formaldehyde steriliser, 8 STU, 2 doors pass-through, left-hand service access |  | 1
| 1 | A sterilizer for general-purpose steam sterilization of instruments, textiles and other medical goods at health care facilities. The temperature range is 121-134°C (or wider, including this range). Sterilizer performs sterilization by means of dry-saturated steam. Pre-vacuum, sterilization and post-vacuum drying are main cycle phases. In addition, sterilizer has the option for sterilization of heat-sensitive materials by formaldehyde in temperature range 55-80°C (or wider, including this range).
| 2 | The device is free-standing unit with, or without side panels, positioned between 2 walls. The unit is serviced from the front and left side of the sterilizer. Two sterilizers standing one beside another can use the same service area.
| 3 | The unit has two doors (pass-through model), which is vertically sliding (up-down). The doors movement is automatic, achieved by means of pneumatic cylinder. The door plate surface is in the same level with the front sterilizer panel (no depth distance between the door plate and front panel), which enables easy cleaning (hygienic design). The door interlock prevents the door opening during the cycle. The door stops in case of any obstacle during closing (prevents hands squeezing).
| 4 | Sterilization chamber has a rectangular shape with rounded corners (not round chamber) for optimal chamber space usage. Chamber is polished. Internal surface roughness: max. 0.5 µm. Nominal internal chamber volume: 600-650 lit Chamber capacity: 8 STU (STU = standard sterilization unit of nominal size 600x300x300 mm)
| 5 | The sterilizer is equipped with integrated steam generator with electrical heaters. It is positioned bellow or above the chamber. Integrated steam generator installed power: 50-55 kW.
| 6 | The unit has integrated multi-stage (double-stage or better) water-ring vacuum pump for pre and post-vacuum cycle phases.
| 7 | The process is completely controlled by micro-processor control system (PLC). It controls all sterilizer’s functions and monitors its operation. It visually and audibly alerts in case of any malfunction, or process irregularity. The sterilizer has its own independent monitoring and documentation system. This system evaluates the sterilizer operation independently from main Control system and it has its own sensors for temperature and pressure.
| 8 | The user interface (operating panel) is in the form of color touch-screen, positioned above, or beside the sterilization chamber, on the front panel of the machine (display with soft buttons is not acceptable). Touch screen diagonal is minimum 10”. The same touch screen (type, size) is positioned both on loading and unloading side of the sterilizer.
| 9 | The user identification (for accessing sterilizers functions) is done by typing in personal password over the user interface and by using some “contact-less” identification system, such as RFID personal tag, bar-code etc.
| 10 | The external communication is achieved over integrated RS232, or Ethernet port. The unit can be connected to the central information system for collecting process data in a common data base. The sterilizer is equipped with hardware and software which enables connecting the sterilizer to “cloud”- based data base. The users are able to access the data base by using their smart phones or computers and to check certain machines functions and parameters (depending on the access level).
| 11 | Pressure and temperature sensors control the sterilization process. The chamber pressure gauge is mounted on the front sterilizer’s panel.
| 12 | Programs: Sterilizer has following sterilization programs: - Sterilization at 121°C with pre and post vacuum (B-cycle) for wrapped goods – type-tested - Sterilization at 134°C with pre and post vacuum (B-cycle) for wrapped goods – type-tested - Rapid sterilization at 134°C for non-porous, non-wrapped goods (N-cycle) - Sterilization at 50-55°C with formaldehyde for heat-sensitive wrapped goods – type-tested - Sterilization at 60-65°C with formaldehyde for heat-sensitive wrapped goods – type-tested - Sterilization at 80-85°C with formaldehyde for heat-sensitive wrapped goods – type-tested - Pre-heating cycle. The cycle used to warm up the sterilizer after long non-operational period (like the first morning cycle). The time of automatic pre-heating cycle start can be set. - Program for fast adaptation (temperature decrease) from high temp sterilization (121-134°C) to low-temperature cycles (55-80°C), with jacket cooling. - Test cycles: “Bowie and Dick”, “leak test”
| 13 | Formaldehyde application: Formaldehyde is in the liquid stage inside the bottle or similar container, which is positioned in the compartment within the sterilizer before the cycle start. The sterilizers which use solid stage formaldehyde (tablets, powder, paraformaldehyde etc.) are not acceptable. At the end of the cycle the device will suck out and remove all remaining formaldehyde out of the bottle, so it can safely be disposed.
| 14 | The unit is equipped with a printer for recording process parameters in alpha-numerical and graph (chart) format. The printer is positioned on the front of the machine. The printer is 4” or bigger. The sterilizer is equipped with USB connection for recording process cycles in PDF file format. The USB memory stick can easily be disconnected from the sterilizer and taken to any PC with printer for printing.
| 15 | Operating Performance: - Softened (cold) water consumption: max. 250 lit per cycle - Highly purified water consumption (RO quality), only if integrated steam generator is activated: max. 20 lit per cycle - Clean steam consumption, only if steam supply from hospital’s central system is activated: max. 20 kg per cycle - Electrical power consumption, if integrated steam generator is activated: max. 15 kWh per cycle - Process time: max. 35 minutes per cycle All performance points above refers (applies) to the steam sterilization cycle which is defined in EN 285 standard: 8 STU containers are placed in the chamber (full chamber), there is 15 kg of metal load in each container, cycle sterilization temperature is 134°C, vacuum level 85 mbar. - Process time for formaldehyde cycle at 60-65°C: max. 3 hours
| 16 | Utilities/power connections: - Softened (cold) water and purified water (RO quality) - Drainage (to be connected in the floor drain) - Compressed air (6-8 bar provided by user) - Electrical power: 3x400V, 50 Hz
| 17 | Chamber/door surface: stainless steel AISI 316L, or AISI 316Ti Side/front/top panels: stainless materials Frame construction: stainless steel, or other corrosion protected material Process pipes and valves: stainless steel
| 18 | Installation Width (left to right) without service door: max. 130 cm Installation Height (bottom to top): max. 220 cm Installation Depth (front to back): max. 170 cm Internal chamber height: min. 69 cm Chamber loading height (bottom internal chamber surface): 700-800 mm
| 19 | Standards: - Sterilizer is MDD 93/42/EC compliant and has proper CE mark. - Sterilizer chamber and steam generator are manufactured in accordance to PED 97/23/EC directive for pressure vessels and have proper CE mark. In addition, both pressure vessels are manufactured in accordance to local regulations in Serbia. - Sterilizer is EN 285 and EN 14180 compliant. - Sterilizer is registered in Serbian Agency for Medical Goods (ALIMS)
8.33 |  | H2O2 Plasma Steriliser, 100 liter |  | 1
| 1 | unit for sterilisation of thermo sensible medical elements made of metal or other materials
| 2 | the steriliser uses hydrogene-peroxyde as sterilant which is conducted into plasma during the sterilisation process
| 3 | sterilisation temperature 50-60*C
| 4 | self standing cabinet model, with lateral panels
| 5 | 1 door, non pass-through model, the chamber is loaded and unloaded on the same side
| 6 | automatic vertical sliding doors, with pneumatic or electro-mechanic drive
| 7 | chamber volume usable: 100-140 liters
| 8 | horizontal chamber with square cross section with rounded edges
| 9 | chamber material: AISI 316L or better
| 10 | inside the chamber is the RF antenna for generating plasma made of stainless steel 316 L or better. Sterilisers which generate plasma outside the chamber are not allowed.
| 11 | the unit is accesible for service from the front side and from the left or right side, if necessary, certain service activities can be performed from the top of the unit
| 12 | integrated vacuum pump, min. 2 stages
| 13 | integrated HEPA filter, hydrophobicfor air filtration of the air which is introduced into the chamber during a cycle for pressure control inside the chamber
| 14 | integrated catalyser for the decomposition of H2O2, placed between the sterilization chamber and the vacuum pump in order to prevent harmful emissions and for the protection of the vacuum pump
| 15 | integrated concentrator of H2O2 in order to achieve faster cycles and higher efficiency of the cycle by increasing the concentration of H2O2 to 85-90%
| 16 | sterilant is in a bottle or a cartridge and it is positioned before the start of a cycle at the foreseen position at the front of the unit which is easy accessible. The concentration of the H2O2 in the bottle is lower than 60%.
| 17 | 1 bottle/ cartridge is sufficient for at least 10 cycles
| 18 | safety system for opening of doors during cycle and preventing closing of doors in case of any obstacles, i.e. hands of the operator etc.
| 19 | microprocessor controlled PLC system controls all functions of the unit, control of performance/ cycle, alarms and gives visual informations about the actual cycle
| 20 | unit has colour touch-screen with min. 7" diagonal, placed at the front of the unit, above the sterilsation chamber
| 21 | integrated thermal printer for printing process parameters placed beside/ above the door of the sterilisation chamber. paper width min. 2"
| 22 | standard program for the sterilisation of instruments, flexible and rigid lumens with sterilisation time of max. 50 min
| 23 | fast program for surface sterilisation of smaller quantities of material (3-4 kg) with sterilisation time of max. 35 min
| 24 | intensive program especially for longer rigid lumens with sterilisation time of max. 60 min
| 25 | sterilisation of rigid lumens made of stainless steel with inner diameter bigger than 1 mm and length of less than 800 mm
| 26 | sterilisation of flexible lumens made of plastic material with inner diameter bigger than 2 mm and length of less than 1.100 mm
| 27 | electrical connection 220 V, 50 Hz
| 28 | installed power 3.5 kW
| 29 | dimensions: width max. 700 mm, depth max. 900 mm, height max. 2.000 mm
| 30 | with 2 inner sliding shelves (baskets) for the positioning of material to be sterilised
| 31 | Standards applied: MDD 93/42/EC ISO 14937 Plasma steriliser registered as medical device according to local regulations with local authorities
8.34 |  | Trolley for steam steriliser, 8 STU |  | 1
| 1 | trolley for charging/ discharging of steriliser
| 2 | appropriate for steam steriliser 8 STU
| 3 | 4 wheels, 2 with brakes
| 4 | docking mechanism for fixing position to steriliser
| 5 | material stainless steel; handles, wheels made of other appropriate materials
| 6 | adjusted to inner chassis for steriliser with 8 STU
| 7 | length max. 1.700 mm
8.35 |  | Shelf rack, open, 5 shelves, 120x60x200 cm |  | 4
| 1 | material: stainless steel, AISI 304
| 2 | wire/ perforated shelves, 5 pcs
| 3 | shelf dimension 120x60 cm (+/- 5 cm)
| 4 | shelf rack height 180 - 200 cm
8.36 |  | Vertical sliding window, pass through, Aluminum/ AISI 304, 110x65 cm |  | 1
| 1 | vertical sliding window
| 2 | implementation into wall
| 3 | aluminum/ stainless steel frame, glass window
| 4 | manual or automatic
| 5 | dimensions 110x65 cm (+/- 10 cm)
8.37 |  | Water softener |  | 1
| 1 | unit for softening of tap water, type duplex
| 2 | capacity of duplex softener adjusted to capacities of 2 steam sterilisers with 8 STU and 2 washing machines with 10 DIN
| 3 | consists of 2 columns with appropriate resin for demineralisation, automatic control head with control panel and salt-tank for regeneration of columns
| 4 | duplex unit - one column in function, the other column in regeneration mode
| 5 | time or volume controlled
| 6 | tap water connection R 6/4"
| 7 | sewerage connection diameter 50 mm
| 8 | electrical connection 220 V, 50 Hz
8.38 |  | Reverse Osmosis (RO) unit |  | 1
| 1 | unit for demineralisation of softened water
| 2 | capacity of RO unit adjusted to capacities of 2 steam sterilisers with 8 STU and 2 washing machines with 10 DIN
| 3 | quality of treated water appropriate for use with steam sterilisers and washing machines
| 4 | consists of housing with RO membrane, all necessary automation and tank for permeate
| 5 | softened water connection adjusted to outlet connection from water softener
| 6 | sewerage connection diameter 50 mm
| 7 | electrical connection 220 V, 50 Hz
8.39 |  | Transport trolley, closed, 6 STU |  | 4
| 1 | transport of medical sterile or non-sterile material placed in wire baskets or containers with STU dimensions (STU = 600x300x300 mm)
| 2 | closed transport trolley means that the trolley is closed from all sides, in order to prevent cross-contamination
| 3 | 4 wheels with diameter 100-150 mm, 2 with brakes
| 4 | material stainless steel; handles, wheels made of other appropriate materials
| 5 | inside the trolley there are non-fixed L-profiles which allow the positioning of STU wire baskets or containers
| 6 | nominal capacity of closed trolley is 6 STU
| 7 | includes 9 pairs of L-profiles for different positioning possibilities of wire baskets and containers
| 8 | dimensions 950 (+/-100) x 750 (+/-100) x 1.450 (+/-100) mm
8.40 |  | Transport trolley, open, 6 STU |  | 4
| 1 | transport of medical sterile or non-sterile material placed in wire baskets or containers with STU dimensions (STU = 600x300x300 mm)
| 2 | open transport trolley means that the trolley is open from all sides
| 3 | 4 wheels with diameter 100-150 mm, 2 with brakes
| 4 | material stainless steel; handles, wheels made of other appropriate materials
| 5 | the trolley has non-fixed L-profiles which allow the positioning of STU wire baskets or containers
| 6 | nominal capacity of closed trolley is 6 STU
| 7 | includes 9 pairs of L-profiles for different positioning possibilities of wire baskets and containers
| 8 | trolley has block-bar, which prevents the falling out of transported material
| 9 | dimensions 850 (+/-50) x 650 (+/-50) x 1.400 (+/-50) mm
8.41 |  | Sterilisation container, 1 STU, 58x28x26 cm |  | 10
| 1 | for storing and sterilisation of sterile and non-sterile material
| 2 | consisting of bottom and lid, perforated or non-perforated
| 3 | material aluminum
| 4 | with sterile filter
| 5 | with silicone gasket between bottom and lid
| 6 | with handles
| 7 | dimensions 580 (+/-15) x 280 (+/-15) x 260 (+/-15) mm
8.42 |  | Sterilisation container, 1/2 STU, 58x28x13.5 cm |  | 10
| 1 | for storing and sterilisation of sterile and non-sterile material
| 2 | consisting of bottom and lid, perforated or non-perforated
| 3 | material aluminum
| 4 | with sterile filter
| 5 | with silicone gasket between bottom and lid
| 6 | with handles
| 7 | dimensions 580 (+/-15) x 280 (+/-15) x 135 (+/-15) mm
8.43 |  | Sterilisation container, 1/2 STU, 28.5x28x26 cm |  | 10
| 1 | for storing and sterilisation of sterile and non-sterile material
| 2 | consisting of bottom and lid, perforated or non-perforated
| 3 | material aluminum
| 4 | with sterile filter
| 5 | with silicone gasket between bottom and lid
| 6 | with handles
| 7 | dimensions 285 (+/-15) x 280 (+/-15) x 260 (+/-15) mm
8.44 |  | Wire basket for sterilisation, 1 STU, 58x28x26 cm |  | 50
| 1 | for storing and sterilisation of sterile and non-sterile material
| 2 | made of wire, with a raster that prevents falling out of standard material/ standard packages for sterilisation
| 3 | material stainless steel
| 4 | dimensions 580 (+/-15) x 280 (+/-15) x 260 (+/-15) mm
8.45 |  | Wire basket for sterilisation, 1/2 STU, 58x28x13.5 cm |  | 50
| 1 | for storing and sterilisation of sterile and non-sterile material
| 2 | made of wire, with a raster that prevents falling out of standard material/ standard packages for sterilisation
| 3 | material stainless steel
| 4 | dimensions 580 (+/-15) x 280 (+/-15) x 135 (+/-15) mm
8.46 |  | Bedpan washer |  | 1
| 1 | to be used for washing of very non-clean material, i.e. bedpans, suction bottles, bowls, etc.
| 2 | for washing, thermal disinfection and partial drying of material
| 3 | self-standing unit in cabinet version
| 4 | loading of washing chamber from the front or top of the unit
| 5 | microprocessor controlled , with control panel at the front of the unit
| 6 | 3 standard programs
| 7 | chamber volume 70 liters
| 8 | connections to cold water, hot water and sewerage
| 9 | 2 dosing pumps for detergents
| 10 | internal steama generator for final disinfection
| 11 | electrical connection 3x380-400 V, 50 Hz
| 12 | installed power max. 5 kW
| 13 | special program for the washing/ rinsing of the interior of the unit
| 14 | accessories: 3 inner chassis/ adapters for various loads
| 15 | dimensions 600x650x1800 (+/- 100) mm
| 16 | chamber material stainless steel AISI 304
| 17 | pipes and valves: copper, brass, stainless steel
| 18 | panels: stainless steel, appropriate plastic materials
| 19 | standards applied:
EN ISO 13485
MDD 93/42/EEC (CE)
EN 61000-3-2 and EN 50081-2
EN 61010-1 and EN 61010-2
EN 15583
8.47 |  | Working table, 80x60x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | lower shelf at 40 cm height from floor
| 3 | dimensions 80x60x90 cm (+/- 5 cm)
8.48 |  | Working table with 1 sink, 100x60x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | 1 sink, dimension 50x40x20 cm
| 3 | combined tap for cold/ hot water
| 4 | lower shelf at 30 cm height from floor
| 5 | dimensions 100x60x90 cm (+/- 5 cm)
8.49 |  | Working table, 120x60x90 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | lower shelf at 40 cm height from floor
| 3 | dimensions 120x60x90 cm (+/- 5 cm)
8.50 |  | Cabinet for drying of hoses |  | 1
| 1 | cabinet for drying primarily anaesthetic accessories such as hoses and other medical equipment, e.g. drying of bags and bellows and other utensils
| 2 | drying temperatures 70-90*C
| 3 | drying time adjustable 0-90 min
| 4 | 1 door model
| 5 | cabinet has six cassette holders and nine removable shelves
| 6 | processes up to 36 hoses in the same drying cycle
| 7 | program for drying of anaesthetic hoses and bellows
| 8 | program for drying of instruments in trays
| 9 | program for drying of 12 hoses simultaneously and up to 18 SPRI II trays
| 10 | double fans for recirculation of air with 200 m3/h
| 11 | control unit at the front of the unit for control of programs and other settings
| 12 | electrical connection 230 V, 50 Hz, power 3.5 kW
| 13 | dimensions 500x400x1600 mm (+/- 5%)
| 14 | Standards: UL 3101-1; EN 50081-1; EN 50082-2
8.51 |  | Shelf rack, open, 5 shelves, 120x50x200 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | non-perforated shelves, 5 pcs
| 3 | shelf dimension 120x50 cm (+/- 5 cm)
| 4 | shelf rack height 180 - 200 cm
8.52 |  | Shelf rack, closed - cabinet type, 5 shelves, 90x50x200 cm |  | 1
| 1 | material: stainless steel, AISI 304
| 2 | non-perforated shelves, 5 pcs
| 3 | shelf dimension 90x50 cm (+/- 5 cm)
| 4 | shelf rack height 180 - 200 cm
8.53 |  | Unit for decontamination of working surfaces by means of fogging chemicals without silver ions |  | 1
| 1 | mobile unit for disinfection of rooms and material
| 2 | disinfection by means of hydrogenperoxyde fog
| 3 | effective destruction of pathogenes and multiresistant microrganisms
| 4 | reduction rate of micfroroganisms: bacteria and mycobacteria >= 5 log, viruses, yeast and funghi >= 4 log, spores >= 3 log
| 5 | contact time min. 15 min
| 6 | 2 programs for diffusion
| 7 | fast diffusion treating up to 28 m3/min with 2 jets/ nozzles
| 8 | secure installation of chemical bottle with drip-free connector
| 9 | traceability for up to 500 records
| 10 | accurate metering of dispensed solution
| 11 | touch-screen controls
| 12 | USB port for data download
| 13 | light signal at each step of the processing
| 14 | integral printer
| 15 | remote control
| 16 | NF-C tag reader programmer
| 17 | automatic connection to unit for ventilation process after performed decontamination
| 18 | performance validated to standard NF-T 72-281
8.54 |  | Unit for rapid neutralisation/ desorption of chemicals used for decontamination of working surfaces |  | 1
| 1 | small mobile air chemical decontamination unit for oxydising agents
| 2 | fast removal of biocide after automated surface disinfection procedure
| 3 | fast recovery of disinfected rooms, very high chemical removal
| 4 | flow rate 2.400 m3/h
| 5 | remote control
8.55 |  | Reconstruction works at Central Sterilisation Department - 1st floor - Preparation/ finishing works |  | 1
| 1 | PVC foil protection of the floor, furniture, joinery, radiators, etc.
Billing price per m2 - 200 m2
| 2 | Mounting and demounting of the auxiliary scaffold for works in the rooms higher than 2m.
Billing price per m2 - 50 m2
| 3 | Final cleaning of the rooms after finished work.
Billing price per m2 - 120 m2
8.56 |  | Reconstruction works at Central Sterilisation Department - 1st floor - Dismount and demolition works |  | 1
| 1 | Demolishing of partition brick walls, up to 12cm thick, with loading and outfreight of debris to the city waste
Billing price per m2 - 80 m2
| 2 | Punching openings in the brick product wall d=25cm and thicker, with loading and outfreight of debris to the city waste
Billing price per m3 - 1 m3
| 3 | Wall and floor recess cutting for revealing and demounting sanitary engineering pipes
Billing price per m1 - 20 m1
| 4 | Drilling holes with a diamond-drill for the installations, in up to 20cm thick reinforced concrete constructions.
Billing price per one hole.
cross section up to Ø 100mm - 20 pcs
cross section from Ø 100 - 200mm - 5 pcs
| 5 | Cutting grips in the brick wall for the passage of installations.
Billing price per m1 - 60 m1
| 6 | Demounting of the door with a surface of up to 2 m2.
Billing price per one piece - 20 pcs
| 7 | Demounting of the window with a surface over 2 m2.
Billing price per one piece - 25 pcs
| 8 | Demounting of the existing panel partitions.
Billing price per m2 - 28 m2
| 9 | Stripping off the plaster wherever it is damaged.
Billing price per m2 - 30 m2
| 10 | Stripping off floor and wall ceramic tiles.
Billing price per m2 - 60 m2
| 11 | Demolishing - stripping off damaged cement screeded bed to the concrete construction
Billing price per m2 - 120 m2
| 12 | Demounting of the worn off PVC floors, with loading and outfreight of debris to the city waste.
Billing price per m2 - 120 m2
| 13 | Manual loading, manual transportation on the building site to a disposal area or to a truck and outfreight of debris to the city waste.
Billing price per m3 - 5 m3
8.57 |  | Reconstruction works at Central Sterilisation Department - 1st floor - Masonry, insulation, plaster, tiling and flooring works |  | 1
| 1 | Plaster work on the walls where plaster is stripped off, with material procurement. The price includes jamb works around windows and doors, i.e. jambs are not billed separately.
Billing price per m2 - 90 m2
| 2 | Patching of openings up to Ø 200mm after installations have passed through.
Billing price per one piece - 15 pcs
| 3 | Patching grips, on the spot conducted installations, with cement material, material procurement included.
Billing price per m1 - 60 m1
| 4 | Making the rapid-hardening cement screeded bed with the addition of Isomat Screed material or similar, 5 - 6cm thick, material procurement included.
Billing price per m2 - 120 m2
| 5 | Material procurement and jamb works done with cement plaster around 30cm wide.
Billing price per m1 - 100 m1
| 6 | Making horizontal and vertical toilet waterproofing with one of the waterproof polymer-cement coating of the type "SIKA", "MAPEI" and similar, material procurement included.
Billing price per m2 - 15 m2
| 7 | Making and installation of AL profile inner door. Doors at the toilet entrance have to have a butterfly cylinder lock and an occupancy indicator.
Billing price per m2 - 20 m2
| 8 | Making and installation of AL profile windows with thermopane glass d=4+16+4 mm. Minimum billing measure 0,90 m2/pc. Equivalent thermal conductivity coefficient of the inbuilt window is К=1,4W/(m2*K)
Billing price per m2 - 25 m2
| 9 | Making of monolith suspended ceiling out of plaster - cardboard panels d = 12.5 mm at the corresponding sub-construction, material procurement included. Suspending height up to 150 cm.
Billing price per m2 - 120 m2
| 10 | Making plaster-cardboard partition walls d=10cm, at the corresponding sub-construction, with thermal insulation and single plaster-cardboard panel veneer d=12.5mm
Billing price per m2 - 74 m2
| 11 | Material procurement and making wall and ceiling veneer of plaster-cardboard panels d=12mm
Billing price per m2 - 60 m2
| 12 | Procurement and installation of inspections for installations in the plaster-cardboard panel wall and ceiling veneers,
Billing price per one piece - 4 pcs
| 13 | Installing ceramic wall tiles 1st class, of a domestic manufacturer, gluing and pointing with anti-bacterial pointing compound, material procurement included.
Billing price per m2 - 50 m2
| 14 | Procurement and installation of angular aluminium fillets on all external angles where the wall surfaces meet ceramic tiling veneer.
Billing price per m1 - 20 m1
| 15 | Floor ceramic tiling 1st class, domestic manufacture, anti-slip (slip resistance coefficient is P9), gluing or cement plastering, and pointing with anti-bacterial pointing compound, material procurement included.
Billing price per m2 - 16 m2
| 16 | Making plinths by cutting floor tiles, minimum height 15cm оf ceramic tiles 1st class, gluing, and pointing with anti-bacterial pointing compound, material procurement included. Plinths are made out of the same ceramics like the floor.
Billing price per m1 - 22 m1
| 17 | Material procurement and making of layer for reinforcement and fine leveling of the ground with self-leveling compound UZIN NC 151-161 2-3mm thick with previous application of primer UZIN PE 310.
Billing price per m2 - 120 m2
| 18 | Procurement of PVC floor covering material, class tarket optima (or equivalent) 1.5mm thick by gluing on dispersive glue. The cost of works includes milling and welding of joints with PVC electrodes.
Billing price per m2 - 120 m2
8.58 |  | Reconstruction works at Central Sterilisation Department - 1st floor - Painting works, sanitary engineering, sewerage works, electrical installations and mechanical works |  | 1
| 1 | Scraping all layers of old paint from walls and ceilings.
Billing price per m2 - 340 m2
| 2 | Skimming of walls and ceilings in two layers, with installation of edge aluminum fillets for achieving flat surfaces and edges, material procurement included.
Billing price per m2 - 460 m2
| 3 | Procurement of material and painting of walls and ceilings with dispersion paint. Painting is done in all color tones according to User's choice.
Billing price per m2 - 460 m2
| 4 | Painting of heating pipes, with oil paint, material procurement included.
Billing price per m1 - 34 m1
| 5 | Radiator painting, material procurement included.
Billing price per m2 of an extended surface - 32 m2
| 6 | Dismantling of sanitary elements.
Flat - rate calculation - 1 lump sum
| 7 | Dismantling existing water pipes.
Billing price per m1 - 26 m1
| 8 | Dismantling existing sewage pipes.
Billing price per m1 - 15 m1
| 9 | Cutting the plumbing network and dismantling existing MS valves with exhaust with union joints and other galvanized fittings for replacement.
Billing is per joint - closing place Ø to 1'' - 4 pcs
| 10 | Procurement and installation of PVC sewer pipes with fittings.
Billing price per m1
DN 75 mm - 18 m1
DN 50 mm - 14 m1
| 11 | Supply, delivery and installation of PVC floor
drains with inox cover.
Billing price per one piece
DN 75 mm - 4 pcs
| 12 | Supply, delivery and installation of PP-R water pipes, for working pressure of 10 bar, for indoor installation with thermal insulation
Billing price per m1 less than Ø 1'' - 40 m1
| 13 | Procurement, transport and installation of a valve with a cap.
Billing price per one piece
Ø 3/4'' - 3 pcs
Ø 1/2'' - 4 pcs
| 14 | Procurement, transport and installation of an EK valve with a chromium plated cap. The valves must match the PP-R tubes.
Billing price per one piece
Ø 13 mm (Ø 1/2'' / 3/8'') - 8 pcs
| 15 | Testing the water supply network at a test pressure of 10 bar for 30 minutes.
Billing price per m1 - 40 m1
| 16 | Testing of the sewerage network passability and water retention.
Billing price per m1 - 32 m1
| 17 | Purchase and installation of stand-alone WC cups, simplon and baltic.
Billing price per one piece - 2 pcs
| 18 | Procurement and installation of toilet seats
Billing price per one piece
PVC, with nickel-plated hinges - 1 pcs
| 19 | Procurement and installation of a lavatory cistern, with a downcomer
Billing price per one piece - 1 pcs
| 20 | Procurement and installation of wash basins
Billing price per one piece
55cm wide - 1 pcs
| 21 | Procurement and installation of a single-lever battery for hot and cold water for the wash basin.
Billing price per one piece
standing - 1 pcs
| 22 | Dismantling the existing and making the new electrical installation for the equipment and lighting power supply.
Billing price per 1m2 of a surface - 120 m2
| 23 | Dismantling radiators, drainage and re-assembly, with previous system charge, irradiation and commissioning.
Billing price per one piece - 8 pcs
| 24 | Procurement and replacement of the faulty radiator valves, coves, screws
Billing price per set - 8 sets
8.59 |  | Reconstruction works at Area for washing and disinfection of incubators - Basement - Preparation/ finishing works |  | 1
| 1 | PVC foil protection of the floor, furniture, joinery, radiators, etc.
Billing price per m2 - 40 m2
| 2 | Final cleaning of the rooms after finished work.
Billing price per m2 - 40 m2
8.60 |  | Reconstruction works at Area for washing and disinfection of incubators - Basement - Dismount and demolition works |  | 1
| 1 | Punching openings in the brick product wall d=25cm and thicker, with loading and outfreight of debris to the city waste
Billing price per m3 - 2.5 m3
| 2 | Wall and floor recess cutting for revealing and demounting sanitary engineering pipes
Billing price per m1 - 10 m1
| 3 | Drilling holes with a diamond-drill for the installations, in up to 20cm thick reinforced concrete constructions.
Billing price per one hole.
cross section up to Ø 100mm - 4 pcs
cross section from Ø 100 - 200mm - 2 pcs
| 4 | Cutting grips in the brick wall for the passage of installations.
Billing price per m1 - 20 m1
| 5 | Demounting of the door with a surface of up to 2 m2.
Billing price per one piece - 2 pcs
| 6 | Stripping off the plaster wherever it is damaged.
Billing price per m2 - 10 m2
| 7 | Stripping off floor and wall ceramic tiles.
Billing price per m2 - 60 m2
| 8 | Manual loading, manual transportation on the building site to a disposal area or to a truck and outfreight of debris to the city waste.
Billing price per m3 - 4 m3
8.61 |  | Reconstruction works at Area for washing and disinfection of incubators - Basement - Masonry, insulation, plaster, tiling and flooring works |  | 1
| 1 | Plaster work on the walls where plaster is stripped off, with material procurement. The price includes jamb works around windows and doors, i.e. jambs are not billed separately.
Billing price per m2 - 20 m2
| 2 | Patching of openings up to Ø 200mm after installations have passed through.
Billing price per one piece - 4 pcs
| 3 | Patching grips, on the spot conducted installations, with cement material, material procurement included.
Billing price per m1 - 15 m1
| 4 | Making the rapid-hardening cement screeded bed with the addition of Isomat Screed material or similar, 5 - 6cm thick, material procurement included.
Billing price per m2 - 40 m2
| 5 | Making and installation of AL profile inner door. Doors at the toilet entrance have to have a butterfly cylinder lock and an occupancy indicator.
Billing price per m2 - 2.2 m2
| 6 | Installing ceramic wall tiles 1st class, of a domestic manufacturer, gluing and pointing with anti-bacterial pointing compound, material procurement included.
Billing price per m2 - 20 m2
| 7 | Procurement and installation of angular aluminium fillets on all external angles where the wall surfaces meet ceramic tiling veneer.
Billing price per m1 - 8 m1
| 8 | Floor ceramic tiling 1st class, domestic manufacture, anti-slip (slip resistance coefficient is P9), gluing or cement plastering, and pointing with anti-bacterial pointing compound, material procurement included.
Billing price per m2 - 40 m2
8.62 |  | Reconstruction works at Area for washing and disinfection of incubators - Basement - Painting works, sanitary engineering, sewerage works, electrical installations and mechanical works |  | 1
| 1 | Scraping all layers of old paint from walls and ceilings.
Billing price per m2 - 140 m2
| 2 | Skimming of walls and ceilings in two layers, with installation of edge aluminum fillets for achieving flat surfaces and edges, material procurement included.
Billing price per m2 - 140 m2
| 3 | Procurement of material and painting of walls and ceilings with dispersion paint. Painting is done in all color tones according to User's choice.
Billing price per m2 - 140 m2
| 4 | Painting of heating pipes, with oil paint, material procurement included.
Billing price per m1 - 10 m1
| 5 | Radiator painting, material procurement included.
Billing price per m2 of an extended surface - 15 m2
| 6 | Painting of metalwork with oil paint, material procurement and all the necessary prerequisites.
Billing price per m2 - 10 m2
| 7 | Dismantling of sanitary elements.
Billing price flat rate calculation - 1 lump sum
| 8 | Cutting the plumbing network and dismantling existing MS valves with exhaust with union joints and other galvanized fittings for replacement.
Billing is per joint - closing place Ø to 1'' - 2 pcs
| 9 | Procurement and installation of PVC sewer pipes with fittings.
Billing price per m1
DN 75 mm - 15 m1
DN 50 mm - 10 m1
| 10 | Supply, delivery and installation of PVC floor
drains with inox cover.
Billing price per one piece
DN 75 mm - 2 pcs
| 11 | Supply, delivery and installation of PP-R water pipes, for working pressure of 10 bar, for indoor installation with thermal insulation
Billing price per m1 less than Ø 1'' - 20 m1
| 12 | Procurement, transport and installation of a valve with a cap.
Billing price per one piece
Ø 3/4'' - 2 pcs
Ø 1/2'' - 4 pcs
| 13 | Procurement, transport and installation of an EK valve with a chromium plated cap. The valves must match the PP-R tubes.
Billing price per one piece
Ø 13 mm (Ø 1/2'' / 3/8'') - 2 pcs
| 14 | Testing the water supply network at a test pressure of 10 bar for 30 minutes.
Billing price per m1 - 20 m1
| 15 | Purchase and installation of waste sinks with high level cisterne, connection pipe and corresponding grate.
Billing price per one piece - 1 pcs
| 16 | Procurement and installation of a single-lever battery for hot and cold water for the wash basin.
Billing price per one piece
standing - 1 pcs
| 17 | Dismantling the existing and making the new electrical installation for the equipment and lighting power supply.
Billing price per 1m2 of a surface - 40 m2
| 18 | Dismantling radiators, drainage and re-assembly, with previous system charge, irradiation and commissioning.
Billing price per one piece - 2 pcs
| 19 | Procurement and replacement of the faulty radiator valves, coves, screws
Billing price per set - 2 sets
Kitchen
WHITE DISHES WASHING
| 8.63 | Tray holder trolley for 10 trays, with 4 pivoting wheels, 2 of which with brakes, fitted with plastic bumpers, constructed in 304 AISI stainless steel, dim. 50x60x175 cm±5cm |  | 1 pcs
| 8.64 | Worktable with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs dim. 120x50x90 cm±5cm |  | 1 pcs
| 8.65 | Waste bin with pedal with wheels, made in stainless steel AISI304, 50litres capacity, with lid, pedal, mechanism for opening and two handles, dim. Ø39x60 cm |  | 1 pcs
| 8.66 | Double-bowl sink with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop and lower shelf thickness 40mm, with rounded edges, worktop and sink stainless steel thickness 1mm, lower shelf stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 200x70x90 cm |  | 1 pcs
| 8.67 | Hand operated mixer tap with spray arm |  | 1 pcs
| 8.68 | Double plastic drain siphon |  | 1 pcs
| 8.69 | Dishwasher The theoretical capacity of 40/32/28 buckets / hour, as well as the fast program min 65 buckets / hour, Three basic washing programs and an intensive program, a cost-effective program (ECO-silent and economical operation), a silent program, an individual-specific program (adapted to requirements) user) and scale removal program, self-washing program and automatic machine shutdown Dimensions 600x640x725-760mm±5cm, Voltage 380V-415V / 3N, - total connection value 6,0kW / tank heating 2,0kW, cylinder heater 5,4kW - 10A (capacity depends from electrical connection and inlet water temperature) Standard equipment: wire basket for plates - VIII-part, dim 500x500 - wire basket for dishes and utensils - flat, smoke. 500x500 - wire basket for cups / cups - V-piece 500x500mm - PVC accessory basket - Part VIII total connection power 6,9kW, ISO 14001 |  | 1 pcs
| 8.70 | Storage cabinet with sliding doors made in stainless steel AISI 304 18/10, with sliding doors, internal shelves thicknes 40mm, stainless steel thicknes 0.8mm, supporting structure square profiles 40x40x1.2mm, sliding doors on sliders made in double layer stainless steel 0.6mm thickness on the inside and 0.8mm on the outside, with opening handle along entire door height, complete with adjustable PVC legs, dim. 120x60x200 cm |  | 1 pcs
| 8.71 | Wash basin with hand operated mixer tap and siphon, additions for wall mounting, knee operated pedal, upstand height of 100 mm, rounded upstand and basin edges, made in stainless steel AISI 304 18/10, dim. 40x40x32 cm |  | 1 pcs
| 8.72 | Waste bin with pedal with wheels, made in stainless steel AISI304, 50litres capacity, with lid, pedal, mechanism for opening and two handles, dim. Ø39x60 cm |  | 1 pcs
| 8.73 | Two level platform trolley framework in 25mm tubing, with handle and 4 pivoting Ø 125mm wheels, 2 of which with brakes, fitted with plastic bumpers, constructed in 304 AISI stainless steel, dim. 90x60x95 cm±5cm |  | 1 pcs
| 8.74 | Refrigerated table Refrigerated counters feature a single-piece ventilated cooling unit with tropical compressor, easy to remove from its casing. The external working environment can reach a temperature of +42°C with 65% relative humidity. The structure is made of stainless steel (including the external back panel), except for the external bottom panel, which is in galvanized steel. Reversible RHS/LHS opening doors, with automatic closure and opening position fixed at 100°, easily removable magnetic gasket, rounded internal corners to facilitate thorough cleaning. 60mm thick polyurethane insulation, injected with HFO blowing agent at high pressure, without CFCs or HCFCs. Latest generation capacitive digital touch controller, temperature detection using NTC probes. Refrigerant gas R134a used for normal temperatures, R507 (R404a) for low temperatures. Evaporator with catephoric painting treatment. Automatic defrosting and evaporation of water condensate through AISI 316L stainless steel coil. Adjustable and removable guides and grill racks. Adjustable stainless steel feet Professional and compact handles Worktops are in AISI 304 stainless steel, thickness 10/10. Power cord supplied standard without plug. Temperature regime 0/+10°C, 318W power, dim. 130x70x85 cm±5cm |  | 1 pcs
| 8.75 | Worktable with single-bowl sink with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, worktop and sink stainless steel thickness 1mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 150x70x85 cm |  | 1 pcs
| 8.76 | Hand operated mixer tap |  | 1 pcs
| 8.77 | Plastic drain siphon |  | 1 pcs
| 8.78 | Refrigerated table Refrigerated counters feature a single-piece ventilated cooling unit with tropical compressor, easy to remove from its casing. The external working environment can reach a temperature of +42°C with 65% relative humidity. The structure is made of stainless steel (including the external back panel), except for the external bottom panel, which is in galvanized steel. Reversible RHS/LHS opening doors, with automatic closure and opening position fixed at 100°, easily removable magnetic gasket, rounded internal corners to facilitate thorough cleaning. 60mm thick polyurethane insulation, injected with HFO blowing agent at high pressure, without CFCs or HCFCs. Latest generation capacitive digital touch controller, temperature detection using NTC probes. Refrigerant gas R134a used for normal temperatures, R507 (R404a) for low temperatures. Automatic defrosting and evaporation of water condensate through AISI 316L stainless steel coil. Adjustable and removable guides and grill racks. Adjustable stainless steel feet (wheels on request). Professional and compact handles featuring an innovative design. Available without worktop or with upstand. Worktops are always in AISI 304 stainless steel, thickness 10/10. Power cord supplied standard without plug. Temperature regime 0/+10°C, 318W power, dim. 170x70x85 cm±5cm |  | 1 pcs
| 8.79 | Worktable with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 170x70x85 cm±5cm |  | 1 pcs
| 8.80 | Gravity slicer manual gravity slicer, knife diameter 275mm, slice thickness up to 15mm, with cut thickness-adjusting knob (millimeter adjustment), top-mounted stone blade sharpener, made in high polished satin finished anodized aluminium, blade is made of hollow ground hard chrome with blade guard ring, dim. cca 65x50x45 cm±5cm |  | 1 pcs
| 8.81 | Vegetable cutter An appliance for cutting/slicing fruit and vegetables in various shapes and sizes, grating cheese, dry bread, chocolate, nuts; shredding carrots, celery, potatoes, apples, turnips, beetroots and Mozzarella. VEGETABLE CUTTER: AISI 430 stainless steel body and inlet - safety micro switch on the handle - dual product loading inlet. Accessories: cutting discs of various types. Power 0.37kW, 230V 1N 50Hz. Speed 255rpm. Delivers with pack of 5 different cutting discs, CE |  | 1 pcs
| 8.82 | Digital desktop scales with display measurement range up to 20 kg, measurement capacity 2kg/1g, 5kg/2g, 10kg/5g,20kg/10g, with 5-digit LCD display |  | 1 pcs
| 8.83 | Planetary mixer with three accessories: shaker, hook and "G" palette, and they are integral machine parts. fine-tuning speed vessel capacity: 7 lit, power: 0,28kW The INOX mixing vessel is removed for work faster and easier to clean, CE |  | 1 pcs
| 8.84 | Crepe machine cooking top dim. 400mm, power 2,75kw, CE |  | 1 pcs
| 8.85 | Hot Dog machine Three peaks for the heating roll, apparatus for heating and hot dog rolls, made of stainless steel, has a thermostat of 30-90 ° C, a glass container for cooking hot dogs, power 1,10kw, CE |  | 1 pcs
| 8.86 | Cast iron electrical grills Stainless steel body, long lasting, safe to use with ergonomic design. Ribbed plate, with thermostat 0-300ºC, dimensions of heating plate: 350x230mm, power 2,20kw, CE |  | 1 pcs
| 8.87 | Tabletop Bain-Marie capacity 1 GN1 / 1 h 150mm, power 1,20kw, temperature controller 0 - 90 C, stainless steel dim. 37x57x20-22 cm, CE |  | 2 pcs
| 8.88 | Lunch counter with infrared lamps capacity 1 GN1 / 1 h 150mm, with bowl, with two infrared lamps, power 0,25kw, temperature regulator 30 - 90 C, made of stainless steel, CE |  | 1 pcs
| 8.89 | El.fryer, tabletop oil volume: 8lit, pool and lid are made of stainless steel, it has a safety thermostat with reset, knob, thermostat range 50 -180 ° C, oil drain tap, basket dimensions: 22x25x10cm, power 3,0kw, CE |  | 1 pcs
| 8.90 | Neutral worktop made in stainless steel AISI 304, on open base, with rounded edges, worktop stainless steel thicknes 1,2mm, complete with adjustable legs dim. 35x70x85 cm |  | 1 pcs
| 8.91 | El.cookers with 4 plates and oven made of INOX, stainless steel 1.2mm thick, with 4 square plates 220x220mm of 2.6kW in the same plane, power regulators with 7 levels for each ring separately, indicator light when the ring is turned on, fuse for interruption of power supply in case of overheating , oven power: 4.7kW, oven capacity: 4x GN 2/1, oven dimensions: 57,5x65x30cm. Total power: 15.1kW 400V, 3N dim. 80x70x85 cm, CE |  | 1 pcs
| 8.92 | El. Fry top made of stainless steel, the stainless steel of 1.2mm thickness, with flat 2/3 I 1/3 ribbed board thickness of 12mm, from steel Fe510 dim: 74,8x56,4cm, 2 heating zones, and the power control button for each heating zone with , plate slightly tilted to drain grease in collecting drawer, breaker to break in case of overheating, power 7,8kw dim. 80x70x85 cm, CE |  | 1 pcs
| 8.93 | El. steam- convection oven capacity 6 trays 600x400 mm, standard function digital controls, convection baking (30-260 ° C), convection cooking with humidity (30-260 ° C), maximum temp. 260 ° C chamber preheating, 5 humidity levels, 99 programs. Equal temperature at every point in the chamber. Preheat oven. 2 fan speeds, possibility to change fan direction. moisture convection baking technology (20-40-60-80-100%). remove moisture from the baking chamber. 5 min to reach a temperature of 60 ° to 260 ° C. ENERGI STAR certificate. Maximum thermal efficiency and safety at work, the engine stops when the door is opened. Tray holder distance: 75mm Total power: 10.3kV 400V, 3N + water softener, manual dim: 800x830x680mm, CE |  | 1 pcs
| 8.94 | Base for convection oven made in stainless steel AISI 304, open, with GN tray sliders, complete with adjustable legs |  | 1 pcs
| 8.95 | El. Pizza oven Refractory, chamotte baking surfaces, thermal insulation of the sides, visual control of the firing possible through the built-in glass openings in the doors, interior lighting and two thermostats for each floor individually, the possibility of separately regulating the temperature of the upper and lower heaters, PIROMETAR for controlling the temperature for each floor separately. Operating Temperature: 50 ÷ 500 ° C, cooking chamber dimension (mm): 2 x 610x610x140h, Power 8,4kw dim: 900x735x750mm, CE |  | 1 pcs
| 8.96 | Base for pizza oven made in stainless steel AISI 304, open, complete with adjustable legs dim. 80x80x85 cm |  | 1 pcs
FOOD DELIVERY LINE
| 8.97 | Warm showcase made of AISI304 18/10 stainless steel, display unit closed with flat glass and sunroof on the work side, with two glass shelves, 2.5kw power, dim. 120x80x85-135 cm |  | 1 pcs
| 8.98 | Refrigerated showcase made of AISI304 18/10 stainless steel, display unit closed with flat glass and sunroof on the work side, with two glass shelves, in the lower part with built-in unit and one cold box, 1.5kw power dim. 120x80x85-135 cm |  | 1 pcs
| 8.99 | Cash desk unit with drawer and bottom shelf, made of acid-resistant AISI304 18/10 stainless steel, worktop and bottom shelf 40mm thick, supporting structure stainless steel box profiles 40x40x1.2mm, complete with PVC sliding legs, dim. 48x80x85 cm |  | 1 pcs
| 8.100 | Tubular tray slide pipes made of stainless steel AISI304 18/10, the support structure stainless steel channel sections 40x40x1.2mm, dim. 288x30 cm |  | 1 pcs
| 8.101 | Table for trays Made of AISI304 18/10 stainless steel, worktop and bottom shelf 40mm thick, supporting structure stainless steel box profiles 40x40x1.2mm, complete with PVC sliding legs, dim. 55x35x60 cm |  | 1 pcs
STORAGE
| 8.102 | Beverage Fridge capacity.350 lit, with glass door, interior lighting, power 0,5kw, dim. 60x60x190cm |  | 1 pcs
| 8.103 | Horizontal freezer capacity 400 liters, power 0.5kw dim. 130x70x85cm |  | 1 pcs
DISHES WASHING
| 8.104 | Two level platform trolley framework in 25mm tubing, with handle and 4 pivoting Ø 125mm wheels, 2 of which with brakes, fitted with plastic bumpers, constructed in 304 AISI stainless steel, dim. 90x60x94 cm |  | 1 pcs
| 8.105 | Prewash table with single-bowl sink with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, worktop and sink stainless steel thickness 1mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 120x75x85-115 cm |  | 1 pcs
| 8.106 | Hand operated mixer tap with spray arm |  | 1 pcs
| 8.107 | Single plastic drain siphon |  | 1 pcs
| 8.108 | Waste bin with pedal with wheels, made in stainless steel AISI304, 50litres capacity, with lid, pedal, mechanism for opening and two handles, dim. Ø39x60 cm |  | 1 pcs
| 8.109 | Hood-type dishwasher With built-in dispenser for flushing agent and medium Theoretical capacity: 44/32 / 22pc / hour + special program, short program 72 hours / hour With the possibility of adjusting the length of the program depending on the circumstances at the place of installation Value of electricity -(Primary power depending on the conditions on the object) From 400V / 3N- / 50-60Hz, fuse 16A, total power of connection 7,9-9,1kW, 25A signal, total power of connection 10,2-13,2kW or 240V / 1N- / 50-60Hz, fuse 25A, total power of connection 5,1kW Dim.635x750x1500/2000±50mm Dishwashing height 440 mm (dish insert) Hygienic carrier basket, basket size 500x500 mm Thin volume 35lit Required quantity of rinse water 2,0lit Thermal block for hygienic safety (60 ° C temperature, rinse temperature 85 ° C Machine weight 165 / 195kg (net-gross), ISO 14001 |  | 1 pcs
Double-walled hood with automatic start Hygienic tanks made with deep drawing with hygienic heaters Built-in water drain pump, lifting pump pressure washer, leak sensor Standard machine equipment: Elliptical washing fields with special geometric nozzles, adapting the washing pressure to the level of dishwashing, soft start of the washing pump, full flow filter (tank cover, cylinder sieve, suction pump, fine fine filter) , blur sensor, built-in dispensers for washing and rinse aid, active power management, flushing system, special ECO program, special Silence program Touchscreen management - easy handling marked with colors, cycle progress cycle, basic cleaning program, self-cleaning program , a program for removing calculus in the machine, automatically switching on and off the machine at a given time, active signaling of the event, indicating errors or impossibility of correct operation (e.g. Blocking of the washing box or lack of detergents), the level of the kitchen chef is protected by a PIN, a level for a PIN-protected service, an integral log and hygiene, complete with water softener "Monosoft", power 16,5kw
| 8.110 | Unloading table from dishwasher made in acid-resistant stainless steel AISI304, no lower shelf, complete with adjustable PVC legs dim. 120x60x90 cm |  | 1 pcs
| 8.111 | Double-bowl sink with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop and lower shelf thickness 40mm, with rounded edges, worktop and sink stainless steel thickness 1mm, lower shelf stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 140x70x85 cm |  | 1 pcs
| 8.112 | Hand operated mixer tap |  | 1 pcs
| 8.113 | Double plastic drain siphon |  | 1 pcs
| 8.114 | Worktable with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 60x60x85 cm |  | 1 pcs
| 8.115 | Dishwasher The theoretical capacity of 40/32/28 buckets / hour, as well as the fast program 66 buckets / hour, Three basic washing programs and an intensive program, a cost-effective program (ECO-silent and economical operation), a silent program, an individual-specific program (adapted to requirements) user) and scale removal program, self-washing program and automatic machine shutdown Dimensions 600x635x725-760mm, Voltage 380V-415V / 3N, - total connection value 6,0kW / tank heating 2,0kW, cylinder heater 5,4kW - 10A (capacity depends from electrical connection and inlet water temperature) Insert height 404mm Standard equipment: wire basket for plates - VIII-part, dim 500x500 - wire basket for dishes and utensils - flat, smoke. 500x500 - wire basket for cups / cups - V-piece 500x500mm - PVC accessory basket - Part VIII total connection power 6,9kW, ISO 14001 |  | 1 pcs
| 8.116 | Table-Machine stand made of AISI304 18/10 stainless steel, 40x40x1.2mm stainless steel box profiles, complete with PVC sliding legs,m dim. 60x60x55 cm |  | 1 pcs
| 8.117 | Worktable with single-bowl sink with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, worktop and sink stainless steel thickness 1mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 120x70x85 cm |  | 1 pcs
| 8.118 | Hand operated mixer tap |  | 1 pcs
| 8.119 | Plastic drain siphon |  | 1pcs
| 8.120 | Waste bin with pedal with wheels, made in stainless steel AISI304, 50litres capacity, with lid, pedal, mechanism for opening and two handles, dim. Ø39x60 cm |  | 1 pcs
| 8.121 | Storage cabinet with sliding doors made in stainless steel AISI 304 18/10, with sliding doors, internal shelves thicknes 40mm, stainless steel thicknes 0.8mm, supporting structure square profiles 40x40x1.2mm, sliding doors on sliders made in double layer stainless steel 0.6mm thickness on the inside and 0.8mm on the outside, with opening handle along entire door height, complete with adjustable PVC legs, dim. 100x60x180 cm |  | 4 pcs
VEGETABLES AND MEAT PREPARATION
VEGETABLES AND MEAT PREPARATION
| 8.122 | One level platform trolley circular profile frame, with handle and 4 swivel wheels 2 of which with brake, with rubber bumpers, stainless steel AISI304, dim. 120x60x90 cm |  | 1
| 8.123 | Worktable with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 130x70x85 cm |  | 2
| 8.124 | Waste bin with pedal with wheels, made in stainless steel AISI304, 50litres capacity, with lid, pedal, mechanism for opening and two handles, dim. Ø39x60 cm |  | 2
| 8.125 | Double-bowl sink with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop and lower shelf thickness 40mm, with rounded edges, worktop and sink stainless steel thickness 1mm, lower shelf stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 160x70x85 cm |  | 1
| 8.126 | Hand operated mixer tap |  | 1
| 8.127 | Double plastic drain siphon |  | 1
| 8.128 | Wash basin with hand operated mixer tap and siphon, additions for wall mounting, knee operated pedal, upstand height of 100 mm, rounded upstand and basin edges, made in stainless steel AISI304 18/10, dim. 40x40x32 cm |  | 2
| 8.129 | Worktable with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 180x70x85 cm
| 8.130 | Sink trolley for vegetables with handle and 4 swivel wheels 2 of which with brake, with rubber bumpers, made in AISI304 stainless steel, dim. 70x60x60 cm |  | 1
| 8.131 | Vegetable peeler POTATO PEELER An appliance for peeling potatoes and carrots. Stainless steel structure - upper cover in food-safe polycarbonate - safety microswitch on the upper cover and on the outlet - button for the ejection of the food product - timer 0 - 4 minutes - automatic jet of water to favour the elimination of the peel residues - an extractable abrasive plate is included as standard. Max load per cyclus 18kg, hour production 220kg/h, power 1.1kW 400V, 3N, 50Hz, CE |  | 1
| 8.132 | UV Knife Sterilizer wall, capacity 10 blades, power 0,5kw, CE |  | 2
| 8.133 | Worktop cupboard with sliding doors with upstand, made in stainless steel AISI304 18/10, worktop and central shelf thickness 40mm, rounded edges, stainless steel thickness 0.8mm, supporting structure square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, sliding doors made in double layer stainless steel, 0.6mm in thickness on the inside and 0.8mm on the oustide, with opening handles formed along entire door height, complete with adjustable PVC legs, dim. 200x70x85 cm |  | 1
| 8.134 | Worktable with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 200x70x85 cm |  | 1
| 8.135 | Worktop cupboard with sliding doors with upstand, made in stainless steel AISI304 18/10, worktop and central shelf thickness 40mm, rounded edges, stainless steel thickness 0.8mm, supporting structure square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, sliding doors made in double layer stainless steel, 0.6mm in thickness on the inside and 0.8mm on the oustide, with opening handles formed along entire door height, complete with adjustable PVC legs, dim. 230x70x85 cm |  | 2
| 8.136 | Refrigerated table Refrigerated counters feature a single-piece ventilated cooling unit with tropical compressor, easy to remove from its casing. The external working environment can reach a temperature of +42°C with 65% relative humidity. The structure is made of stainless steel (including the external back panel), except for the external bottom panel, which is in galvanized steel. Reversible RHS/LHS opening doors, with automatic closure and opening position fixed at 100°, easily removable magnetic gasket, rounded internal corners to facilitate thorough cleaning. 60mm thick polyurethane insulation, injected with HFO blowing agent at high pressure, without CFCs or HCFCs. |  | 2
Latest generation capacitive digital touch controller, temperature detection using NTC probes. Refrigerant gas R134a used for normal temperatures, R507 (R404a) for low temperatures. Evaporator with catephoric painting treatment. Automatic defrosting and evaporation of water condensate through AISI 316L stainless steel coil. Adjustable and removable guides and grill racks. Adjustable stainless steel feet (wheels on request). Professional and compact handles featuring an innovative design. Worktops are always in AISI 304 stainless steel, thickness 10/10. Power cord supplied standard without plug. Temperature regime 0/+10°C, 318W power, dim. 128x70x85 cm
| 8.137 | Worktable with single-bowl sink with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, worktop and sink stainless steel thickness 1mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 120x70x85 cm |  | 1
| 8.138 | Hand operated mixer tap |  | 1
| 8.139 | Plastic drain siphon |  | 1
| 8.140 | Chopping block PVC tray 10cm thick, stainless steel base parts with adjustable PVC legs, dim. 70x70x85 cm |  | 1
| 8.141 | Worktable with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 90x70x85 cm |  | 1
| 8.142 | Planetary mixer with three accessories: shaker, hook and "G" palette, and they are integral machine parts. fine-tuning speed vessel capacity: 7 lit, power: 0,28kW The INOX mixing vessel is removed for work faster and easier to clean, CE |  | 1
| 8.143 | Vegetable cutter An appliance for cutting/slicing fruit and vegetables in various shapes and sizes, grating cheese, dry bread, chocolate, nuts; shredding carrots, celery, potatoes, apples, turnips, beetroots and Mozzarella. VEGETABLE CUTTER 3000: AISI 430 stainless steel body and inlet - safety micro switch on the handle - dual product loading inlet. Accessories: cutting discs of various types. Power 0.37kW, 230V 1N 50Hz. Speed 255rpm. Delivers with pack of 5 different cutting discs, CE |  | 1
| 8.144 | Cutter Tank capacity 8lit, Apparatus for homogenization, mixing, grinding and mixing, Body made of stainless steel, Lid made of polycarbonate, Pressing machine start buttons, Safety microswitches located on lid and base of pan, Two speeds, Equipped with one double knife, Speed 1400/2800 rpm, Max. pouring into a vessel 2,4lit, power 0,80kw, CE
| 8.145 | Meat mincer capacity 300kg / h, aluminum housing with stainless steel grinder unit, knife, washer and other accessories, power 1,1kw, CE |  | 1
| 8.146 | Digital desktop scales with display measurement range up to 20 kg, measurement capacity 2kg/1g, 5kg/2g, 10kg/5g,20kg/10g, with 5-digit LCD display |  | 2
| 8.147 | Magnetic knife rack, L=45 cm |  | 2
THERMAL BLOCK AND BAKERY / PASTRY PREPARATION
| 8.148 | Worktop cupboard with sliding doors with upstand, made in stainless steel AISI304 18/10, worktop and central shelf thickness 40mm, rounded edges, stainless steel thickness 0.8mm, supporting structure square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, sliding doors made in double layer stainless steel, 0.6mm in thickness on the inside and 0.8mm on the oustide, with opening handles formed along entire door height, complete with adjustable PVC legs dim. 200x70x85 cm |  | 1 pcs
| 8.149 | Worktable with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 200x70x85 cm |  | 1 pcs
| 8.150 | Ingredient bin for dry bulk ingredients, capacity 115 lt,with internal separator, made in acid-resistant stainless steel AISI304 18/10, 4 wheels |  | 1 pcs
| 8.151 | Wash basin with hand operated mixer tap and siphon, additions for wall mounting, knee operated pedal, upstand height of 100 mm, rounded upstand and basin edges, made in stainless steel AISI304 18/10, dim. 40x40x32 cm |  | 2 pcs
| 8.152 | Waste bin with pedal with wheels, made in stainless steel AISI304, 50litres capacity, with lid, pedal, mechanism for opening and two handles, dim. Ø39x60 cm |  | 2 pcs
| 8.153 | Planetary mixer Planetary mixer of 10 liters, with dough and cream additions included: cream wire, dough hook and dough spatula, three speeds. Power 0,45kW, CE |  | 1 pcs
| 8.154 | Transportation trolley framework in square 25x25 mm tubing, completely welded, runners, with upturn at each end, for 14 tray 600x400mm, pivoting Ø 125mm wheels, 2 with brakes, fitted with plastic bumpers, constructed entirely in 304 AISI stainless steel, dim. 50x70x180±5 cm |  | 1 pcs
| 8.155 | Transportation trolley framework in square 25x25 mm tubing, completely welded, runners, with upturn at each end, for 14x1/1GN containers, pivoting Ø 125mm wheels, 2 with brakes, fitted with plastic bumpers, constructed entirely in 304 AISI stainless steel, dim. 45x65x180 cm±5 cm |  | 2 pcs
| 8.156 | El.steam-convection oven Capacity: 10 x GN 1/1, distance 67mm, quick and easy control Touch Screen, Manual baking modes: Convection baking 30-260 ° C baking technology 35-260 ° C with 30-90% humidity Combined convection baking mode with humidity up to 20% Steam 48-130 ° C with moisture baking dry with dry air up to 260 ° C with a humidity of 10-100% MULTI.Point probe Advanced and automatic cooking programs: possibility to set up to 10 different timers container loading technology so that all containers are ready at the same time |  | 1 pcs
Preheating oven preheating technology Smart humidity sensors Cooking intelligent sensors recognize the type and quantity of food in the oven and automatically adjust the cooking parameters to obtain perfect cooking results smart sensors gently regulate heat distribution when it is necessary to rapidly increase the temperature in the oven, in order to gently bake sensitive products possibility to estimate the oven dishness and suggests a suitable washing program reversible fans, 4 speeds, programmable Moisture and vapor production of 35 ° C in steps of 10% Moisture removal from the chamber in steps of 10% technology of power consumption rationalization Automatic cleaning: 4 automatic washing cycle Total power: 18,5kW dim.750x775x1000mm
| 8.157 | Open stand with rach supports made in AISI 304 stainless steel, open, with slides for trays, complete with PVC sliding legs, dim.730x550x750±50mm |  | 1 pcs
| 8.158 | Automatic water softener , 8 l, CE |  | 1 pcs
| 8.159 | Three level oven Capacity of each oven 3 x GN 2/1, oven size 540x730x300 mm, all in stainless steel AISI304, power 17,7kw, dim. 80x90x170±5 cm, CE |  | 1 pcs
| 8.160 | Worktable with single-bowl sink with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, worktop and sink stainless steel thickness 1mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 180x70x85 cm |  | 1 pcs
| 8.161 | Hand operated mixer tap |  | 1 pcs
| 8.162 | Plastic drain siphon |  | 1 pcs
| 8.163 | Worktable central, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 170x70x85 cm |  | 2 pcs
| 8.164 | Worktable central, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 160x80x85 cm |  | 1 pcs
| 8.165 | El. cookers with 4 plates and oven made of INOX, stainless steel 1.5mm thick, with 4 square plates 300x300mm , power regulators with 7 levels for each ring separately, indicator light when the ring is turned on, fuse for interruption of power supply in case of overheating, oven capacity: 4x GN 2/1, oven dimensions: 540x730x300mm. Total power: 18,9kW 400V, 3N, dim. 80x90x90 cm, CE |  | 2 pcs
| 8.166 | Steam boiling pan 100 l tank dimensions 540x640x360 mm, steam consumption 30kg / h, all in stainless steel AISI304, indirect heating, without condensate separator, dim. 80x90x90 cm |  | 4 pcs
| 8.167 | Neutral worktop closed on three sides, open on work side, with drawer and bottom shelf, all in AISI 304 stainless steel, dim. 80x90x90 cm |  | 1 pcs
| 8.168 | El. bratt pan 80 l. rectangular stainless steel bowl, all in AISI304 stainless steel, manual lifting of the vessel using a rotary lever, power 9,0 kw, dim. 80x90x90 cm |  | 1 pcs
ISSUE FOOD
| 8.169 | Worktop cupboard with sliding doors central, made in stainless steel AISI304 18/10, worktop and central shelf thickness 40mm, rounded edges, stainless steel thickness 0.8mm, supporting structure square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, sliding doors made in double layer stainless steel, 0.6mm in thickness on the inside and 0.8mm on the oustide, with opening handles formed along entire door height, complete with adjustable PVC legs dim. 160x70x85 cm |  | 3 pcs
STORAGE OF BREAD
| 8.170 | Bread slicer It cuts the bread in all sizes into the slices of equal thickness. Work sites and blades are manufactured from durable stainless sheet. Complies with health conditions. Automatic shut-off at the end of the slicing process. Wheels. It has 5 mm thick reinforced body and stainless steel top cover. It has capped bearings which don't require lubrication and maintenance. It operates silently and without vibration. Machine body is painted with durable electrostatic paint. Three-phase or single-phase (230V- 400V, 1N-3N 50 Hz) Motor 0.37kW. Bread capacity 2-8kg, slice thickness 16mm, number of knives 30pcs, capacity 500 breads/hour. |  | 1 pcs
| 8.171 | Worktable with upstand, made in acid-resistant stainless steel AISI304 18/10, worktop thickness 40mm, with rounded edges, stainless steel thickness 0.8mm, supporting structure stainless steel square profiles 40x40x1.2mm, upstand thickness 15mm, and 100mm in height, complete with adjustable PVC legs, dim. 160x70x85 cm |  | 1 pcs
WAREHOUSE
| 8.172 | Digital refrigerator 1400 l Refrigerated cabinets equipped with a single-piece ventilated cooling unit, featuring a tropical compressor easy to remove from its casing. The external working environment can reach a temperature of +42°C with 65% relative humidity. The structure is made of stainless steel, except for the external back panel, top panel and bottom panel, which are in galvanized metal. Reversible RHS/LHS opening doors , with automatic closure and opening position fixed at 100°, easily removable magnetic gasket, rounded internal corners to facilitate thorough cleaning. Cabinets supplied standard with 60 mm thick polyurethane insulation, injected with HFO blowing agent at high pressure, without CFCs or HCFCs. |  | 4 pcs
Capacitive digital touch controller, temperature detection using NTC probes. Ventilated cooling with R134a gas for normal temperatures. Evaporator with catephoric painting treatment. Automatic defrosting and evaporation of water condensate through AISI 316L stainless steel coil. Adjustable and removable guides and grill racks. Adjustable stainless steel feet. Power cord supplied standard without plug. Supplied standard with: no.2x3 plastic-coated grills GN2/1 and no.2x6 stainless steel “U” guides . Temperature regime 0/+10°C. Power 0,51kw
FOOD CABINET
| 8.173 | Peristaltic enteral feeding pump |  | 5 pcs
Volumetric pump for intermittent application of drugs, blood, parenteral and enteral nutrition
Standard Flow Rate: 0,1 ml/h - 1.200 ml /h
Max Bolus Flow Rate: 999 ml/h
Flow rate precision : +/- 5%
Bolus application: - bolus application by continuous pressure on bolus command button - bolus application by defining volume to bi infused (0,1 - 99,9 ml) - bolus application by calculation of flow rate
Programing of infusion time (from 00h:01min. - 99h59min. in steps of 1min.)
Automatic calculation of flow rate (ml/h) derived from medication dose (µg, mg, ../kg/min )
Automatic calculation of flow rate (ml/h) delivered from: - concentration or quantity of medication (µg, mg, IU, mmol, mEq / ml), - quantity of medication in specific volume, - patient weight
Pump should be able to perform infusion: - parenteral infusion of drugs, - parentaral nutrition, - enteral nutrition* , - blood
* if offered pump can not be used for aplication of enteral nutrition potential bider can offer additional 5 unites of dedicated enetral pumps free of charge)
Automatic calculation of flow rate increasement / decreasement according to parameters: total volume to be infused, time and time of increasement / decreasement
Interval / Intermitent infusion - programming of min. 12 infusion sequences according to parameters (time, volume, flow rate)
Drug List with capacity of minimum of 1000 drugs and option of drug grouping.
Safety systems
Visual and sound alarms for: - deviation of infused volume , - pressure increase in infusion system (in perfusor syringe or perfusor tube), - occlusion alarm, - infusion stop / end, - battery low alarm, - incorrect IV system placement, - therapy close to end (can be set based on end user request)
KVO (Keep Vain Open) - option of keeping vain access open after infusion is ended
"Air in line" detection
Two (2) pressure check points (upstrem and downstream)
Automatic priming of IV lines with no need to remove IV line from the pump
User inteface and software in Serbian language
Option of bare-cod use (upgrade)
Integrated 1000ml nutrition bag
Needle free injection port
Female inverse luer-lock patient conector
Male scaled cone adapter
| 8.174 | Peristaltic enteral feeding pump |  | 5 pcs
rotary peristaltic enteral feeding pump designed to deliver programmed doses of enteral nutrition solutions (botlle wrapping)
SIZE Pump Dimensions : (35 mm H x 140 mm W x 95 mm D)
Pump Weight : 392 g, Battery: internal rechargeable Lithium ion battery, 3.6 V DC 2000 mAh
Flow rate: 1-400 ml/h
Accuracy flow rate: ±5,0% +bottle sets enable the connection between an ENFit feeding tube and feeding reservoirs with a wide neck (40mm) screw fitting or crown cork
Max 24 hours usage
Single patient usage
Components:
Universal connector for wide neck or crown cork bottles
Drip chamber
Pump insert
Medication port stopcock ENFit
Female ENFit connector with dustcap
Bottle hanger Language: English
Medical grade PVC tubing
Latex-free
DEHP-free +Single use enteral syringe
Intended for one patient only
PEGs and medical ports, 60 mll
Body composition analyzer:
Measurement Method: Direct Segmental Multi-frequency Bioelectrical Impedance Analysis Method (DSM-BIA) Simultaneous Multi-frequency Impedance Measurement (SMF-BIA)
Measurement technology: 30 Impedance Measurements by Using 6 Different Frequencies (1kH, 5kHz, 50kHz, 250kHz, 500kHz, 1000kHz) at Each of 5 Segments (Right Arm, Left Arm, Trunk, Right Leg, and Left Leg)
Electrode Method: Tetrapolar 8-Point Tactile Electrodes with Thumb Electrodes
Hand electrodes : They consist of two electrodes, electrode for thumb is separted
Calculation Method: No Empirical Estimation, choice Male/Female does not influence measured results.
Voice Guidance: Provides audible indication for test in progress, test complete, and successfully saved settings changes
Database: The device can save up to 100,000 results.
Applied Rating Current : 80µA A(±10µA)
Display Type: 800 × 480 Color TFT LCD
| 8.175 | Body composition analyzer |  | 1 pcs
Internal Interface: Touchscreen, Keypad
External Interface: RS-232C 4EA, USB HOST 2EA, USB SLAVE 1EA, LAN (10T) 1EA, Bluetooth 1EA, Wi-Fi 1EA
Testing Time: About 60 seconds
Testing Age Range: 3~99 years
Height Range: 95 ~ 220cm (3ft. 1.40in. ~ 7ft. 2.61in.)
Testing Weight Range: 10 ~ 270kg (22.0 ~ 595lbs)
Outots reports: Results Sheet, Result Sheet for Children and Body Water Results Sheet
Outputs (Results Sheet) parameters:
Total Body Water,
Protein,
Soft Lean Mass,
Minerals,
Fat Free Mass,
Weight,
Skeletal Muscle Mass,
Body Fat Mass,
Body Mass Index,
Percent Body Fat,
Based on ideal weight/Based on current weight: Right Arm, Left Arm, Trunk, Right Leg, Left Leg
ECW Ratio Analysis
Percent Body Fat,
ECW Ratio,
Visceral Fat Area (Graph),
Based on BMI/Percent Body Fat: Athletic Shape, Slightly Obese, Obesity, Muscular Shape, Average, Slightly Obese, Slim Muscular, Slim Sarcopenic Obesity, Thin, Slightly Thin
Intracellular Water,
Extracellular Water,
Skeletal Muscle Mass,
Basal and Metabolic rate,
Waist-Hip Ratio,
Visceral Fat Level,
Obesity Degree,
Bone Mineral Content,
Body Cell Mass,
Arm Circumference,
Arm Muscle,
Reactance (5kHz, 50kHz, 250kHz)
Whole Body Phase Angle (50kHz),
Segmental Phase Angle (50kHz, Right Arm, Left Arm, Trunk, Right Leg, Left Leg,
LAUNDRY - WASHING AND DRYING
| 8.176 | Washer extractor High extraction force for eficient dewatering • fully programmable, microprocessor with up to 150 programs • Large door opening for easy loading/unloading • Low noise level, for pleasant working environment • One frequency controlled motor • Prepared for 9 liquid detergent signals • Eight liquid supply connections • Stainless and galvanized steel in all vital parts for highdegree of rust protection • Stainless steel front panels and powder painted sidepanels • Electrical or direct steam heating Max. capacity filling factor 1:9 : 45 kg . Drum volume : 400 lit. Drum diameter: Ø 920 mm Extraction : 825 rpm G factor : 350 Total time : 50 min. Total power : 38 kw CE |  | 1 pcs
| 8.177 | Washer extractor Microprocessor: – Large display and one control knob for easy program selection – Language selection – Quick selection and option buttons for the most frequently used wash programs and options – Wash program packages optimized on Economy, Perfomance and Time. (Available on general programs) • AS – Automatic Savings: – The machine weighs the linen and adjusts the amount of water and energy according to the load. This leads to considerable savings in case the machine is not fully loaded.(Available on general programs) • Extremely low water and energy consumption – see table below • Four compartment detergent box for manual dosing of powder or liquid detergent • High extraction force for efficient dewatering • Door with gasket and strong hinges built for heavy duty operation • Stainless and galvanized steel in all vital parts for high degree of rust protection • Power Balance: – Maximises the extraction force, leading to highest dewatering and savings in drying – Minimizes the stress on the machine, leading to trouble-free operation and increased machine lifetime – Reduces vibrations, making it possible to install the machine in all locations |  | 2 pcs
Max. capacity filling factor 1:9 : 27 kg . Drum volume: 240 lit. Drum diameter: Ø 795 mm Extraction : 1005 rpm G factor : 450 Total time : 50 min. Total power : 19,9 kw CE
| 8.178 | Tumble dryer Axial airflow and tight construction gives low energy consumption • 2 full loads per hour • Large door opening for easy loading and unloading • Easy and ergonomic access to the lint screen • Excellent water evaporation efficiency per kWh • The coin version with Ecopower to avoid over drying of the garments and get a lower energy consumption • With program control - Large and clear display with control knob for easy program selection - Easy access with user-friendly interface - Language selection - Drying program packages optimized for Economy, Care and Time - Service program for adjustment of parameters - USB connection Rated capacity, filling factor (1:22) : 30,6 kg . filling factor (1:18) : 37 kg . Drum volume : 675 lit. Drum diameter : Ø 913 mm Total time : 28 min. Total power : 34,5 kw CE |  | 1 pcs
| 8.179 | Tumble dryer Rated capacity, filling factor (1:22) : 15,5 kg . filling factor (1:18) : 19,4 kg linen Drum volume : 360 lit. Drum diameter : Ø 755 mm Total time : 40 min. Total pover : 14,5 kw CE |  | 2 pcs