## Price Schedule
Bidder: |  | Offered price:
LOT 2 - Equipment for Internal Medicine Department
LOT ID | Equipment name | QTY | Model, Manufacturer and Country of Origin | Unit price | Total Price per line item
1 | MICROSCOPE WITH ADAPTERS | 6 |  |  | 0
2 | MICROSCOPE WITH CAMERA AND ADAPTERS | 2 |  |  | 0
3 | BLOOD CELL SEPARATOR WITHOUT PHOTOPHERESIS | 1 |  |  | 0
4 | BLOOD CELL SEPARATOR WITH PHOTOPHERESIS | 1 |  |  | 0
5 | Kidney biopsy gun, autocut | 4 |  |  | 0
6 | VVG/needle for 14G automatic gun | 60 |  |  | 0
7 | VVG/needle for 16G automatic gun | 20 |  |  | 0
8 | VVT/introducer | 80 |  |  | 0
9 | DXA apparatus for measuring bone density and body composition | 2 |  |  | 0
10 | Small Intestinal Capsule system, Endocapsule | 1 |  |  | 0
11 | Proctoscope set | 2 |  |  | 0
12 | Video colonoscope | 1 |  |  | 0
13 | Endo-thermal disinfector | 2 |  |  | 0
14 | Video gastroscope | 3 |  |  | 0
15 | Video colonoscope- dual focus | 3 |  |  | 0
16 | Energy platform for GI procedures | 1 |  |  | 0
17 | Endoscopic tower for EUS procedures | 1 |  |  | 0
18 | Endoscopic tower | 2 |  |  | 0
19 | Endoscopic video tower with colonoscope and gastroscope | 1 |  |  | 0
20 | Microscope automated imaging and analysis systems | 1 |  |  | 0
21 | Automated platform for FISH | 1 |  |  | 0
22 | Voltage stabilizer 3000VA / 2700W on line | 12 |  |  | 0
23 | Voltage stabilizer 3000VA/3000W on line | 4 |  |  | 0
24 | Laboratory refrigerator | 5 |  |  | 0
25 | Real-time PCR with notebook computer and software | 1 |  |  | 0
26 | PCR Thermocycler with block capacity: 96 tubes of 0.2ml, microplate with 96 seats | 1 |  |  | 0
27 | Microcentrifuge with 96-tube plate rotor | 1 |  |  | 0
28 | Vortex mixer | 1 |  |  | 0
29 | Mixing device, for 1.5 ml tubes with a minimum of 24 seats | 1 |  |  | 0
30 | Pipette set (0.1-2.5 µl; 0.5-10 µl; 10-100 µl; 100-1000 µl and 500-5000 µl) + extension sets for the specified pipette set (Dual filter 0.1-10 µl, Dual filter 0,5-20 µl, Dual filter 2-100 µl, Dual filter 50-1000 µl) | 2 |  |  | 0
31 | Spectrophotometer for determining the concentration of nucleic acids | 1 |  |  | 0
32 | Precise analytical balance | 1 |  |  | 0
33 | Rotating pipette stand | 2 |  |  | 0
34 | Voltage stabilizer 3000VA/2700W | 1 |  |  | 0
35 | Hibridization system SDH-12 | 1 |  |  | 0
36 | Freezer at -60 C to -86 C | 1 |  |  | 0
37 | Table centrifuge with 4x6 rotor | 1 |  |  | 0
38 | Digestor | 1 |  |  | 0
39 | UV-VIS Spectrophotometer, range 195-1050 nm | 1 |  |  | 0
40 | Microcentrifuge refrigerated with rotor 24x1.5 ml | 1 |  |  | 0
|  |  |  | Total Price per Lot | 0

## LOT2
Bidder:
Line item No. |  | Description | Technical Specification Offered Yes/No | QTY | Unit price | Total Price per line item
1 |  | MICROSCOPE WITH ADAPTERS |  | 6
| 1.1 | Light microscope for transmitted light with halogen or LED illumination
| 1.2 | Ergonomic binocular tube with adjustable interpupillary distance min. (55-75)mm
| 1.3 | High resistance ceramic coated table, possibility for left- and right-handed operation
| 1.4 | Ergonomic height adjustable focus knobs
| 1.5 | Standard condenser with Koehler illumination
| 1.6 | Eyepieces min 10X/20
| 1.7 | Plan Objectives min. 4x,10x,20x,40x, 100x oil
2 |  | MICROSCOPE WITH CAMERA AND ADAPTERS |  | 2
| 2.1 | Light microscope for transmitted light with halogen or LED illumination
| 2.2 | Ergonomic binocular tube with adjustable interpupillary distance min. (55-75)mm
| 2.3 | High resistance ceramic coated table, possibility for left- and right-handed operation
| 2.4 | Ergonomic height adjustable focus knobs
| 2.5 | Standard condenser with Koehler illumination
| 2.6 | Eyepieces min 10X/20
| 2.7 | Plan Objectives min. 4x,10x,20x,40x, 100x oil
| 2.8 | Microscope WiFi digital color camera
| 2.9 | Live image full HD resolution min 1920x1080 pixel, 30 frame/sec
| 2.10 | Photo resolution min 5 Mpixel
| 2.11 | Possibilty to take photos and videos directly to SD card
| 2.12 | Possibility of remote control, independently of PC
3 |  | BLOOD CELL SEPARATOR WITHOUT PHOTOPHERESIS |  | 1
| 3.1 | Mobile automated blood cell separator for collecting of blood components
| 3.2 | Possibility of collecting 1-2-3 platelets dose
| 3.3 | Possibility of platelet collection with one or two needles
| 3.4 | Lowered Leukocytes count in final product witout filter usage
| 3.5 | Possibility of automated adding of PAS solution in final product (platelets)
| 3.6 | Regulation of maximal Cytrate infusion (from 0,5 to 1,5 mg/kg/min)
| 3.7 | Therapeutic programs: Therapeutic Plasma exchange; MNC collection; Exchange of red blood cells;
| 3.8 | Medium loss of platelets during collection of MNC cells: 10%
| 3.9 | Extracorporeal set volumen of 163ml during collection of MNC
| 3.10 | Plasma exchange with efficacy over 80%
| 3.11 | Extracorporeal set volumen of 160ml for plasma exchange procedure
| 3.12 | Medium loss of platelets during therapeutic plasma exchange: 2,3%
| 3.13 | Continuous cells collection, with simultaneous separation in collection bag out of separator chamber in order to prevent of cells damage
| 3.14 | System consists of 6 pumps, 5 scales for multiple security control of returned fluid
| 3.15 | System consists of 3 cassettes for prompt and easy set charging
| 3.16 | Accurate fluid balance control with real-time volume measurement and option for maintenance of device passability with saline solution (NaCl) gravity aplication
| 3.17 | Antimicrobial filter (0.2 micrones) in the place of injection of Methoxsalen
| 3.18 | Possibility of changing position of equipment after finishing `prime and go` option
| 3.19 | LCD color dislay with touch screen
| 3.20 | CE sign
| 3.21 | Automated cuff for hand
| 3.22 | Possibility of collecting of red blood cells with efficiency over >40g Hb/unit
| 3.23 | Numerical control of blood collection and return from 20 - 150ml/minute
| 3.24 | Possibility of plasma collecting as an additional product with the number of erythrocytes <6x10 on 9, WBC <1x10 on 6 / and PLT< 50x10 on 9
4 |  | BLOOD CELL SEPARATOR WITH PHOTOPHERESIS |  | 1
| 4.1 | Mobile automated blood cell separator for collecting of blood components
| 4.2 | Possibility of collecting 1-2-3 platelets dose
| 4.3 | Possibility of platelet collection with one or two needles
| 4.4 | Lowered Leukocytes count in final product witout filter usage
| 4.5 | Possibility of automated adding of PAS solution in final product (platelets)
| 4.6 | Regulation of maximal Cytrate infusion - at least from (0,5 to 1,5) mg/kg/min
| 4.7 | Therapeutic programs: Therapeutic Plasma exchange; MNC collection; Exchange of red blood cells; extracorporeal photopheresis
| 4.8 | Medium loss of platelets during collection of MNC cells: max. 10%
| 4.9 | Extracorporeal set volumen of min. 160ml during collection of MNC
| 4.10 | Plasma exchange with efficacy over 80%
| 4.11 | Extracorporeal set volumen of min.160ml for plasma exchange procedure
| 4.12 | Medium loss of platelets during therapeutic plasma exchange: 2,3%
| 4.13 | Continuous cells collection, with simultaneous separation in collection bag out of separator chamber in order to prevent of cells damage
| 4.14 | Two-way communication possibility between separator and extracorp device
| 4.15 | Device for extracorporeal photoimunotherapy designed to deliver min. 1,5J/cm2 of UVA light
| 4.16 | Accurate fluid balance control with real-time volume measurement and option for maintenance of device passability with saline solution (NaCl) gravity aplication
| 4.17 | Antimicrobial filter (max. 0.2 micrones) in the place of injection of Methoxsalen
| 4.18 | Possibility of changing position of equipment after finishing `prime and go` option
| 4.19 | LCD color dislay with touch screen
| 4.20 | CE sign
5 |  | Kidney biopsy gun, autocut |  | 4
| 5.1 | Variable throw min. range (16-22)mm
| 5.2 | Safety button
| 5.3 | Reusable gun
| 5.4 | Loaded status indicator
| 5.5 | Compatible with biopsy histological needle for model V tek
| 5.6 | Usable for histological biopsies including kidney, liver, breast, etc.
6 |  | VVG/needle for 14G automatic gun |  | 60
| 6.1 | Compatible with introducer/coaxial (VVT) and biopsy gun
| 6.2 | Available gauge 14G
| 6.3 | Ultrasound enhanced visibility of a needle tip
| 6.4 | Color coded needle depending on gauge
7 |  | VVG/needle for 16G automatic gun |  | 20
| 7.1 | Compatible with introducer/coaxial (VVT) and biopsy gun
| 7.2 | Available gauge 16G
| 7.3 | Ultrasound enhanced visibility of a needle tip
| 7.4 | Color coded needle depending on gauge
8 |  | VVT/introducer |  | 80
| 8.1 | Introducer/coaxial, one gauge smaller than needle and 5cm shorter than histological needle), Compatible with histological needle
9 |  | DXA apparatus for measuring bone density and body composition |  | 2
| 9.1 | Dual Energy X-ray Absorptiometry (DXA) bone densitometer with body composition analysis
| 9.2 | Scanning along X and Y axis - narrow angle fan beam technology
| 9.3 | X-ray beam created with the K-edge filter (dual energy beam), not with the energy switching technique
| 9.4 | X-ray generator max. 80 kV
| 9.5 | Active scan area min. 195 cm x 60 cm
| 9.6 | Technology which reduces scan time and X-ray dosage byidentifying bone regions after each transverse sweep and estimating where to begin scanning on the subsequent sweep
| 9.7 | Maximum patient weight supported min. 155 kg
| 9.8 | SQL server database
| 9.9 | Multiuser database with access from min. 3 computers connected to scanner via network
| 9.10 | Function which provides guidelines to minimize operator error through identification of potential measurement and/or analysis errors
| 9.11 | Customized reports and templates
| 9.12 | Performing AP spine and dual femur exam without repositioning between scans
| 9.13 | Feature which allows to set up multiple measurements in one exam
| 9.14 | Types of scanning:
| 9.14.1 | AP spine
| 9.14.2 | Lateral BMD
| 9.14.3 | Femur
| 9.14.4 | Dual femur
| 9.14.5 | Seated forearm
| 9.14.6 | Nonseated forearm
| 9.14.7 | View of the entire spine (min. T7-L4) in AP and/or lateral position by dual-energy technique (with the possibility of displaying either single-energy and double-energy images) for visualization of vertebral fractures and morphometry
| 9.14.8 | Total body BMD
| 9.14.9 | Fracture risk tool (FRAX)
| 9.14.10 | Pediatric AP spine
| 9.14.11 | Pediatric femur
| 9.14.12 | Pediatric total body
| 9.14.13 | Total body composition
| 9.14.14 | Estimating visceral and subcutaneous adipose tissue mass, volume and area within the android region
| 9.15 | Trabecular Bone Score – integrated TBS which provides TBS score based on assessment of trabecular region of bone, including FRAX-adjusted TBS
| 9.16 | Estimated skin entrance dose (μGy) for standard (most commonly used) scan mode:
| 9.16.1 | AP spine max. 45 μGy
| 9.16.2 | Femur max. 45 μGy
| 9.16.3 | Forearm max. 2 μGy
| 9.16.4 | Total Body max. 0.5 μGy
| 9.17 | Desktop computer with monitor min. 24" and printer
| 9.18 | DICOM
10 |  | Small Intestinal Capsule system, Endocapsule |  | 1
| 10.1 | Endocapsule
| 10.2 | Optics Field of view min.160 degrees
| 10.3 | Depth of field min. range (0–20) mm
| 10.4 | Recorder unit
| 10.5 | Battery pack
| 10.6 | Real-Time View/Capture, allows to play back and capture images any time during the procedure
| 10.7 | Patient Guidance Function- as text messages preceded by beep and vibration alerts
| 10.8 | Battery Life min. 12 hours
| 10.9 | Antenna Unit- belt type
| 10.10 | Antenna unit – lead type
| 10.11 | PC with installed Endocapsule reading software with a variety of functions to detect images requiring closer inspection, providing the means for the fast reviewing of results to ultimately speed up diagnosis
| 10.12 | Red Color Detection - identify individual images that are suspected of containing an active bleed, angioectasia, red spots, ulcers or erosions.
| 10.13 | 3D Track Function - allowing the system to display the capsule track in 3D
| 10.14 | Adjust Mode - Change playback speed depending on differences in images Omni -selected Mode - recognize that an image is identical, even when the capsule is displaying the same section of small intestine from a different angle, images that overlap with previous ones are skipped Red assist mode - Playback speed slows down at the first image detected by the Red Color Detection and is automatically adjusted according to the number of successive red lesion images displayed
11 |  | Proctoscope set |  | 2
| 11.1 | Led light source
| 11.2 | Lamp service life min.30.000h
| 11.3 | Color temperature: 6500K
| 11.4 | Brightness control: 0-100%
| 11.5 | Proctoscope tube Ø 24mm, WL 60mm
| 11.6 | Obturator Ø 24mm, WL 60mm
| 11.7 | Proctoscope tube Ø 21mm, WL 90mm
| 11.8 | Handle with cold light illumination and removable attachment
| 11.9 | Fiber light cable Ø 2.5mm, WL approx. 2300mm
| 11.10 | Sponge forceps, WL approx.300mm
| 11.11 | Disinfection box including lid
12 |  | Video colonoscope |  | 1
| 12.1 | Field of view min. 140°
| 12.2 | Depth of field min. range (3-100) mm
| 12.3 | Insertion tube outer diameter 12.8 mm
| 12.4 | Distal end outer diametar 12.8 mm
| 12.5 | Instrument channel inner diameter 3.7 mm
| 12.6 | Working length min. 1330 mm
| 12.7 | Angulation range: Up/Down 180°/180° Left/Right 160°/160°
| 12.8 | Compatible with Olympus CV-165 processor
13 |  | Endo-thermal disinfector |  | 2
| 13.1 | Possibility of simultaneous disinfection of two endoscopes
| 13.2 | PAA process
| 13.3 | Flexibile endoscope disinfection process with drying with air temperature 57°C
| 13.4 | Self disinfection program
| 13.5 | EN ISO 15883 -1 i -4
| 13.6 | Mechanical and color coded containers for chemicals (cleaner, desifier and activator)
| 13.7 | Automatic leakage test
| 13.8 | Electronic security door lock
| 13.9 | LC display 4 x 40 characters
| 13.10 | Outer casing Stainless steel
| 13.11 | Aditional basket and adapters for one disinfector for EUS endoscope
| 13.12 | One set of chemistry per disinfector
14 |  | Video gastroscope |  | 3
| 14.1 | Field of view min. 140°
| 14.2 | HDTV image quality
| 14.3 | Depth of field min. range (2-100) mm
| 14.4 | Insertion tube outer diameter 8,9 mm or less
| 14.5 | Distal end outer diameter 8,9 mm or less
| 14.6 | Instrument channel inner diameter 2,8 mm or more
| 14.7 | Water jet channel
| 14.8 | Working length min. 1030 mm
| 14.9 | Angulation range: Up/Down 210°/90° Left/Right 100°/100°
| 14.10 | Enhancement of the visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light
| 14.11 | White light Texture and Color enhacement for enhanced visibility of potentially suspicious tissue which includes inflammations, flat or depressed lesions
| 14.12 | Observation method for easier detection of bleeding spots and visualization of deep blood vessels by using specific light wavelenghts od red light
| 14.13 | Single use cleaning brush
| 14.14 | Compatible with existing EXERA III CV/CLV 190 system
15 |  | Video colonoscope- dual focus |  | 3
| 15.1 | Dual focus colonoscope
| 15.2 | Field of view in normal focus mode170°; in near focus mode 160°
| 15.3 | Depth of field in normal focus mode 5-100 mm; in near focus mode 2-6 mm
| 15.4 | HDTV image quality
| 15.5 | Insertion tube outer diameter 12,8 mm
| 15.6 | Distal end outer diametar 13,2 mm
| 15.7 | Instrument channel inner diameter 3,7 mm
| 15.8 | Water jet chanel
| 15.9 | Working length min. 1680 mm
| 15.10 | Angulation range: Up/Down 180°/180° Left/Right 160°/160°
| 15.11 | Variable stiffnes
| 15.12 | Enhancement of the visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light
| 15.13 | White light Texture and Color enhacement for enhanced visibility of potentially suspicious tissue which includes inflammations, flat or depressed lesions
| 15.14 | Observation method for easier detection of bleeding spots and visualization of deep blood vessels by using specific light wavelenghts
| 15.15 | Waterproof One-touch Connector (no need for aditional water resistant caps)
| 15.16 | Single use cleaning brush
| 15.17 | Compatible with existing EXERA III CV/CLV 190 system
16 |  | Energy platform for GI procedures |  | 1
| 16.1 | Electrosurgical Generator
Power min. 120W
Intuitive touch screen
Monopolar: Pure Cut, Blend Cut,Pulse Cut slow,Pulse Cut fast
Bipolar Cut
Minimum 4 different monopolar coagulation modes
Minimum 4 different bipolar coagulation modes
| 16.2 | Argon Gas Unit
Automatic Argon probe recognition
Argon probe socket
MONOPOLAR socket min: 3-pin plug with 4 mm pin diameter 1-pin plug with 8 mm pin diameter 1-pin plug with 4 mm pin diameter
BIPOLAR socket min: 2-pin plug with 4 mm pin diameter, fixed pin spacing 28.6 mm 2-pin plug with 4 mm pin diameter and a 3rd coding pin Coaxial plug with 8/4 mm pin diameter
| 16.3 | Energy Cart
| 16.4 | Single use HF snares
| 16.5 | Single Use Plasma Coagulation Probe Axial
17 |  | Endoscopic tower for EUS procedures |  | 1
| 17.1 | Xenon light source - 300 W xenon lamp, expected lamp life mion. 500h
| 17.2 | Video processor
Analog HDTV signal output: (RGB, YPbPr),
Digital signal output: HD-SDI, SD-SDI, DV, i DVI
Image enhancement setting up to 3 levels, 8 possible modes
Enhancement of the visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light
Portable Memory to store endoscopic images
| 17.3 | Ultrasound gastrovideoscope Linear
Field of view 100°/ forward oblique view 55°
Depth of field 3-100 mm
Insertion tube outer diameter 12,6 mm
Channel inner diameter 3,7 mm
Angulation range Up/Down 130°/90° left/right 90°/90°
Scanning Direction: Parallel to the insertion direction
Frequency: 5/6/7.5/10/12 MHz
Scaning range 180°
Contacting Method: Balloon method, Direct Contact method
Operation mode: B-mode, FLOW mode, PW mode or equivalent
Connection with US processor via datachable cable
| 17.4 | Ultrasound gastrovideoscope Radial
Field of view 100°; Direction of view-forward oblique viewing 50°
Depth of field 3-100 mm
Insertion tube outer diameter 10,9 mm
Distal end outer diameter ø 13.4 mm
Channel inner diameter 3,7 mm
Angulation range Up/Down 130°/90° left/right 90°/90°
Scanning Direction: Perpendicular to the insertion direction
Frequency: 5/6/7.5/10/12 MHz
Scaning range 360°
Contacting Method: Balloon method,
Operation mode: B-mode, FLOW mode, PW mode or equivalent
Connection with US processor via datachable cable
| 17.5 | Enoscopic ultrasound center
Mechanical scanning: Display mode: B-mode Scanning: Radial scanning Usable frequencies: C5, C7.5, C12, C20, 7.5, 12, 20 MHz Display range: 2, 3, 4, 6, 9, 12 cm Measurement Distance, area, circumstance
Electronic scanning: Display mode: B-mode, FLOW mode, PW mode or equivalent Scanning: Radial scanning, curved linear array scanning. Usable frequencies: 5, 6, 7.5, 10, 12 MHz Display range: 2, 3, 4, 5, 6, 7, 8, 9, 12 cm
Measurement Distance, area
| 17.6 | LCD Monitor
Display size min diagonale 31,5 inches 4K upscaling function or equivalent Minimal resolution 3840x2160 pixels (UHD) Viewing angle 178° /178°
Input: 4K in: 12G-SDI ×2, Display Port ×1, HDMI ×1SDI (1); SDI (2) 2K in: 3G-SDI ×1, DVI-D x 1
Output: 4K, 2K, Clone out
| 17.7 | Trolley
Sliding Keyboard Tray LCD Monitor arm with swivel and tilt funcionality Scope Pole Kit Integrated Separation Transformer Main ON/OFF switch
| 17.8 | Suction pump
Nominal vacuum: 95 kPa pre-set flow speed settings 40l/min, 50l/min and 60l/min Thermal cut-out protection Suction jar minimum capacity 2l Disposable bacteria filter minimum 10 pcs
| 17.9 | Leakage tester
18 |  | Endoscopic tower |  | 2
| 18.1 | Video processor
Integrated 5 LEDs light source
Operation via touch-sensitive panel
Enhanced oservation methods for compatible endoscopes:
Enhanced visibility of capillaries and other structures on the mucosal surface by using the blue and green narrow band observation light
Analog signal output: VBS composite; Y/C
Digital signal output: 12G-SDI, 3G-SDI, HD-SDI, SD-SDI
Electronic zoom
Iris (auto, peak and average)
Registration of patient data (50 patients)
User settings: 20 different user settings
Portable Memory to store endoscopic images
Compatible with existing Olympus CF/GIF-H185 endoscope
| 18.2 | LCD Monitor
Display size 31,5 inches 4K upscaling function Resolution 3840x2160 pixels (UHD) Viewing angle 178° /178°
Input: 4K in: 12G-SDI ×2, Display Port ×1, HDMI ×1 2K in: 3G-SDI ×1, DVI-D x 1
Output: 4K, 2K, Clone out
| 18.3 | Trolley
Sliding Keyboard Tray LCD Monitor arm with swivel and tilt funcionality Scope Pole Kit Integrated Separation Transformer Main ON/OFF switch
| 18.4 | Suction pump
Nominal vacuum: 95 kPa
pre-set flow speed settings 40l/min, 50l/min and 60l/min
Thermal cut-out protection
Suction jar 2l
Disposable bacteria filter
| 18.5 | Flushing pump
Peristaltic pump for additional rinsing in endoscopic procedures
Automatic cut-off timer for safe control of irrigation
Flow rate 230-750 ml/min
Two liters water container
| 18.6 | Leakage tester
19 |  | Endoscopic video tower with colonoscope and gastroscope |  | 1
| 19.1 | Video processor
| 19.2 | Integrated 5 LEDs light source
Operation via touch-sensitive panel
Enhanced oservation methods for compatible endoscopes:
Enhanced visibility of capillaries and other structures on the mucosal surface by using the blue and green narrow band observation light
Analog signal output: VBS composite; Y/C
Digital signal output: 12G-SDI, 3G-SDI, HD-SDI, SD-SDI
Electronic zoom
Structural enhancement: Enhancement of contrast of the fine patterns in the image
Iris (auto, peak and average)
Registration of patient data (50 patients)
User settings: 20 different user settings
Portable Memory to store endoscopic images
| 19.3 | Video colonoscope
Field of view in normal focus mode170°; in near focus mode 160°
Depth of field 2-100 mm
HDTV image quality
Insertion tube outer diameter 12,8 mm
Distal end outer diametar 13,2 mm
Instrument channel inner diameter 3,7 mm
Water jet chanel
Working length 1680 mm
Angulation range: Up/Down 180°/180° Left/Right 160°/160°
Variable stiffnes
Enhancement of the visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light
Waterproof One-touch Connector (no need for aditional water resistant caps)
| 19.4 | Video gastroscope
Field of view min. 140°
HDTV image quality
Depth of field min. 2-100 mm
Insertion tube outer diameter 8,9 mm or less
Distal end outer diameter 8,9 mm or less
Instrument channel inner diameter 2,8 mm or more
Water jet channel
Working length min. 1030 mm
Angulation range: Up/Down 210°/90° Left/Right 100°/100°
Enhancement of the visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light
White light Texture and Color enhacement for enhanced visibility of potentially suspicious tissue which includes inflammations, flat or depressed lesions
Observation method for easier detection of bleeding spots and visualization of deep blood vessels by using specific light wavelenghts od red light
Waterproof One-touch Connector (no need for aditional water resistant caps)
| 19.5 | LCD Monitor, 2 pcs.
LCD Panel Screen size 26 inches Display devices a-Si TFT active matriResolution 1,920 × 1,080 pixels (full HD) Viewing angle 178° (horizontal and vertical)
Input: SDI (1); SDI (2) VIDEO BNC connector Y/C 4-pin mini-DIN HD15 D-sub 15-pin DVI (1) DVI-I connector DVI (2) DVI-I connector Aux in (HD/SD-SDI) BNC connector
Output: HD/SD SDI (1) SDI (2) BNC, VIDEO BNC connector, Clone Out (3G/HD-SDI)
| 19.6 | Trolley
Sliding Keyboard Tray LCD Monitor arm with swivel and tilt funcionality Scope Pole Kit Integrated Separation Transformer Main ON/OFF switch
| 19.7 | Suction pump
Nominal vacuum: 95 kPa
pre-set flow speed settings 40l/min, 50l/min and 60l/min
Thermal cut-out protection
Suction jar 2l
Disposable bacteria filter
| 19.8 | Flushing pump
Peristaltic pump for additional rinsing in endoscopic procedures
Automatic cut-off timer for safe control of irrigation
Flow rate 230-750 ml/min
Two liters water container
Tubing
| 19.9 | Leakage tester
20 |  | Microscope automated imaging and analysis systems |  | 1
| 20.1 | Fully automated imaging platform comprised of fully equipped motorized fluorescent microscope, motorized 8 slides stage, digital camera, computer and monitor
| 20.2 | Possibility to perform automated imaging as well as review and analysis on the same station
| 20.3 | Hallogen bright field light source
| 20.4 | Metal halide long life fluorescent light source, computer controled power off at end of scanning
| 20.5 | Motorized condenser with dry to lens NA 0.9
| 20.6 | Automated oil disenzer
| 20.7 | Response time 8ms (gray to gray)
| 20.8 | High resolution camera, min. 5 mega pixel resolution
| 20.9 | Microscope fluorescence illuminator equipped with a Fly’s-eye lens
| 20.10 | Full web review and analysis of FISH cases using standard web browsers from any location
| 20.11 | 7 position motorized nosepiece
21 |  | Automated platform for FISH |  | 1
| 21.1 | Automated and standardized slide specimen processing
| 21.2 | Possibility to perform deparaffinization, pretreatment, histology/cytology staining, special stains and routine slide washing with a single system
| 21.3 | Slide capacity per run: 150 mL basin up to 8 slides 250 mL basin up to 20 slides 500 mL basin up to 50 slides
| 21.4 | 12 ambient reagent basins
| 21.5 | 3 heated reagent basins
| 21.6 | Operating temperature 15-30⁰C
| 21.7 | Capability of batch processing, with 1-50 slides per run.
| 21.8 | Maximum throughput of the processor 50 slides in 15 min
| 21.9 | The system is capable of executing preprogrammed protocols
| 21.10 | The system is capable for user-programmed protocols including basin locations, dwell times, heated basin temoeratures, washing station mode, sequence of operations.
| 21.11 | The system will not allow robotic movement while the access door is open.
22 |  | Voltage stabilizer 3000VA / 2700W on line |  | 12
| 22.1 | Output power capacity: 3kVA/3kW
| 22.2 | Efficiency at 100%: > 93.3%
| 22.3 | Output Voltage Note: 208/220/230/240VAC +/-1% (configurable)
| 22.4 | THD: < 2% (linear load), < 5% (non-linear load)
| 22.5 | Output Frequency: 50 - 60Hz +/-0.2% (‘’battery mode’’)
| 22.6 | Output Voltage Distortion: +/- 1.5%
| 22.7 | Topology: Online Double Conversion
| 22.8 | Voltage: 230 Vac, 1PH
| 22.9 | Input voltage: 120-295 V
| 22.10 | Freaquency: 50 - 60 Hz +/-10%
| 22.11 | THDi: < 3%
| 22.12 | Internal batteries (6 x 12V, 9Ah), expandable
| 22.13 | Autonomy at 100% load and pf 1: 10 min.
| 22.14 | Communication: USB port, RS232
| 22.15 | Multifunction LCD panel, on the front of the device, with navigation keys for manual setting and monitoring of device parameters
| 22.16 | Audible and visual alarms
| 22.17 | Emergency Power Off
| 22.18 | SNMP, RS485 and dry contact optional
| 22.19 | ECO mode: high efficiency mode
| 22.20 | Bypass: static
| 22.21 | Protection: from overload, short circuit and overheating
| 22.22 | Dimensions (HxWxD): 330mm x 190mm x 400mm
| 22.23 | Working environment conditions: -5-40°C, rel.humidity 0-95%
| 22.24 | Standards: EN50091-1-1/IEC62040-1-1/AS 62040-1:2009, EN50091-2/IEC62040-2/AS 62040-2, EN50091-3/IEC62040-3/AS 62040-3(VFI SS 111), EN50171 (CEI 34-102)
23 |  | Voltage stabilizer 3000VA/3000W on line |  | 4
| 23.1 | Output power capacity: 3kVA/3kW
| 23.2 | Efficiency at 100%: > 93.3%
| 23.3 | Output Voltage Note: 208/220/230/240VAC +/-1% (configurable)
| 23.4 | THD: < 2% (linear load), < 5% (non-linear load)
| 23.5 | Output Frequency: 50 - 60Hz +/-0.2% (‘’battery mode’’)
| 23.6 | Output Voltage Distortion: +/- 1.5%
| 23.7 | Topology: Online Double Conversion
| 23.8 | Voltage: 230 Vac, 1PH
| 23.9 | Input voltage: 120-295 V
| 23.10 | Freaquency: 50 - 60 Hz +/-10%
| 23.11 | THDi: < 3%
| 23.12 | Internal batteries (6 x 12V, 9Ah), expandable
| 23.13 | Autonomy at 100% load and pf 1: 10 min.
| 23.14 | Communication: USB port, RS232
| 23.15 | Multifunction LCD panel, on the front of the device, with navigation keys for manual setting and monitoring of device parameters
| 23.16 | Audible and visual alarms
| 23.17 | Emergency Power Off
| 23.18 | SNMP, RS485 and dry contact optional
| 23.19 | ECO mode: high efficiency mode
| 23.20 | Bypass: static
| 23.21 | Protection: from overload, short circuit and overheating
| 23.22 | Dimensions (HxWxD): 330mm x 190mm x 400mm
| 23.23 | Working environment conditions: -5-40°C, rel.humidity 0-95%
| 23.24 | Standards: EN50091-1-1/IEC62040-1-1/AS 62040-1:2009, EN50091-2/IEC62040-2/AS 62040-2, EN50091-3/IEC62040-3/AS 62040-3(VFI SS 111), EN50171 (CEI 34-102)
24 |  | Laboratory refrigerator |  | 5
| 24.1 | Biomedical Refrigerator with glass door
| 24.2 | Temperature range (°C): min range (2 - 8)
| 24.3 | Max. ambient temperature (°C): min. 32
| 24.4 | Capacity: min 430 l
| 24.5 | Exterior dimensions WxDxH (mm): 600x700x1865
| 24.6 | Interior dimensions WxDxH (mm): 475x500/365x1723
| 24.7 | Weight (kg): max. 110
| 24.8 | Visual / Acoustic Alarm
| 24.9 | Ajustable high/low temperature alarm:yes / yes
25 |  | Real-time PCR with notebook computer and software |  | 1
| 25.1 | Fully automated amplification and possibility of following aplications: gene expression, SNP genotyping, microRNA, copy number variation, mutation detection
| 25.2 | Capability to run up to 96 reactions
| 25.3 | Amplification and detection: 5 excitation, 5 emission filters (detection of FAM, ROX, SYBR Green, JOE, TAMRA, Cy5, NED, VIC)
| 25.4 | Only one set of controls required for any batch size
| 25.5 | Standardized interface with LIS
| 25.6 | Data archiving (calibrator, control, and patient data logs)
| 25.7 | Automated quality checks and calibration
| 25.8 | Weight not greater than 35 kg
| 25.9 | Power source 100-240V
26 |  | PCR Thermocycler with block capacity: 96 tubes of 0.2ml, microplate with 96 seats |  | 1
| 26.1 | Compact design—fits on every benchtop
| 26.2 | Cloud-enabled—conveniently access your instrument anytime, anywhere
| 26.3 | Block format: 96-well, 0.2 mL isothermal block
| 26.4 | Max. block ramp rate: 3.0°C/sec
| 26.5 | Max. sample ramp rate: 2.2°C/sec
| 26.6 | Temperature accuracy :±0.25°C (35–99.9°C)
| 26.7 | Temperature range: 0–100.0°C
| 26.8 | Temperature uniformity: <0.5°C (30 sec after reaching 95°C)
| 26.9 | Temperature calibration: Calibrated to standards traceable to the National Institute of Standards and Technology
| 26.10 | PCR volume range: 10–100 μL
| 26.11 | Instrument memory : 2,000 MB onboard memory (capacity for >1,000 protocols); USB port for additional external storage
| 26.12 | Display interface: 5-inch color TFT LCD
| 26.13 | Power: 100–240 V, 50–60 Hz, max. 500 W
| 26.14 | Data connectivity: Cloud or mobile via Ethernet or WiFi
27 |  | Microcentrifuge with 96-tube plate rotor |  | 1
| 27.1 | Swing-out rotor cat no 11123 with 2 rectangular buckets for microtitar plates cat no. 13223/ max speed with this rotor - 4000 rpm / 2,111 x g
| 27.2 | Max. speed of centrifuge 15.300 rpm
| 27.3 | Automatic rotor identification
| 27.4 | Timer 10 s, 11h 59 min; continuous run, short run, increment 1s
| 27.5 | Speed increment 100 rpm/ 10 rpm
| 27.6 | 50 programs
| 27.7 | 10 acceleration and braking curves
| 27.8 | Angle 90°
| 27.9 | High-performance cooling
| 27.10 | Temperature setting range: -10 °C to +40 °C, increment 1°C
| 27.11 | Rotor cooling also possible at standstill
| 27.12 | Pre-cooling program "Rapid Temp"
| 27.13 | Refrigerant R134a
| 27.14 | Digital LCD display showing setpoints and actual values
| 27.15 | Backlit buttons for start, stop and lid opening
| 27.16 | Intuitive Spincontrol Universal controller
| 27.17 | Motorised lid lock
| 27.18 | Automatic lid opening with audible signal at the end of run (must be activated)
| 27.19 | Stainless steel chamber
| 27.20 | Induction motor
| 27.21 | Dimensions : (310 x 550 x 570) mm ±10%
| 27.22 | Power max. 750 W
| 27.23 | Weight without rotor max 60kg
| 27.24 | Produced in compliance with the latest national and international standards EN 61010-2-020
28 |  | Vortex mixer |  | 1
| 28.1 | Speed: from 0 to 3000 rpm
| 28.2 | Material: Zinc alloy and technopolymer
| 28.3 | Movement: Orbital, radius 4.5 mm
| 28.4 | Operational mode: Touch (pulse) or continuous
| 28.5 | Speed setting: Analog
| 28.6 | Speed control: Electronic
| 28.7 | Protection CEI EN 60529: IP 42
| 28.8 | Fuse/power: 100-240 V / 50-60 Hz / 15 W
| 28.9 | Dimension: max 180x70x220mm /2,2 Kg
29 |  | Mixing device, for 1.5 ml tubes with a minimum of 24 seats |  | 1
| 29.1 | ThermoMixer C, basic device without thermoblock, 220-240V
| 29.2 | Max. speed 3,000rpm
| 29.3 | Mixing and vortexing radius1.5mm
| 29.4 | Mixing frequency 300– 3,000 rpm
| 29.5 | Mixing orbit 3 mm
| 29.6 | Temperature accuracy Max. ±0.5 °C at 20 – 45 °C
| 29.7 | Temperature settings from 1 °C / 100 °C
| 29.8 | Sample capacity: All common tubes and plate formats from 5 µL - 50mL
| 29.9 | Timer5 sec to 99:30 h, continuous
| 29.10 | Heating rate Max. 7 °C/min (when using SmartBlock plates) Cooling rate Max. 2.5 °C/min between 100 °C and RT
| 29.11 | Interfaces USB interface
| 29.12 | Power supply 220 – 240 V, 50 – 60 Hz
| 29.13 | Max. power consumption 200 W
| 29.14 | Dimensions max 20,6 × 30,4 × 13,6 cm
| 29.15 | Weight w/o accessories 6,3 kg / 13,9 lb
| 29.16 | Temperature range Min: 15 °C below RT, max: 100 °C
| 29.17 | SmartBlock 1.5mL, Thermoblock for 24 tubes 1.5 mL
30 |  | Pipette set (0.1-2.5 µl; 0.5-10 µl; 10-100 µl; 100-1000 µl and 500-5000 µl) + extension sets for the specified pipette set (Dual filter 0.1-10 µl, Dual filter 0,5-20 µl, Dual filter 2-100 µl, Dual filter 50-1000 µl) |  | 2
| 30.1 | Pipette single-channel, variable 0.1 - 2.5 µL, dark gray, incl. epT.I.P.S.® box
| 30.2 | Pipette single-channel, variable 0.5–10 µL, medium gray incl. epT.I.P.S.® box
| 30.3 | Pipette single-channel, variable 10 - 100 µL, yellow, incl. epT.I.P.S.® box
| 30.4 | Pipette single-channel, variable 100 - 1000 µL, blue incl. epT.I.P.S.® box
| 30.5 | Pipette single-channel, variable 0.5 - 5 mL, purple, incl. epT.I.P.S.® sample bag
| 30.6 | Pipette Dualfilter TIPS 0,1-10µl M, sterile and PCR clean, 10 racks of 96 tips (960)
| 30.7 | Pipette Dualfilter TIPS 0,5-20 µL L, sterile and PCR clean, 10 racks of 96 tips (960)
| 30.8 | Pipette Dualfilter TIPS 2-100µl, sterile and PCR clean, 10 racks of 96 tips (960)
| 30.9 | Pipette Dualfilter TIPS 50-1000µl, sterile and PCR clean, 10 racks of 96 tips (960)
| 30.10 | Pipette Dualfilter TIPS 100-5000µl, sterile and PCR clean, 5 racks of 24 tips (120)
31 |  | Spectrophotometer for determining the concentration of nucleic acids |  | 1
| 31.1 | Absorption measuring range from 0 A – 3.0 A (260 nm)
| 31.2 | BSA concentration range (UV 280 nm) cuvette G1.0 758 ng/µL – 45,450 ng/µL
| 31.3 | BSA concentration range (UV 280 nm) cuvette 10 mm 76 ng/µL – 4,545 ng/µL
| 31.4 | BSA concentration range (UV 280 nm) cuvette 2 mm 379 ng/µL – 22,725 ng/µL
| 31.5 | BSA concentration range (for 1 to 10 mm path length, calculation factor for BSA = 1.515 µg/µL) 76 ng/µL – 45,450 ng/µL
| 31.6 | Cuvette shaft requred 12.5 mm × 12.5 mm
| 31.7 | Detector Type CMOS photodiode array, 1024 pixel
| 31.8 | Light source absorption Xenon flash lamp
| 31.9 | Measuring principle absorption Single-beam absorption spectrophotometer with reference beam
| 31.10 | Method memory>100 method programs
| 31.11 | Operator guidance language required Spanish, Italian, French, English, German, Japanese
| 31.12 | Power consumption Approx. 15 W during operating step
| 31.13 | Approx. 5 W with dimmed display
| 31.14 | Random error absorption≤0.002 at A = 0 ≤0.005 (0.5 %) at A = 1
| 31.15 | Random error wavelength≤0.5 nm
| 31.16 | Sample light path length 8.5 mm
| 31.17 | Scanning is required
| 31.18 | Smallest step size1 nm
| 31.19 | Spectral bandwidth ≤4 nm
| 31.20 | Systematic error absorption±1 % (A = 1)
| 31.21 | Systematic error wavelength ±1 nm
| 31.22 | Wavelength Selection Increment1 nm
| 31.23 | Wavelength range from 200 nm – 830 nm
| 31.24 | Wavelength range absorption Scan (nm): 200 – 830 at 1 nm increments dsDNA concentration range (UV 260 nm) cuvette G1.025 ng/µL – 1,500 ng/µL dsDNA concentration range (UV 260 nm) cuvette 10 mm2.5 ng/µL – 150 ng/µL dsDNA concentration range (UV 260 nm) cuvette 2 mm12.5 ng/µL – 750 ng/µL dsDNA concentration range (for 1 to 10 mm path length)2.5 ng/µL – 1,500 ng/µL
| 31.25 | Interfaces› USB master: for USB stick and thermal printer DPU-S445 › USB slave for connecting to a PC (All functions available without PC) › Serial interface RS-232: for thermal printer DPU-414 › Ethernet interface RJ45: For connecting to a network printer or emailing data outputs directly off instrument
| 31.26 | Power supply 230 V, 50 – 60 Hz
| 31.27 | Dimensions max 29,5 × 40,0 × 15,0 cm / 11,6 × 15,7 × 6 in
| 31.28 | Weight w/o accessories max 5,4 kg / 12 lb
| 31.29 | Kuvette 220 nm – 1,600 nm, individually packed, (including RNase-free, DNase-free) and protein-free, PCR clean, 50 – 2,000 µL, 200 pcs. individually packaged
| 31.30 | PCR-clean and protein-free quality
| 31.31 | UV transparent plastic, free of fluoropolymers and other halogenated hydrocarbons
| 31.32 | Two built-in optical path lengths in a single cuvette - just turn the UVette 90° to change from 10 mm and 2 mm
| 31.33 | Volume markings for 500 µL and 1,000 µL
| 31.34 | Individually packaged and certified PCR clean and protein-free quality for sensitive (e.g. RNA) and valuable samples
| 31.35 | Deep-lying optical window prevents scratches
| 31.36 | Tapered cuvette base for optimal filling, frosted grip for labeling
| 31.37 | Routine pack for convenient access to each UVette and safe storage in a reclosable box
32 |  | Precise analytical balance |  | 1
| 32.1 | Easy to use: All primary functions have their own key on the keypad
| 32.2 | Automatic internal adjustment time-controlled every 2 hours
| 32.3 | Guarantees high degree of accuracy and makes the balance independent of its location
| 32.4 | Capacity display: A bar lights up to show how much of the weighing range is still available
| 32.5 | Draught shield standard on all models with [Max] = 600 g weighing space W×D×H 134×128×80 mm
| 32.6 | RS-232 data interface, standard (not for EWJ-SM)
| 32.7 | USB data interfaces for transferring weighing data to the PC, printer, USB sticks, etc. (not for EWJ-SM)
| 32.8 | Protective working cover included with delivery
| 32.9 | Weighing capacity max 600 gr
| 32.10 | Weighing capacity max 600 gr
| 32.11 | Readability (d) 0.01 gr
| 32.12 | Verific. Value (e) 0.1 gr
| 32.13 | Minimum load (min) 0.5 gr
| 32.14 | Linearity 0.03 gr
33 |  | Rotating pipette stand |  | 2
| 33.1 | Carries up to six manual single or multi-channel pipettes
| 33.2 | Pipette holders are interchangeable, no tools required
| 33.3 | Robust design that securely holds and protects pipettes
| 33.4 | Handle for easy transport from bench to bench
| 33.5 | When used with optional pipette holder, the carousel can carry the Multipette
34 |  | Voltage stabilizer 3000VA/2700W |  | 1
| 34.1 | Output power capacity: 3kVA/3kW
| 34.2 | Efficiency at 100%: > 93.3%
| 34.3 | Output Voltage Note: 208/220/230/240VAC +/-1% (configurable)
| 34.4 | THD: < 2% (linear load), < 5% (non-linear load)
| 34.5 | Output Frequency: 50 - 60Hz +/-0.2% (‘’battery mode’’)
| 34.6 | Output Voltage Distortion: +/- 1.5%
| 34.7 | Topology: Online Double Conversion
| 34.8 | Voltage: 230 Vac, 1PH
| 34.9 | Input voltage: 120-295 V
| 34.10 | Freaquency: 50 - 60 Hz +/-10%
| 34.11 | THDi: < 3%
| 34.12 | Internal batteries (6 x 12V, 9Ah), expandable
| 34.13 | Autonomy at 100% load and pf 1: 10 min.
| 34.14 | Communication: USB port, RS232
| 34.15 | Multifunction LCD panel, on the front of the device, with navigation keys for manual setting and monitoring of device parameters
| 34.16 | Audible and visual alarms
| 34.17 | Emergency Power Off
| 34.18 | SNMP, RS485 and dry contact optional
| 34.19 | ECO mode: high efficiency mode
| 34.20 | Bypass: static
| 34.21 | Protection: from overload, short circuit and overheating
| 34.22 | Dimensions (HxWxD): 330mm x 190mm x 400mm
| 34.23 | Working environment conditions: -5-40°C, rel.humidity 0-95%
| 34.24 | Standards: EN50091-1-1/IEC62040-1-1/AS 62040-1:2009, EN50091-2/IEC62040-2/AS 62040-2, EN50091-3/IEC62040-3/AS 62040-3(VFI SS 111), EN50171 (CEI 34-102)
35 |  | Hibridization system SDH-12 |  | 1
| 35.1 | Capacity: 12 Slides
| 35.2 | Temperature range: up to 100 deg C
| 35.3 | Operation mode: Denaturation/Hybridization Hybridization, fixed temperature
| 35.4 | Timer: 1min-99h59min
| 35.5 | Heating Time (37OC to 95OC): ≤3min
| 35.6 | Cooling Time (95OC to 45OC): ≤5min
| 35.7 | Capacity: 12 Slides
| 35.8 | Overall dimensions: L420×W225×H143mm
| 35.9 | Net Weight: 6kg
36 |  | Freezer at -60 C to -86 C |  | 1
| 36.1 | Temperature range: min (-60°C to -86°C)
| 36.2 | Net volume min. 90 L
| 36.3 | Dimensions (HxWxD): (830x600x650 mm) ± 5%
| 36.4 | Two adjustable shelves, stainless steel
| 36.5 | Insulation 80mm
| 36.6 | Inner Stainless steel
| 36.7 | High/Low temperature alarm and open door alarm
| 36.8 | Battery Backup min. 48h for display and data logging
| 36.9 | Data logging min. 35.000 entries
| 36.10 | Porthole for external probe Ø12 mm
| 36.11 | Integrated lock with key
| 36.12 | Integrated USB port
| 36.13 | Digital touch display
| 36.14 | Noise level max. 60dB
37 |  | Table centrifuge with 4x6 rotor |  | 1
| 37.1 | Swing-out rotor max capacity 24 tubes for blood and urine max Ø 16 x 108 mm , Max. speed with rotor: 4,000 rpm Max. RCF: 2,379 x g
| 37.2 | Centrifuge speed range up to 4,000 rpm (increment 100 rpm) / RFC 2.540 (increment10 x g)
| 37.3 | Digital LCD display showing operating parameters
| 37.4 | 10 editable progrmans, mikroprocessor control
| 37.5 | Two acceleration and braking curves
| 37.6 | SOFT start /stop function
| 37.7 | Memory of last cycle
| 37.8 | Timer (s; min:s) 10 – 99:59
| 37.9 | Stainless steel chamber
| 37.10 | Automatic imbalance monitoring
| 37.11 | Automatic overspeed monitoring
| 37.12 | Emergency lid release
| 37.13 | Window in lid for external speed monitoring
| 37.14 | LED running indicator, at the bottom of the centrifuge, readily visible at a distance
| 37.15 | Dimensions max 293 x 378 x 535 mm
| 37.16 | Produced in compliance with the latest national and international standards EN 61010-2-020
| 37.17 | Weight without rotor max 23kg
38 |  | Digestor |  | 1
| 38.1 | Dimensions: max 120x80x250 cm
| 38.2 | Ceramic acid-resistant tile worktop with built-in ceramic acid-resistant bath
| 38.3 | Laboratory water tap
| 38.4 | The base of the digester is required to be a metal plasticized base made of steel smoke pipes max 30x30x2 mm.
| 38.5 | The upper part of the digester is required to be formed on both sides by melamine plates coated on the inside with ceramic acid-resistant tiles or Max compact board
| 38.6 | The digester door is required to be sliding. Safety glass in a frame is required to be made of aluminum profiles.
| 38.7 | The digester hood is required to be made of polypropylene with a fi 250 connection ventilation duct.
| 38.8 | The digester is required to be equipped with three single-wire power connections.
| 38.9 | The digester is required to have waterproof LED lighting 2x10W.
| 38.10 | All edges edged with ABS edging tape 2 mm thick, glued with PU waterproof glue.
| 38.11 | Acid-resistant fan for digester.
| 38.12 | Fluid flow: Q = 900 m3 / h
| 38.13 | Motor power required: P = 0.55 kW
| 38.14 | Rotation speed: V = 1,450 rpm
| 38.15 | Damping branches = 250 mm
| 38.16 | Pressure branches = 200 mm
| 38.17 | Mounting kit
| 38.18 | Fan bracket
| 38.19 | Acid-resistant ventilation duct with a diameter of 200 mm.
| 38.20 | DM switch - bimetallic protection for three-phase electric motor
39 |  | UV-VIS Spectrophotometer, range 195-1050 nm |  | 1
| 39.1 | Cuvette Capability: Compatible with most quartz, glass, and plastic cuvettes Linear up to 2.6 Abs
| 39.2 | Full scan: 2.5 - 4 seconds per reading 200 to 900 nm Resolution better than 1.8 nm
| 39.3 | Endless Connectivity: Flexible Unit Control and Ultimate Data Security Built-in File Server for data access from Windows and Mac computers Print to Airprint™ and HP Universal Driver compatible printers as well as DYMO Label printers REST API for LIMS integration
| 39.4 | Battery Powered Up to 8 hours battery operation
| 39.5 | Footprint: (20 x 20 x 12) cm ±5%; for nucleic acids, protein and samples in most organic solvents; upgradable to NanoVolume capability; No reconditioning, no recalibration and no regular maintenance ever; Stand-alone operation with built-in 7 inch ±5% glove compatible touch screen Universal data output: Excel and PDF
| 39.6 | Multi Language User Interface ¦ Barcode ready 32 GB of onboard memory
| 39.7 | Detection Range dsDNA 0.1 - 130 ng/µl Detection Range BSA 0.003 - 3.7 mg/ml Photometric Range 0 - 2.6 A Center Height (Z-Height) 8.5 mm Cell Types Outside dimension 12.5 x 12.5 mm Heating 37 °C ± 0.5 °C
| 39.8 | Operating System Linux based NPOS Onboard Processor Intel Celeron dual core 2.4 GHz Internal Data Storage C40: 128 GB In & Output Ports 2x USB A, USB B, HDMI, Ethernet, WiFi
| 39.9 | Optical Specifications Wavelength Scan Range C40: 200 - 900 nm
| 39.10 | Measure Time For Full Scan Range C40: 2.5 - 4.0 sec
| 39.11 | Wavelength Reproducibility C40: ± 0.2 nm
| 39.12 | Wavelength Accuracy C40: ± 0.75 nm
| 39.13 | Bandwidth C40: < 1.8 nm
| 39.14 | Absorbance Reproducibility C40: < 0.002 A @ 0 - 0.3 A @ 280 nm CV < 1% @ 0.3 - 2.0 A @ 280 nm
| 39.15 | Absorbance Accuracy < 1.75% @ 0.7 A @ 280 nm of the reading Stray Light C40: < 0.5% @ 240 nm using NaI
| 39.16 | Optical Arrangement 1 x 3648 CCD Array
| 39.17 | Lamp ¦ Lifetime Xenon flash lamp ¦ 109 flashes, up to 10 years
| 39.18 | Main Body Size: (200 x 200 x 120) mm ±5%
| 39.19 | Weight (3.5 - 5.5 kg) ±5%
| 39.20 | Operating Voltage 90 - 250 V, 50/60 Hz, 90 W, 18/19
| 39.21 | VDC Display 1024 x 600 pixels; glove compatible touchscreen
| 39.22 | Built-in Battery Pack
| 39.23 | Min. charging cycles: 800
| 39.24 | Certification CE, IEC 61010-1:2012 and EN 61326-1:2013
| 39.25 | Battery Certification IEC 62133 and UN38.3 transport test
| 39.26 | Security Slot for Kensington lock
40 |  | Microcentrifuge refrigerated with rotor 24x1.5 ml |  | 1
| 40.1 | Rotor capacity: 24 x 1.5/2.0 mL vessels
| 40.2 | Max. speed: 15,000 rpm
| 40.3 | Design for quiet opeartion, even without a rotor lid
| 40.4 | Temperature range from -10 °C to +40 °C
| 40.5 | Ensured temperature of 4 °C at maximum rotational speed
| 40.6 | Built-in condensation drain to eliminate water accumulation in the rotor chamber
| 40.7 | 230V/50-60 Hz