## Price Schedule
Bidder: |  | Offered price:
LOT 1 - Imaging
LOT ID | Equipment name | QTY | Model, Manufacturer and Country of Origin | Unit price | Total Price per line item
Hospital Common needs
1 | MOBILE DIGITAL X-RAY DEVICE | 6 |  |  | 0
2 | PORTABLE HIGH CLASS ULTRASOUND | 9 |  |  | 0
3 | PORTABLE HIGH CLASS ULTRASOUND WITH VASCULAR PROBE | 5 |  |  | 0
4 | PREMIUM CLASS ULTRASOUND - Radiology | 2 |  |  | 0
5 | PREMIUM CLASS ULTRASOUND - Internal dpt. | 3 |  |  | 0
6 | Gamma camera with site preparation works | 1 |  |  | 0
7 | CT scanner with site preparaton works | 1 |  |  | 0
8 | CT injector | 1 |  |  | 0
|  |  |  | Total Price per Lot | 0

## LOT1
Bidder:
LOT 1 |  | Imaging |  |  | Date:
Total price:
Line item No. |  | Description | Technical Specification Offered Yes/No | QTY | Unit price | Total Price per line item
1 |  | MOBILE DIGITAL X-RAY DEVICE |  | 6
| 1.1 | High frequency generator, power at least 32 kW
| 1.2 | Voltage in the range of at least 40-125 KV or more
| 1.3 | Maximum current at least 400 mA
| 1.4 | mAs range of at least 0.1-500 mAs or more
| 1.5 | Minimum exposure time 1 msec or less
| 1.6 | Number of possible anatomical programs at least 700
| 1.7 | Exposure from own battery power
| 1.8 | Battery charge level indicator
| 1.9 | Manual exposure with extension cable or wireless
| 1.10 | Charging the battery from a standard 230V socket
| 1.11 | X-RAY TUBE
| 1.12 | Rotating anode, heat capacity of the anode, at least 150 kHU
| 1.13 | Dual focus, small max. 0.6mm / large min. 1.2mm
| 1.14 | Manual or motorized collimator
| 1.15 | Collimator rotation at least +/- 120°
| 1.16 | Buttons on collimator for fine forward / reverse movement for easy positioning ("inch-mover")
| 1.17 | WIRELESS PORTABLE FLET PANEL DETECTOR
| 1.18 | Scintillator material: CsI (cesium iodide)
| 1.19 | Effective detector area min 34x42cm
| 1.20 | A/D conversion minimum 16 bits
| 1.21 | DQE: minimum 63% at 0 lp/mm
| 1.22 | Pixel size: maximum 150μm
| 1.23 | Image matrix at least 2300 x 2800
| 1.24 | Maximum FPD load capacity with evenly distributed load of at least 150kg
| 1.25 | Battery capacity: minimum 5 hours (Full Performance)
| 1.26 | ACQUISITION WORKSTATION BUILT-IN MOBILE X RAY
| 1.27 | Touch-screen monitor
| 1.28 | Monitor screen size at least 17"
| 1.29 | Image storage capacity: at least 20,000 images
| 1.30 | Magnifier
| 1.31 | Noise reduction
| 1.32 | Contrast
| 1.33 | Distance and Angle measurement tool
| 1.34 | Zoom
| 1.35 | Image rotation
| 1.36 | L / R markers
| 1.37 | Annotation
| 1.38 | Images can be send manually or automatically in PACS in DICOM format via "wireless" interface
| 1.39 | DICOM functions at least: Storage, Print, MWL, MPPS
| 1.40 | MECHANICAL PARAMETERS
| 1.41 | Motorized forward / backward on its own battery-powered
| 1.42 | Telescopic tube positioning system
| 1.43 | The distance from the focus of the tube to the floor in the vertical direction range of at least 14000mm +/- 10mm
| 1.44 | Horizontal movement of the tube at least 550mm +/- 10mm
| 1.45 | Column rotation at least +/- 325°
| 1.46 | Rotation of the tube around the telescopic arm at least +/-180 °
| 1.47 | Tube rotation - tilt + 90°/-30°
| 1.48 | System width maximum 600mm
2 |  | PORTABLE HIGH CLASS ULTRASOUND |  | 9
| 2.1 | General Characteristics
| 2.2 | Laptop style portable ultrasound system
| 2.3 | System weight maximum 5 kg (with battery)
| 2.4 | Cart for the system, height adjustable, with min. 3 probe ports
| 2.5 | Recent beamforming technology
| 2.6 | Systemic dynamic range min. 265 dB
| 2.7 | Max. frame rate 1700 fps, probe and mode dependent
| 2.8 | Frequency range 2 - 18 MHz or wider
| 2.9 | High-resolution (min.1920 x 1080 pixels) LCD monitor min.15.6'' with brightness, contrast and color temperature adjustment
| 2.10 | Integrated SSD min. 256 GB
| 2.11 | Min. 8 TGC pods
| 2.12 | Max. number of focal points min. 8, probe and application dependent
| 2.13 | Coded harmonic imaging
| 2.14 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 2.15 | Spatial compounding
| 2.16 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 2.17 | Full alphanumeric keypad covered with washable protection film
| 2.18 | Maximum depth of field min. 30 cm, probe dependent
| 2.19 | Technique which provides a convex field of view on linear and sector transducers
| 2.20 | HD write zoom and read zoom min. 50X
| 2.21 | Real-time automatic Doppler calculations
| 2.22 | Voice comments recording
| 2.23 | Transducer type:
| 2.23.1 | Convex array
| 2.23.2 | Linear array
| 2.23.3 | Microconvex array
| 2.23.4 | Sector phased array
| 2.23.5 | Volume Probe (4D)
| 2.24 | Operating modes:
| 2.24.1 | B-Mode
| 2.24.2 | M-Mode
| 2.24.3 | Color M-Mode
| 2.24.4 | PW Doppler with high PRF
| 2.24.5 | Color Flow mode (CFM)
| 2.24.6 | Power Doppler Imaging (PDI) and Directional PDI
| 2.24.7 | CW Doppler Mode (CWD)
| 2.25 | Auto IMT
| 2.26 | Method for showing hemodynamics in grayscale using B-Mode technology (not the Doppler technology) which is based on reflection of blood cells (direct, real-time visualization of blood flow echoes)
| 2.27 | Hardware and software options which can be added later:
| 2.27.1 | Tissue Velocity Imaging (TVI)
| 2.27.2 | 3D/4D option with the following acquisition and visualization modes: - Static 3D - Real-time 4D - Tomographic imaging (TUI or equivalent)
| 2.28 | Inputs and outputs: min. 3 x USB, HDMI output (1920 x 1080), Ethernet (RJ45)
| 2.29 | Storage formats: - DICOM – compressed/uncompressed, single/multi-frame, with/without raw data - Export JPEG, WMV (MPEG 4) and AVI formats
| 2.30 | DICOM 3.0 connectivity
| 2.31 | B/W thermal printer
| 2.32 | Probes:
| 2.32.1 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 2.32.2 | Sector probe Frequency range: 2 – 4 MHz or wider Field of view (FOV): min. 120°
| 2.32.3 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
3 |  | PORTABLE HIGH CLASS ULTRASOUND WITH VASCULAR PROBE |  | 5
| 3.1 | General Characteristics
| 3.2 | Laptop style portable ultrasound system
| 3.3 | System weight maximum 5 kg (with battery)
| 3.4 | Cart for the system, height adjustable, with min. 3 probe ports
| 3.5 | Recent beamforming technology like Quad beamforming
| 3.6 | Systemic dynamic range min. 265 dB
| 3.7 | Max. frame rate 1700 fps, probe and mode dependent
| 3.8 | Frequency range 2 - 18 MHz or wider
| 3.9 | High-resolution (min.1920 x 1080 pixels) LCD monitor min.15.6'' with brightness, contrast and color temperature adjustment
| 3.10 | Integrated SSD min. 256 GB
| 3.11 | Min. 8 TGC pods
| 3.12 | Max. number of focal points min. 8, probe and application dependent
| 3.13 | Coded harmonic imaging
| 3.14 | High definition speckle reduction imaging which provides multiple levels (min. 8) of speckle reduction
| 3.15 | Spatial compounding
| 3.16 | Function which continuously optimize the brightness, contrast and uniformity of B-Mode images and in PW/CW Doppler mode on live image auto adjusts the velocity scale (PRF), baseline shift, and invert
| 3.17 | Full alphanumeric keypad covered with washable protection film
| 3.18 | Maximum depth of field min. 30 cm, probe dependent
| 3.19 | Technique which provides a convex field of view on linear and sector transducers
| 3.20 | HD write zoom and read zoom min. 50X
| 3.21 | Real-time automatic Doppler calculations
| 3.22 | Voice comments recording
| 3.23 | Transducer type:
| 2.23.1 | Convex array
| 2.23.2 | Linear array
| 2.23.3 | Microconvex array
| 2.23.4 | Sector phased array
| 2.23.5 | Volume Probe (4D)
| 3.24 | Operating modes:
| 3.24.1 | B-Mode
| 3.24.2 | M-Mode
| 3.24.3 | Color M-Mode
| 3.24.4 | PW Doppler with high PRF
| 3.24.5 | Color Flow mode (CFM)
| 3.24.6 | Power Doppler Imaging (PDI) and Directional PDI
| 3.25 | Auto IMT
| 3.26 | Method for showing hemodynamics in grayscale using B-Mode technology (not the Doppler technology) which is based on reflection of blood cells (direct, real-time visualization of blood flow echoes)
| 3.27 | Hardware and software options which can be added later:
| 3.27.1 | Elastography on linear probe
| 3.28 | Inputs and outputs: min. 3 x USB, HDMI output (1920 x 1080), Ethernet (RJ45)
| 3.29 | Storage formats: - DICOM – compressed/uncompressed, single/multi-frame, with/without raw data - Export JPEG, WMV (MPEG 4) and AVI formats
| 3.30 | DICOM 3.0 connectivity
| 3.31 | B/W thermal printer
| 3.32 | Probes:
| 3.32.1 | Convex probe Frequency range: 2 – 5 MHz or wider Field of view (FOV): min. 55°
| 3.32.2 | Linear probe Frequency range: 4 – 13 MHz or wider Field of view (FOV): 40 mm (+/- 5 mm)
4 |  | PREMIUM CLASS ULTRASOUND - Radiology |  | 2
| 4.1 | General Characteristcs
| 4.2 | Ultrasound system, not portable, with wheels and front and rear handles
| 4.3 | Four active probe ports, not counting pencil probe connector
| 4.4 | OLED monitor min. 22'', wide screen, high-resolution (min. 1920 x 1080) with translation (independent of console), swivel and fold down
| 4.5 | Wide, high-resolution, color touch LCD screen, min. 10,” with interactive menu, brightness adjustment and user-configurable layout
| 4.6 | Operating keyboard adjustable in height and rotation
| 4.7 | Backlit alphanumeric keyboard
| 4.8 | Integrated SSD with disk encryption (encrypts patient data archive partition)
| 4.9 | Integrated DVD multi drive
| 4.10 | Integrated speakers
| 4.11 | Integrated gel warmer
| 4.12 | Maximum depth of field min. 30 cm (probe dependent)
| 4.13 | Maximum dynamic range min. 400 dB
| 4.14 | Maximum framerate min. 2900 fps (Hz)
| 4.15 | Coded harmonic imaging
| 4.16 | Supported transducer types:
| 4.16.1 | Convex array
| 4.16.2 | Micro convex array
| 4.16.3 | Bi-plane micro convex probe
| 4.16.4 | Linear array
| 4.16.5 | Sector phased array
| 4.16.6 | Matrix array
| 4.16.7 | Intraoperative probes
| 4.16.8 | Biopsy convex probe
| 4.16.9 | Volume probes (convex array, micro convex array, linear array)
| 4.16.10 | TEE probe
| 4.16.11 | Split crystal probes
| 4.17 | Operating modes:
| 4.17.1 | B-Mode
| 4.17.2 | M-Mode
| 4.17.3 | Anatomical M-Mode
| 4.17.4 | Curved Anatomical M-Mode
| 4.17.5 | Color Flow Mode (CFM)
| 4.17.6 | Power Doppler Imaging (PDI)
| 4.17.7 | PW Doppler
| 4.18 | 3D acquisition with 2D probes including acquisition of color data as Advanced 3D or corresponding
| 4.19 | Convex field of view available on linear and sector probes
| 4.20 | Speckle reduction imaging which provides multiple levels of speckle reduction
| 4.21 | Variable angle spatial compounding
| 4.22 | Automatic optimization of B-Mode image to improve axial and lateral gain uniformity, overall gain level and contrast resolution
| 4.23 | Auto spectral optimization – adjusts baseline, invert, PRF (on live image) and angle correction
| 4.24 | Method for showing hemodynamics in grayscale using B-Mode technology (not the Doppler technology) which is based on reflection of blood cells (direct, real-time visualization of blood flow echoes)
| 4.25 | Contrast imaging with time intensity curve analysis
| 4.26 | Strain elastography on all offered probes
| 4.27 | Elastography quantification
| 4.28 | On-board reporting package for report writing
| 4.29 | DICOM 3.0
| 4.30 | CINE memory min. 700 MB
| 4.31 | Probes:
| 4.31.1 | Convex probe Bandwidth: 1 – 6 MHz or wider Number of elements: min. 192 Field of view (FOV): min. 70°
| 4.31.2 | Matrix linear probe Bandwidth: 4 – 15 MHz or wider Number of elements: min. 1000 Field of view (FOV): 50 mm (+/- 5 mm)
| 4.31.3 | Linear probe Bandwidth: 2 – 8 MHz or wider Number of elements: min. 192 Field of view (FOV): 45 mm (+/- 5 mm)
| 4.31.4 | Micro convex probe Bandwidth: 3 – 10 MHz or wider Number of elements: min. 192 Field of view (FOV): min. 145°
| 4.32 | Multiple USB 2.0 and USB 3.0 ports
| 4.33 | Ethernet
| 4.34 | DVI-D or HDMI connector
| 4.35 | Digital BW thermal printer
5 |  | PREMIUM CLASS ULTRASOUND - Internal dpt. |  | 3
| 5.1 | General Characteristcs
| 5.2 | Ultrasound system, not portable, with wheels and front and rear handles
| 5.3 | Four active probe ports, not counting pencil probe connector
| 5.4 | OLED monitor min. 22'', wide screen, high-resolution (min. 1920 x 1080) with translation (independent of console), swivel and fold down
| 5.5 | Wide, high-resolution, color touch LCD screen, min. 10,” with interactive menu, brightness adjustment and user-configurable layout
| 5.6 | Operating keyboard adjustable in height and rotation
| 5.7 | Backlit alphanumeric keyboard
| 5.8 | Integrated SSD with disk encryption (encrypts patient data archive partition)
| 5.9 | Integrated DVD multi drive
| 5.10 | Integrated speakers
| 5.11 | Integrated gel warmer
| 5.12 | Maximum depth of field min. 30 cm (probe dependent)
| 5.13 | Maximum dynamic range min. 400 dB
| 5.14 | Maximum framerate min. 2900 fps (Hz)
| 5.15 | Coded harmonic imaging
| 5.16 | Supported transducer types:
| 5.16.1 | Convex array
| 5.16.2 | Micro convex array
| 5.16.3 | Bi-plane micro convex probe
| 5.16.4 | Linear array
| 5.16.5 | Sector phased array
| 5.16.6 | Matrix array
| 5.16.7 | Intraoperative probes
| 5.16.8 | Biopsy convex probe
| 5.16.9 | Volume probes (convex array, micro convex array, linear array)
| 5.16.10 | TEE probe
| 5.16.11 | Split crystal probes
| 5.17 | Operating modes:
| 5.17.1 | B-Mode
| 5.17.2 | M-Mode
| 5.17.3 | Anatomical M-Mode
| 5.17.4 | Curved Anatomical M-Mode
| 5.17.5 | Color Flow Mode (CFM)
| 5.17.6 | Power Doppler Imaging (PDI)
| 5.17.7 | CW Doppler
| 5.17.8 | PW Doppler
| 5.18 | 3D acquisition with 2D probes including acquisition of color data as Advanced 3D or corresponding
| 5.19 | Convex field of view available on linear and sector probes
| 5.20 | Speckle reduction imaging which provides multiple levels of speckle reduction
| 5.21 | Variable angle spatial compounding
| 5.22 | Automatic optimization of B-Mode image to improve axial and lateral gain uniformity, overall gain level and contrast resolution
| 5.23 | Auto spectral optimization – adjusts baseline, invert, PRF (on live image) and angle correction
| 5.24 | Method for showing hemodynamics in grayscale using B-Mode technology (not the Doppler technology) which is based on reflection of blood cells (direct, real-time visualization of blood flow echoes)
| 5.25 | Contrast imaging with time intensity curve analysis
| 5.26 | Strain elastography on all offered probes
| 5.27 | Elastography quantification
| 5.28 | On-board reporting package for report writing
| 5.29 | DICOM 3.0
| 5.30 | CINE memory min. 700 MB
| 5.31 | Probes:
| 5.31.1 | Convex probe Bandwidth: 1 – 6 MHz or wider Number of elements: min. 192 Field of view (FOV): min. 70°
| 5.31.2 | Matrix linear probe Bandwidth: 4 – 15 MHz or wider Number of elements: min. 1000 Field of view (FOV): 50 mm (+/- 5 mm)
| 5.31.3 | Sector matrix array probe Bandwidth: 1 – 5 MHz or wider Number of elements: min. 288 Field of view (FOV): min. 120º
| 5.32 | Multiple USB 2.0 and USB 3.0 ports
| 5.33 | Ethernet
| 5.34 | DVI-D or HDMI connector
| 5.35 | Digital BW thermal printer
6 |  | Gamma camera with site preparation works |  | 1
| 6.1 | Dual head gamma camera with flexible positioning including upright, seated, standing and imaging patients on stretchers
| 6.2 | Whole-body, SPECT, dynamic and gated examinations
| 6.3 | Real-time infrared body contouring for whole body examinations
| 6.4 | SPECT scanning with automatic body contouring for bone, neuro and cardiac examinations
| 6.5 | Automatic and manual radial adjustment of the distance between detectors and patient for bone, neuro and cardiac examinations
| 6.6 | Cardiac monitor for gated examinations
| 6.7 | Preprogrammed positions for easy patient positioning
| 6.8 | Monitor on the NM gantry for patient positioning details and scanning parameters
| 6.9 | Patient table interactive touch-ruler for easy setup of scan range
| 6.10 | Remote service diagnostic based on encrypted internet line
| 6.11 | Online application support for protocols optimization before and during ongoing examination
| 6.12 | Patient Table
| 6.13 | Carbon fiber tabletop with patient accessories for comfort
| 6.14 | Whole body scan range, not less than 200cm
| 6.15 | Table load capacity, at least 220kg
| 6.16 | Accessories for patient positioning - Straps and Pads, Head holder, Arm Support, Leg Rest support
| 6.17 | Detectors characteristics
| 6.18 | Real-time corrections: at least for Center of Rotation, Uniformity, Energy, Linearity, Isotope decay
| 6.19 | Crystal thickness, at least 9.5mm (3/8")
| 6.20 | UFOV, at least 53x39 cm
| 6.21 | Energy range, at least 40-600 keV
| 6.22 | Detectors characteristics, according to NEMA NU-1 2007 /or newer version/
| 6.23 | Intrinsic Energy Resolution FWHM - 3/8", maximum 9.5%
| 6.24 | Intrinsic Spatial Resolution - 3/8"
| 6.25 | CFOV: FWHM, max 3.7mm
| 6.26 | CFOV: FWTM, max 7.0mm
| 6.27 | UFOV: FWHM. Max 3.8 mm
| 6.28 | UFOV: FWTM, max 7.0 mm
| 6.29 | Flood Field Uniformity - 3/8"
| 6.30 | CFOV: Differential, max 2.3%
| 6.31 | CFOV: Integral, max 3%
| 6.32 | UFOV: Differential, max 2.3%
| 6.33 | UFOV: Integral, max 3.6%
| 6.34 | Intrinsic Spatial Linearity - 3/8"
| 6.35 | CFOV: Differential, max 0.2 mm
| 6.36 | CFOV: Integral, max 0.5 mm
| 6.37 | UFOV: Differential, max 0.2 mm
| 6.38 | UFOV: Integral, max 0.5 mm
| 6.39 | SPECT console and acquisition
| 6.40 | Dicom 3.0 compatible, print, send, store
| 6.41 | Intercom bidirectional communication with the patient
| 6.42 | Collimators
| 6.43 | LEHR (a pair) with cart
| 6.44 | HEGP/HEAP (a pair) with cart
| 6.45 | Pinhole collimator with cart
| 6.46 | On site field upgradeability
| 6.47 | Upgradable to a high-end SPECT/CT system, CE marked model, with CZT based nuclear detectors and CT system with:
- minimum 17lp/cm high contract resolution
- minimum 6 MHU real anode heat storage capacity
- maximum 1.5mm2 large focal spot size
| 6.48 | Stand-alone diagnostic workstation
| 6.49 | Dual medical monitors with min 24" and min resolution 1680x1050
| 6.50 | Software package for fusion of CT and SPECT images
| 6.51 | Software that allows reconstruction and processing with noise correction, attenuation correction, scatter correction, motion correction and resolution recovery
| 6.52 | Importing SPECT, CT and hybrid images from multi-vendor systems using CD/DVD and network
| 6.53 | Post processing of CT, NM and PETCT exams
| 6.54 | Advanced Cardiac package -analysis and quantification program for gated and ungated SPECT perfusion studies
| 6.55 | Software for analysis of Bone, Nephrology, Gastroenterology, Pulmonary, Neurology and Endocrinology examinations
| 6.56 | CD/DVD archiving with Dicom viewer
| 6.57 | Backup electric power system (UPS) that matches gamma camera
| 6.58 | Site preparation in accordance to the Gamma camera manufacturer requirements
7 |  | CT scanner with site preparaton works |  | 1
| 7.1 | HIGH VOLTAGE GENERATOR
| 7.2 | Max Generator power (nominal hardware): at least 72 kW
| 7.3 | mA range: at least 20-560 mA or more, with 5mA increments
| 7.4 | kV range for radiography: 80 - 140 kV or wider
| 7.5 | X-RAY TUBE
| 7.6 | Number of tube focuses, at least two
| 7.7 | Real heat capacity of the tube anode (not equivalent value), minimum 7 MHU
| 7.8 | X-ray tube cooling rate: min 1 MHU/min
| 7.9 | RADIOGRAPHIC PATIENT TABLE
| 7.1 | Patient table load capacity: at least 220 kg
| 7.11 | Horizontal scanning range min. 200cm
| 7.12 | Coronal imaging head holder.
| 7.13 | Hand holder for easier patient positioning
| 7.14 | GANTRY
| 7.15 | Aperture min. 70cm
| 7.16 | Shortest time for one full scan rotation: max. 0,35sec not only for Cardiac
| 7.17 | DETECTOR SYSTEM
| 7.18 | At least 64 detector rows
| 7.19 | At least 128 reconstruction layers (slice) per rotation
| 7.2 | Minimum layer (slice) thickness: max. 0.625mm
| 7.21 | Spatial resolution: max 0.30mm
| 7.22 | High Contrast Resolution not less than 12 lp/cm (X-Y-Plane at 50%MTF)
| 7.23 | Total number of physical detector elements, at least 47000
| 7.24 | ACQUISITION PARAMETERS / ACQUISITION WORKSTATION - OPERATING CONSOLE
| 7.25 | Field of view for scanning min. 50cm
| 7.26 | Trigger for acquisition when bolus arrives to ROI
| 7.27 | Real Time MPR reconstruction
| 7.28 | CT body/brain 4D perfusion software package
| 7.29 | Iterative reconstruction technology that allows dose reductions with the same diagnostic image quality (iDOSE, SAFIRE, ASIR, AIDR 3D or equivalent).
| 7.3 | System for correction of artifacts from the presence of metals in the body
| 7.31 | Basic CT viewers with tools for fast 3D postprocessing (3D volume rendering, MIP, MPR, VRT)
| 7.32 | Advanced non-invasive CT angiography (CTA) analysis tools with automatic bone removal
| 7.33 | Cardiology study package
| 7.34 | DICOM protocol for network print, sent, query/retrieve
| 7.35 | DICOM Media study archive on CD/DVD with DICOM viewer
| 7.36 | Injector class IV synchronization interface
| 7.37 | Image Reconstruction speed: min 55 images per second
| 7.38 | Technology applicable for Low Dose Screening of lung cancer and high-risk patients
| 7.39 | Body/brain perfusion
| 7.4 | Extended range for 4D perfusion scans: min 140mm coverage
| 7.41 | High performance console host computer with minimum 36GB RAM and minimum 8 CPU cores
| 7.42 | ECG Monitor synchronized with CT system for cardiac prospective gating
| 7.43 | DIAGNOSTIC SERVER WORKSTATION with three (3) client workplaces
| 7.44 | 3 pcs. WORKSTATIONS (Clients) with LCD / TFT, radiological DICOM calibrated monitors with a diagonal of monitor min 24"
| 7.45 | Dynamic licenses distribution between clients (floatable licenses)
| 7.46 | Server image database storage, at least 2 TB
| 7.47 | Server must be equipped with the following permanent (unlimited time of use) licenses:
| 7.48 | 3D anatomy analysis and display tools (3D volume rendering, MIP, MPR, VRT), at least three licenses
| 7.49 | Advanced non-invasive CT angiography (CTA) analysis tools with automatic bone removal, at least three licenses
| 7.5 | Software for assessment and automatic visualization and monitoring of pulmonary nodules and diagnosis -CT LNA (lung nodule analysis), at least two licenses
| 7.51 | Advanced tools for analysis of lung disease (e.g. chronic obstructive pulmonary disease), at least two licenses
| 7.52 | Oncology package for automatic multimodal tumor monitoring, automatic calculation following WHO criteria, RECIST 1.1., at least two licenses
| 7.53 | Multimodality Fusion package (image fusion between CT and MR, SPECT, PET), at least two licenses
| 7.54 | Cardiology study package (calcium score, CT analysis of coronary blood vessels, functional cardiac analysis), at least one license
| 7.55 | Advanced tools for analysis of neurological exams (CT brain perfusion, stroke assessment and analysis of aneurysms), at least two licenses
| 7.56 | CT perfusion for body examinations, at least one license
| 7.57 | Site preparation in accordance to the CT scanner manufacturer requirements
8 |  | CT injector |  | 1
| 8.1 | Automatic CT injector with three head for injecting contrast media and saline solution
| 8.2 | Operation by touch screen with operation modes for CT procedures
| 8.3 | Need to have a system solution that recognizes the air in the system and warns the user with an audible and visual alarm about the appearance of air in order to prevent embolism, possession of air detection sensors
| 8.4 | Integrated heater for contrast media
| 8.5 | Communication with bluetoth or wi-fi technology
| 8.6 | Flow rate between 0.1 to 10 ml/sec