## 1. HR SEMs
Bidder: |  | Offered price:
LOT01 - High Resolution Scanning Electron Microscopes
ID | Equipment name | QTY | Model, Manufacturer and Country of Origin | Unit Price | Total Price per line item
High Resolution Scanning Electron Microscopes
1.1 | HR dual beam system | 1 |  |  | 0
1.2 | HR SEM | 1 |  |  | 0
|  |  | Total Price per Lot |  | 0

## 1.1 HR dual beam system
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
High Resolution Scanning Electron Microscopes
| 1.1 | High Resolution Dual Beam System |  | 1 |  | 0
| 1 | The SEM must be a Field Emission Scanning Electron Microscope (FESEM), with a high brightness gun capable of delivering: • Landing voltage range of 20 eV to 30 keV, continuously variable • A beam current up to at least 100 nA
| 2 | Secondary electron image resolutions must meet the following minimum resolution: • 0.6 nm or better at 15 kV • 0.7 nm or better at 1 kV • 1.0 nm or better at 500 V
| 3 | The system should be equipped with the following detection features: • Everhart-Thornley SE detector • At least 3 In-lens/in column detectors for simultaneous detection of SE and both low loss and no loss BSE • In-lens SE and BSE signals should be detected and displayed simultaneously • Integrated Camera for taking a (color) image of the samples on the stage which can then be used for navigation • Retractable backscatter detector based on solid state technology and consist of a concentric ring design allowing the selection of different contrast and signal based on landing voltage and collection angle • Integrated beam current monitor • A live IR CCD image for viewing the samples and chamber must be simultaneously visible with a live SED image. Images must be displayed within the SEM’s standard graphical user interface and must be able to be shown full screen or within one of four quadrants. Automatic shutdown of IR source illumination in BSE mode is required to allow the light sensitive BSE to work correctly. • Retractable STEM detector with at least 11 segments capable of BF/DF and HAADF imaging in order to check the quality of the TEM lamella in situ. • A FIB optimized, secondary ion and electron detector
| 4 | The system should be equipped with the Energy Dispersive X-ray Spectrometer (EDS). The EDS system, hardware and unique software, should be ready for integration with EBSD and TKD, parallel beam WDS and Micro-XRF. The EDS should have following minimal specification: • Detector - LN2-free high-resolution, high-speed SD Detector - Energy resolution ≤129eV at Mn K - Active area ≥ 60 mm2 - Maximum input count rate > 1.500.000 cps - Slew window, detection from Beryllium (4) to Americium (95) - Motorized retraction system • Base System - Signal Processing Unit for high output count rates - Fully automatic software controlled calibration of hardware - Workstation Computer, 24“ TFT LCD Monitor, incl. Windows 10 • Software - Software for spectra acquisition and qualitative analysis. - Automatic software tools for standardless quantitative spectra analysis. - Editor for user defined spectra evaluation methods. - Automatic multipoint and object analysis. - Fast spectral line scan with spectra data base. - Fast digital element mapping with unlimited number of elements. - Data management and filing system. - Report generation and print formatting
| 5 | The system should be equipped with the EBSD system that can be used in combination with the EDS system to acquire EDS spectra to be considered for phase selection prior to EBSD map acquisition and to make Transmission Kikuchi Diffraction (TKD). The EDS should have followed minimal specification: • Detector - Tapered nose design to allow simultaneous acquisition of EBSD and EDS signals in best geometry possible - In-situ vertical screen positioning using detector tilt by up to ±4.5 degrees - Detector tilt angle automatically read by electronic sensor - 34 x 25.5 mm rectangular Phosphor screen - Motorized insertion and retraction, maximum speed 10 mm/s, position accuracy better than 0.01 mm. - Manual insertion/retraction using key panel on both detector sides - Native image resolution: 640 x 480 pixels, 12bit CCD camera - 630 patterns per second (pps) at 4 x 4 binning, image size 160 x 120 pixels - 930 pps at 8 x 8 binning, image size 80 x 60 pixels • Software - Multithreaded software for simultaneous acquisition and ultrafast data processing. - Automatic optimization of the detector signal considering current image size (binning), automatic camera gain, definition of the region of interest (ROI), static or dynamic background correction and pattern quality meter - Automatic calibration of the pattern center using the current sample, including dynamic pattern center correction - Data acquisition and indexing features: - Quality control feature displaying the acquired and processed information at the mouse cursor position in real time (available during acquisition as well as offline): - Integrated phase editor including visualization of crystal structures - Orientation and phase distribution features - Grain detection and grain boundary features - Orientation distribution (crystallographic texture) representation features - Data storage and export/import options - Databases and search engines
| 6 | The system should be equipped with the following features: • Active vibration isolation system • Cryo-decontamination device to reduce contamination found in the chamber • Integrated plasma cleaner for both cleaning the sample inside the chamber and for cleaning the inside of the microscope chamber.
| 7 | The microscope stage should have the following minimum specifications: • Fully motorized high-precision 5 axis motorized stage • x, y range: 150 mm • z range: 10 mm • Tilt: -10° to +60° • Rotation: 360° • The repeatability of the X and Y movements must be 1 um or better • Stage must have compucentric rotation and tilt
| 8 | The FIB should be capable of delivering up to 65 nA in a tight, perfectly round spot, and guaranteeing a beam current density of at least 39 A/cm2 at 30 kV, 1 nA, for highest quality milling during fast FIB material removal
| 9 | The FIB column should be equipped with an automated aperture strip for controlling the ion beam current. The aperture strip should have a minimum of 15 apertures.
| 10 | The FIB should operate from 500 V to 30 kV acceleration voltage range.
| 11 | The FIB resolution should be 2.5 nm or better at 30 kV using the selective edge method to measure the resolution
| 12 | to ensure optimum low voltage performance the FIB should have 2 stage differential pumping and a time of flight correction.
| 13 | The microscope should allow visualization of the FIB milling process, using • Simultaneous patterning and imaging • Intermittent patterning and imaging • SE/SI signal visible inside the pattern area during ion beam milling
| 14 | The system should be able touse the electron beam to neutralize the charge created by Ga+ ions in automated, software-controlled manner when milling on an insulating sample
| 15 | System must have 5 Gas Injection Systems for C, Pt, Au, Silica, XeF2.
| 16 | The system should come with a fully integrated nanomanipulator for lifting out lamella prepared for SEM. The nanomanipulator should be manufactured by the SEM supplier, port based and have the following minimum specifications to enable simple and repeatable lift out: • Drift < 200 nm/min • True Z movement over 5 um move < 500 nm • Vibration <15 nm • Omnidirectional repeatability < +/- 150 nm • Smallest step size 50 nm • Motorized 360 control of needle rotation
| 17 | Single or four-quadrant frame display of live, line averaged, line integrated, frozen, or frame-averaged digital images. Secondary Electron (SE), Backscattered Secondary Electron (BSE) mixed SE/BSE, CCD images can be viewed live in the four-quadrant display.
| 18 | The microscope must have a comprehensive and versatile scan system to enable appropriate strategies for a variety of samples. More specifically it should enable frame averaging, line averaging and interlaced scanning for dealing with challenging samples.
| 19 | The microscope must also enable drift corrected frame integration to overlay frames with an offset to adjust for any sample or system instability (such as slight image drift due to charging) to yield better images in some situations.
| 20 | It should be possible to capture 4 live images from different detectors or detector segments from a single scan of the electron beam
| 21 | The system should be delivered with the following software: • Software for unattended automatic preparation of multiple TEM samples. • Automation software for automated serial sectioning and imaging through a user-defined volume of a specimen. • Software to enable the user to control columns, detectors, stage motion and site alignment, with logical operation and looping in order to generate protocols for various automated processes. • Software to enable accurate patterning of large and complex nanostructures on multiple sites, supporting the following patterning processes: focused ion beam (FIB) milling, gas-assisted FIB milling, FIB induced deposition, and electron-induced deposition. • Software enabling the automated acquisition of multiple mosaic acquisitions. • 3D visualization and segmentation software
| 22 | SEM must be supplied with a USB Manual User Interface (MUI) allowing adjustment of focus, magnification, contrast, brightness, beam shift, and stigmation via knob controls.
| 23 | System must be supplied with required peripherals such as compressor, chiller and UPS.
| 24 | 5 days onsite training conducted by an employee of the manufacturer.
| 25 | The system should be delivered with the software enabling the automated acquisition of multiple mosaic acquisitions.
| 26 | 1 Spare FEG Source including labor for replacement after warranty expires (this additional Source will be delivered, any time at customer need it, after warranty expires)
| 27 | 1 Spare Liquid Metal Ion Source including labor for replacement after warranty expires (this additional Source will be delivered, any time at customer need it, after warranty expires)
| 28 | The microscope must be supplied with sample manipulation kit, at least next items should be included: • SEM pin stubs aluminum-pack of 100 and stub storage boxes-25psc • Adhesive Carbon Tape 12mm x 20m and Adhesive Carbon Tabs 12mm (Pack of 100) • HP Tweezers 5 Stainless Steel. 0.10 x 0.06mm tip • Tweezers, opening 1/4 - 3/8 inch and Tweezers with flat tip • Copper lift-out grids with 5 posts, pack of 100
| 29 | The microscope must be supplied with next spare parts (these additional parts will be delivered, any time at customer need it, after warranty expires): • Platinum Reservoir for Platinum Gas Injection System • Carbon Reservoir for Carbon Gas Injection System • Gold (Au) Refill for Gold Gas Injection System • Insulator Refill for Insulator Deposition Gas Injection System • Insulator Refill for Insulator Etch Gas Injection System • Extractor for Ion column • Molybdenum Lift-Out Grids with 3 posts • Tip seals for dry pump • Oil cartridge for TMP
| 30 | Anti-vibration Table to reduce room vibration to level acceptable for Microscope.
| 31 | The system is supplied with operating manual, labels for plugs and any necessary instruction for its operation.
| 32 | Availability of the spare parts for 8 years.
| 33 | Warranty: one year for the hardware, from installation and test, including all parts and their performances.
| 34 | Delivery time: Within 6 months
| 35 | Response time for failure notification – 24 (twenty-four) hours
| 36 | Maximum period for repair and replacement from the moment of declaring 6 weeks

## 1.2 HR SEM
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
High Resolution Scanning Electron Microscopes
| 1.2 | High Resolution Scanning Electron Microscope |  | 1 |  | 0
| 1 | The SEM must be a Field Emission Scanning Electron Microscope (FESEM), with a high brightness gun capable of delivering: • 0.1% beam current stability / h or better • Landing voltage range of 20 eV to 30 keV, continuously variable • A beam current up to at least 50 nA at 10 kV
| 2 | Secondary electron image resolutions must meet the following minimum resolution: • High Vacuum SE - 0.9 nm or better at 15 kV - 1.0 nm or better at 1 kV • Low vacuum SE - 1.2 nm or better at 15 kV - 1.8 nm or better at 3 kV
| 3 | The electron beam column should be equipped with an automated sliding aperture system for controlling the beam current.
| 4 | It should be possible to bias the sample positively up to +600 V to help manage charge on non-conductive specimens when performing in-lens detection.
| 5 | Microscope must be capable of operating in low and high vacuum with the following minimum vacuum requirements: • Chamber vacuum (high vacuum) < 1 × 10-5 mbar • Low vacuum 10 to 500 Pa • Evacuation time 3.5 minutes or better • Oil free vacuum system
| 6 | The system should be equipped with the following detection features: • Everhart-Thornley SE detector • Low vacuum SE detector • In-lens/in-columns SE detector • In-lens/in-columns BSE detector • In-lens SE and BSE signals should be detected and displayed simultaneously • 10 MPixel Integrated Camera for taking a (color) image of the samples on the stage which can then be used for navigation without the need for manual alignment • Software controlled retractable backscatter detector • Integrated beam current monitor • A live IR CCD image for viewing the samples and chamber must be simultaneously visible with a live SED image. Automatic shutdown of IR source illumination in BSE mode is required to allow the light sensitive BSE detector to work correctly. • The software controlled retractable backscatter detector should be based on solid state technology and offer flexible segmentation enabling the selection of either concentric or angular sectors. The segmentation mode should be selectable in the user interface. • EDX detector with an active area of at least 30 mm2 and a Mn Ka resolution of 129 eV or better.
| 7 | The microscope stage should have the following minimum specifications: • Fully motorized high-precision 5 axis motorized stage • x, y range: 110 mm • z range: 65 mm • Tilt: -15° to +90° • Rotation: 360° • Max sample weight – 5 kg • The repeatability of the X and Y movements must be 3.0 um or better • The stage should be mechanically tilt eucentric at the analytical working distance • Stage must have compucentric rotation and tilt
| 8 | The microscope must be supplied with sample holder/s capable of mounting: • 18 standard 12 mm stubs simultaneously • 3 pre tilted 12 mm stubs • Clamp for mounting cross sections • Row bar holder for mounting STEM specimens
| 9 | Single or four-quadrant frame display of live, line averaged, line integrated, frozen, or frame-averaged digital images. Secondary Electron (SE), Backscattered Secondary Electron (BSE) mixed SE/BSE, CCD images can be viewed live in the four-quadrant display.
| 10 | The microscope should enable frame averaging, line averaging and interlaced scanning.
| 11 | The microscope must enable drift corrected frame integration to overlay frames with an offset to adjust for any sample or system instability (such as slight image drift due to charging) to yield better images in some situations.
| 12 | It should be possible to capture 4 live images from different detectors or detector segments from a single scan of the electron beam.
| 13 | SEM must be supplied with a USB Manual User Interface (MUI) allowing adjustment of focus, magnification, contrast, brightness, beam shift, and stigmation via knob controls.
| 14 | SEM must be supplied with an integrated Plasma cleaner for cleaning samples and also the specimen chamber.
| 15 | System must be supplied with required peripherals such as compressor, chiller and UPS.
| 16 | 5 days onsite training conducted by an employee of the manufacturer.
| 17 | The system should be delivered with the software enabling the automated acquisition of multiple mosaic acquisitions.
| 18 | 1 Spare FEG Source including labor for replacement after warranty expires (this additional Source will be delivered, any time at customer need it, after warranty expires)
| 19 | The microscope must be supplied with sample manipulation kit, at least next items should be included: • SEM pin stubs aluminium-pack of 100 and stub storage boxes-25psc • Adhesive Carbon Tape 12mm x 20m and Adhesive Carbon Tabs 12mm (Pack of 100) • HP Tweezers 5 Stainless Steel. 0.10 x 0.06mm tip • Tweezers, opening 1/4 - 3/8 inch and Tweezers with flat tip
| 20 | The microscope must be supplied with next spare parts (this additional parts will be delivered, any time at customer need it, after warranty expires): • Oil cartridge for TMP • Tip seals for dry pump • Coolant 10l for chiller
| 21 | Anti-vibration Table to reduce room vibration to level acceptable for Microscope.
| 22 | The system is supplied with operating manual, labels for plugs and any necessary instruction for its operation.
| 23 | Availability of the spare parts for 8 years.
| 24 | Warranty: one year for the hardware, from installation and test, including all parts and their performances.
| 25 | Delivery time: Within 6 months
| 26 | Response time for failure notification – 24 (twenty-four) hours
| 27 | Maximum period for repair and replacement from the moment of declaring 6 weeks