## 2. E-beam lithography
Bidder: |  | Offered price:
LOT02 - E-beam lithography
ID | Equipment name | QTY | Model, Manufacturer and Country of Origin | Unit Price | Total Price per line item
E-beam lithography
2.1 | E-beam lithography system | 1 |  |  | 0
|  |  | Total Price per Lot |  | 0

## 2.1 E-beam lithography system
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
E-beam lithography
| 2.1 | E-beam lithography system |  | 1 |  | 0
| 1 | General description: • Ultra High Resolution Electron Beam Lithography and Imaging Workstation; System is capable of advanced nano-lithography and imaging applications. • Shipment includes all auxiliaries to run the system including an Uninterrupted Power Supply (UPS) and a system water chiller with closed circuit. • Site Survey (EM noise, acoustic and floor vibration measurements) on site included. • Online (remote) factory acceptance test result and machine review before shipping. • Shipment package DAP customer site included. On-site system installation, qualification & training, acceptance included. All materials for acceptance test on site included. Electron beam lithography starter kit with samples provided (for customer exercising).
| 2 | Electron optics system: • Thermal field emission (TFE) Schottky source. • Fully digital electron optics column, beam energy selectable between 20eV - 30keV. • Fast electrostatic beam-blanking. • User defined continuously adjustable write fields from 0,5 micron to 2 mm with automated calibration and selection within seconds. • Beam size 1,6 nm at 20 kV. 4 nm at 1 kV. • Beam current stability: ≤ 0,5% / 8 hours (at room temperature ±0,5°) • Beam position stability: ≤ 300nm/ hour (at room temperature ±0,5°) • Secondary Electron (SE) Detectors: Everhart-Thornley SE Detector and highest efficiency in-lens SE Detector. • Dual-detector for secondary and backscattered electrons (Inlens-Detector) • Tool provides state of the art TFE-SEM mode (e.g. for process control on resist/etch/metal level with CAD design overlay to compare actual features with design and for automated measurement of critical design dimensions with statistical evaluation; or for overlay work on SEM visible nano particles; or/and for navigation on non-industrial samples (even irregular broken piece parts) missing well-defined marks)
| 3 | Stage and chamber, vibration isolation, sample holders: • 100 x 100 mm (or larger) travel range laser interferometer controlled stage (based on a frequency stabilized HeNe laser source) for xy positioning. • Stage controlled in true closed loop with laser-interferometer: position correction feedback signal applied to motor system. • 1 nm XY position resolution (i.e. 1nm drive commands are possible and executed) at any working distance, write-field size and SEM magnification (actual position is reached without beam tracking). • 30 mm variation in working distance (z) at full interferometric control. • Kinematic - stress free - mount of sample holders through front door and pull out stage or manual load-lock; all holders with high precision (e.g.) sapphire inserts for holders´ optimum seating stability, position repeatability and minimum wear and pre-level function. • One universal sample holder for small and large samples up to 2 inch (50mm) square (including the possibility to clamp irregular shaped (broken) piece parts included. • One 4” wafer holder for 4 inch wafers according to SEMI standards included. • One 4” mask holder for 4 inch masks according to SEMI standards included. • CCD Camera (for chamber view) included. • Integrated 3-fold (passive) anti-vibration platform with both vertical and horizontal vibration isolation mechanism included. • Dry vacuum pump system for oil-free chamber vacuum. • Anti-Contaminator (plasma cleaner) attached to the chamber delivered with an Oxygen Radical Source.
| 4 | Manual load-lock for samples up to 100 x 100 mm: • Load lock with high vacuum gate valve • Short loading cycles • Control electronics • Oil free pumping • Safety interlocks
| 5 | Pattern Generator: • 20 MHz DSP controlled high speed pattern processor unit with dynamic scan corrections. • Dwell time increment of scan generator 1 nsec. • Scan generator with minimum 16bit DACs with additional provision for gain/shift/rotation compensation. • Write field size calibration (and so grating (pitch) control e.g.) with sub-nm resolution.
| 6 | Uninterrupted Power Supply: • Autonomy time: 20 min (100% load) • UPS is fully integrated in operating software and drives down the system in a controlled way in case critical time for main power off has expired.
| 7 | Further hardware: • Electron beam lithography starter kit. • Pico-ampere-meter (to measure the beam current with high resolution). • System water chiller (closed circuit).
| 8 | System must have Fixed Beam Moving Stagewriting capability; the following applies: • Stage follows curved paths at constant speed under laser interferometer control with beam feedback loop • Writes smooth, continuous patterns over large distances with no stitching errors. • Curved path and standard scanning beam pattern definitions integrated in GDSII data base. • Dedicated curved paths editor and conversion function from standard GDSII pattern. • Spot mode and programmable beam shape exposure for flexible linewidth control. • Combined Fixed Beam Moving Stage and vector scan exposures in one run. • Variable stage/beam speed between (minimum range) 1 μm/s and 1mm/s • Positioning resolution of <= 1 nm • High speed beam tracking combined with precise stage motion (e.g. piezo motor moves) • Standard free, intrinsic multiple field size/beam deflection calibration mode. • Fixed Beam Moving Stage patterning control fully integrated in GDSII editor and common multi user software interface. • Automatic dwell time and xy interferometer stage speed parameter determination.
| 9 | System must have Modulated Beam Moving Stage writing capability; the following applies: • Writes continuous strip-shaped periodic patterns over large distances with no stitching errors • Predefined periodic patterns database • Scanning beam deflection pattern calculator for user-specific individual periodic patterns • Combined Modulated Beam Moving Stage and vector scan patterning in one run. • Variable stage/beam speed between (minimum range) 1 μm/s and 1mm/s • Positioning resolution of <= 1 nm • High speed beam tracking combined with precise stage motion (e.g. piezo motor moves) • Standard free, intrinsic multiple field size/beam deflection calibration mode. • Modulated Beam Moving Stage patterning control fully integrated in GDSII editor and common multi user software interface. • Automatic dwell time and xy interferometer stage speed parameter determination.
| 10 | Software: • System control software with true multi-user management system (user login/authentication in different levels and corresponding user-specific individual system parameter and file setup -> no user “interference”). • SEM-like sample inspection, analysis and image archiving. • Dimensional metrology software with filters (noise, threshold,…) and algorithms for data processing and evaluation (distance, edge, pitch determination, cross correlation) and highest precise long range distance measurements using laser-interferometer controlled stage. • Integrated professional GDSII editor with large file handling capabilities. • Special patterning modes: including line by line, meander, true circular mode, raster, free angle, parallel axis, parallel element and sub-modes: including meander, inwards, outwards, transversal, longitudinal. Assignable as attributes to a single individual shape/design GDSII element. • GDSII-Patterning on image (POI) functionality (GDSII design overlay on imported BMP/TIFF image or SEM-image for mix and match applications). • Integrated proximity effect correction, postprocessor, resist development simulator and 3D graphic display, Monte Carlo simulation based parameter determination. • Recipe Manager database: User administrated database with freely configurable exposure parameter and strategy sets assignable to single (or a group of) individual GDSII shapes/elements. • Column Control database: User administrated database for most relevant column parameters with freely configurable parameter sets, can be implemented in automated exposure sequence, column parameter change within seconds. • Multi-Process-Automation: Job Sequencer (position-/task list). • GUIs for Macros and Script Control. • High speed PC with Windows (64bit) software environment. • Protocol, Logger, and Debug tool.
| 11 | Lithography specifications (all exposure results to be evaluated (after exposure and development of resist without further processing like metallisation) using the metrology functionality of the tool itself): • Minimum grating periodicity: ≤ 40 nm • Minimum feature size: ≤ 8 nm • Field stitching: (¦mean¦ + 3 σ) ≤ 40 nm • Overlay accuracy, alignment: (¦mean¦ + 3 σ) ≤ 40 nm • Vendor must provide a detailed description of measurement procedures of above specifications (minimum grating periodicity, minimum feature size, field stitching and overlay accuracy and alignment) • Fixed Beam Moving Stage evaluation: A (minimum) 1 mm long and straight line of about 100 nm width to be exposed (e.g. PMMA950K resist on a silicon substrate) in fixed beam moving stage mode. After exposure and development the measurement and evaluation of the exposure result is conducted by placing line scans on a 10 micron pitch along the length of the path. The scans measure the centre position of the path as well as its line width. The straightness has to be less than or equal to 20 nm. The line width uniformity has to be less than or equal to 10 nm. • Modulated Beam Moving Stage evaluation: Two (minimum) 2mm long gratings with a pitch of 200 nm and 200.5 nm respectively are exposed (e.g. PMMA950K resist on a silicon substrate) using modulated beam moving stage mode (grating lines written perpendicular to stage movement direction). Measurement and statistical evaluation demonstrates for pitch control (accuracy and uniformity): - Grating 1: 200 nm ± 0.2 nm - Grating 2: 200.5 nm ± 0.2 nm
| 12 | The tenderer (supplier) should offer: Active anti vibration platform for a solid stiff floor for the offered tool, including installation and setup. Vibration cancellation starting below 1 Hz Full 6 Degree-of-Freedom active isolation design Active Bandwidth: 0.6 to 100 Hz Isolation above 2 Hz: >80%
| 13 | On Site Training: • Five days training course on site for operators (after completed installation).
| 14 | Shipment Package includes: • Packing material • Shipment cost • Transport-Insurance • According to INCOTERMS 2010 DAP (loading bay customer)
| 15 | Warranty/Service: • One year full warranty including all necessary repair and maintenance labor on site, including all spares and all consumables and all engineer travel and shipping costs. • One year of full service including preventative service, all necessary repair and maintenance labor on site, including all spares, consumables and all engineer travel and shipping costs. (Service period to start after expiry of standard warranty (of 12 months)).
| 16 | Delivery time: Within 9 months
| 17 | Remote (phone, email or direct Internet system access) response time on failure notification within 24 hours after failure notification (on workdays, i.e. Monday to Friday, except bank holidays)
| 18 | Maximum period of time for repair (recovery of system) after notification of system down status – 8 weeks