## 4.ICP-CVD
Bidder: |  | Offered price:
LOT04 - Inductively Coupled Plasma-Chemical Vapor Deposition
ID | Equipment name | QTY | Model, Manufacturer and Country of Origin | Unit price | Total Price per line item
Inductively Coupled Plasma-Chemical Vapor Deposition
4.1 | ICP-CVD system | 1 |  |  | 0
|  |  | Total Price per Lot |  | 0

## 4.1 ICP-CVD system
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
Inductively Coupled Plasma-Chemical Vapor Deposition
| 4.1 | ICP-CVD system |  | 1 |  | 0
| 1 | Si-based dielectrics (SiO2, SiN, SiON, etc.) & a-SiH The hardware should be compatible with deposition with doped P- or N-doped Si requiring additional gases like PH3, AsH3, B2H6
| 2 | System assembled in clean-room and clean-room compatible during its operation
| 3 | Single enclosure rack for all system components (except chiller and low vacuum pump)
| 4 | The tool chamber, plasma source, substrate holders, etc. should be compatible with plasma etching processes using Fluorinated and Chlorinated chemistries. Compatible coatings should be applied.
| 5 | The tool should feature engagement for wide substrate temperature range (at least -30…+150°C).
| 6 | The tool need to allow easy adaptation to various substrate dimensions and shapes within 150 mm circumference, preferably without breaking vacuum in the deposition chamber.
| 7 | Sample holder should allow He backside pressure under the wafer for maintaing the temperature during deposition or etching.
| 8 | The tool should be equipped with an interferometer with an integrated camera for insitu deposition rate control. Camera magnification not less then x50
| 9 | RF power force for ICP plasma generation should feature 2 MHz frequency RF power for RF Bias is 13,56MHz.
| 10 | ICP power 1 KW as standard, 2kw as option
| 11 | Deposition uniformity over 15 cm substrate below 3%.
| 12 | Dry pump for low vacuum.
| 13 | Turbomolecular magnetic levitation pump for high vacuum with capacity at least 1500 l/s
| 14 | Gas pod with up to 11 gas lines
| 15 | All gas lines purgeable and bypassed
| 16 | Manual load lock substrate transport system by prevaccumed chamber to avoid frequent opening of system main reactor chamber.
| 17 | Plasma working pressure up to 50 mT
| 18 | Substrate temperature control between -30 and +150 °C.
| 19 | Heated reactor walls
| 20 | Quartz liner for protection of inner chamber walls from contamination as an option
| 21 | Electropneumatic shut-off valves for gas flow management
| 22 | Appropriate mass flow controllers for required gasses
| 23 | Remotely movable (up and down) cathode electrode in reactor
| 24 | Software for tool operation and complete process control
| 25 | Availability of spare parts for minimum 2 years.
| 26 | Response time for failure notification – 48 hours
| 27 | At least 1 year warranty
| 28 | Maximum period for repair and replacement from the moment of declaring - 31 days
| 29 | Delivery time 6 months
| 30 | Warranty: two years for the hardware, from installation and test, including all parts and their performances.