## 15. Laboratory Equipment with V
Bidder: |  | Offered price:
LOTXX - Laboratory Equipment with Ventilation Units
ID | Equipment name | QTY | Model, Manufacturer and Country of Origin | Unit price | Total Price per line item
Laboratory Equipment with Ventilation units
15.1 | Fume Hood | 7 |  |  | 0
15.2 | Glove box - Class III Biological Safety Cabinet | 1 |  |  | 0
15.3 | Chemical storage cabinet | 3 |  |  | 0
15.4 | Fume cupboard 2100 mm for Clean room zone with main fan | 1 |  |  | 0
15.5 | Special Fume cupboard 1800 mm with main fan | 1 |  |  | 0
15.6 | Special Fume cupboard 1500 mm with main fan | 1 |  |  | 0
15.7 | Fume cupboard 1500 mm for Clean room zone with main fan | 1 |  |  | 0
|  |  | Total Price per Lot |  | 0

## 15.1 Fume Hood
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
Laboratory Equipment with Ventilation Units
| 15.1 | Fume Hood |  | 7 |  | 0
| 1.1 | Dimensions up to 1500x900x2300 mm
| 2 | Support structure should be made of steel profiles and sheet metal, with leveling feet in epoxy protection.
| 3 | The inner lining of the working space made of a material resistant to extreme temperatures from -40 °C up to +140 °C, and in the time intervals of 20 minute even up to +180 °C
| 4 | Chemical resistance is necessary
| 5 | Work surface area should be made of monolithic acid-resistant ceramics with thickness up to 33 mm. It should have mechanical resistance of 6 on Mohs scale, as well as chemical and temperature resistance. At the back of the surface area there should be ceramic tub dimensions up to 250x100x150, with PP siphon and one faucet for cold water.
| 6 | Fume hood window should be made out of tempered glass, placed on the steel frame.
| 7 | At the back of the fume hood there should be one compressed air faucet.
| 8 | LED with IP55 protection and power 2x8W
| 9 | At the front of the fume hood there should be three 230 V connectors, and switches for the light and ventilation, as well as for the water faucet.
| 10 | On top panel there should be electric cabinet with motor protection for the ventilator, three-pole switch 25A, light signalization
| 11 | Drawer below the fume hood should be included.
| 12 | Acid-resistant ventilator in Ex version, capacity 1000 m3/h, power 0,25 kW and supply voltage 400 V, ATEX 95 Ec II 3 c IIB T3. The ventilation system must ensure during operation, after the installation of the digester in the user's laboratory, that the amount of air expelled from the digester is - min.750m3 / h; that the speed of air flow is in the plane of the opening of the working space of the digester - min. 0.30 m / s; that there is an air flow towards the digester opening. As a confirmation, submit the Report issued by the laboratory accredited for testing according to EN 14175, on testing of at least one digester, performed after digester installation, in the customer's laboratory, from which it can be determined: / h; that the speed of air flow in the plane of the opening of the working space of the digester is at least 0.30 m / s; that the air flow towards the digester opening has been determined. The motor of the offered fan must comply with the requirements of the Ordinance on equipment and protective systems, intended for use in potentially explosive atmospheres, as well as the requirements of the standard: SRPS EN 60079-0; SRPS EN 60079-7; with established Ex protection II2G; ExeIIT3; and 24 s Proof: A certificate proving that the characteristics of the offered fan motor meet the requirements of the Ordinance on equipment and protective systems, intended for use in potentially explosive atmospheres, as well as the requirements of the standard: SRPS EN 60079-0; SRPS EN 60079-7; and that Ex protection II2G; ExeIIT3 A certificate issued by a designated RS body will be considered acceptable The offered fan must comply with the Electromagnetic Compatibility Ordinance Proof: Certificate of compliance with the Ordinance on electromagnetic compatibility for offered fan of the offered manufacturer A valid certificate of conformity issued by a designated RS body will be considered acceptable
| 13 | Availability of the spare parts for 7 years.
| 14 | Delivery time: Within 3 months
| 15 | Response time for failure notification – 72 hours
| 16 | Maximum period for repair and replacement from the moment of declaring 6 weeks
| 17 | Minimum warranty 12 months
| 18 | Manufacturer's catalog or certificate of workspace materials and ceramics

## 15.2 Glove box - Class III
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
Laboratory Equipment with Ventilation Units
| 15.2 | Glove box - Class III Biological Safety Cabinet |  | 1 |  | 0
| 1.1 | Class III Biological Safety Cabinet totally enclosed and gas-tight with ULPA filtered supply and exhaust air. Work is performed with long-sleeved gloves. The cabinet should be kept under negative pressure of at least 120 Pa, and airflow should be maintained by a dedicated exterior exhaust system in order to protect the operator, product and environment.
| 2 | It should contain the pass box with double doors that have interlock function and UV lamp.
| 3 | UV lamp of 253.7 nm for sterilization decontamination.
| 4 | LCD display
| 5 | Dimensions of external size up to 1550x900x2000 mm; and internal size up to 920x660x670
| 6 | Exhaust air must be double-filtered through high-quality HEPA filters with typical efficiency of 99.999% for 0.3 um particles.
| 7 | Pressure in the cabinet should be ≥ negative 120 Pa
| 8 | Work zone material should be 304 stainless steel
| 9 | Illumination ≥ 1000 Lux
| 10 | Power supply AC220 V ±10%, 50/60 Hz; 110 V ±10%, 60 Hz
| 11 | Availability of the spare parts for 7 years.
| 12 | Delivery time: Within 3 months
| 13 | Response time for failure notification – 72 hours
| 14 | Maximum period for repair and replacement from the moment of declaring 6 weeks
| 15 | Minimum warranty 12 months

## 15.3 Chemical storage cabinet
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
Laboratory Equipment with Ventilation Units
| 15.3 | Chemical storage cabinet |  | 3 |  | 0
| 1.1 | Cabinet class/type 90, EN 14470-1
| 2 | Dimensions up to 900x650x1980 mm
| 3 | 2 swing doors
| 4 | Three-layer sandwich construction; from outside to inside: steel with epoxy protection, fire-resistant layer, panel for shelf mounting
| 5 | Ventilation exhaust Ø75, quickly put into operation
| 6 | Locking should have protection against unauthorized use; lockable doors with cylinder lock
| 7 | Available interior equipment: shelves should have load capacity of 75 kg and bottom collecting sumps should have volume of 23 l
| 8 | It should include ventilation system for full capacity utilization of Safety Storage Cabinet. Ventilator capacity should be between 60-250 m3/h, with ventilation connector Ø75, and power 0,09 kW. Supply voltage 230 V.
| 9 | The cabinet must satisfy: Standard EN 14470-1; Directive 94/9 / EC
| 10 | with the offer: -submit an excerpt from the manufacturer's catalog for the offered cabinets as proof of compliance with the required technical characteristics and the required standard (EN 14470-1); - CE declaration as proof of compliance with the required Directive 94/9 / EC.
| 11 | Availability of the spare parts for 7 years.
| 12 | Delivery time: Within 3 months
| 13 | Response time for failure notification – 72 hours
| 14 | Maximum period for repair and replacement from the moment of declaring 6 weeks
| 15 | Minimum warranty 12 months

## 15.4 Fume cupboard 2100 mm
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
Laboratory Equipment with Ventilation Units
| 15.4 | Fume cupboard 2100 mm for Clean room zone with main fan |  | 1 |  | 0
| 1 | The fume cupboard must comply with the requirements of EN 14175. The Bidder must submit with the Bid a Certificate EN 14175 issued by a certified EU institution and a copy of the test protocol as well as the Declaration of Conformity for standard 61010-1 2001.
| 2 | width [mm] 2100 depth [mm] 900 working height [mm] 900 height [mm] 2700
| 3 | top part 1 x Bench-mounted fume cupboard - upper part 2100 mm wide Internal lining: melamine Glass pane: tempered safety glass Side wall left and right: Internal lining melamine
| 4 | The rear wall of the internal work area should be constructed as a "double wall" with built-in channels for heavy fumes that may eventually be collected at the corners of the worktop. The part of the double wall facing the inner work area should be made of Trespa, 5 mm thick, and the part of the wall that forms the structure of the fume cupboard should be made of hard-pressed chipboard coated with melamine resin coated.
| 5 | The front window sash doors should be designed in such a way that they do not, when fully raised, exceed the maximum height of the fume cupboard 2700 mm. The lower part of the window-door should have horizontal sliding windows.
| 6 | The front door should be connected to the counter via a toothed belt made of plastic materials laced with steel threads. The front window-door frame shall be fitted with security systems that lock the door in the event of a belt break.
| 7 | bottom part 1 x supporting construction for fume cupboard - 2100mm wide as H- frame. Turned frame to have access to the underbench cabinets through the back of the fume cupboard. with service panels 1 x spine for fume cupboard - 2100mm wide 1 x 1,89 m2 panel, 19 mm particle board (cover of front side of bottom part)
| 8 | The supporting metal frames of the tables should be connected by 4 legs with cross bars. The supporting frames should be made of steel profiles (60/25/2 mm), with a load capacity of 200 kg / m². All welds should be polished, smooth, with no bumps or openings. The surface should be coated with epoxy-coated powder, the thickness of the layer should be about 80 µm. The feet should have height adjustable feet
| 9 | Interior lighting should be a minimum of 400 lux. The lamp should be energy saver type. The lamp base should be designed in such a way that it is at the same time a safety pressure relief for the shock wave in the event of an explosion in the work area.
| 10 | Functional monitoring and Air flow control of fume cupboards in accordance with EN 14175 Monitoring air outlet volume based on measurement of pressure difference. Calculation of the amount of exhaust air based on the position of the front window. The green LED on the display should illuminate during normal operation. The red LED on the display should indicate a malfunction and should be accompanied by an audible signal. The visual alarm signal disappears on its own when the fault no longer exists. Raising the front door-window above the safe zone (500 mm) must be accompanied by the activation of an alarm and a red LED on the display. The control panel with visual and audible warning signals should be integrated into the aluminum profile - the edge of the fume cupboard at eye level.
| 11 | extraction manifold the manifold should be made of polypropylene with 2 inlets 200 mm and one outlet 250 mm. The VAV system damper should be connected to the manifold and close pipeline when fume cupboard is off.
| 12 | Underbench solvents disposal unit - access from the rear side of fume cupboard 1 x fume cupboard underbench unit - FWF 90, for 2 container changing system width 600 mm depth 595 mm height 635 mm For disposal of combustible hazardous substances in working rooms according to EN 14 470-1 * 1 hinged door * fastening device with thermo-control * ventilation connections in rear panel * bushing for waste disposal line in the rear panel * interior equipment: 1 floor tray - powder coated sheet steel for disposal of solvents (2 funnel changing system)
| 13 | 1 x electric panel (single) to control 1 waste disposal container, in explosion-proof execution for flammable liquids, with electric level indicator: visual + acoustic
| 14 | 1 x funnel tube with spring cover (worktop installation), with removable sieve for disposal of flammable liquids of electrically conductive polyethylene
| 15 | 2 x PE-HD container 5 l, with PTZ+UN-approval for chemical waste
| 16 | worktop 1 x stoneware worktop 2050 x 708 x 26/33 mm with raised edge on all sides 1 x sink of stoneware 400 X 400 X 250 mm
| 17 | mechanical services 1 x outlet nozzle 90° (Broen) nitrogen - N2 1 x outlet, fume cupboard valve (Broen) nitrogen - N2 1 x outlet nozzle 90° (Broen) vacuum - V 1 x outlet, fume cupboard valve (Broen) vacuum - V 2 x bench connection piece – supply spray points (for „spraying guns“: https://cleanroomworld.com/cleanroom-equipment/nitrogen-spray-guns/90-degree-PTFE-FEP-quarter-inch ; https://cleanroomworld.com/cleanroom-equipment/di-water-spray-guns/PTFE-blow-off-gun-with-8-feet-of-PTFE-tubing-and-quarter-inch-inlet-thread 2 x Spray guns (as described above) 1 x outlet, fume cupboard valve (Broen) demineralized water – WDC 1 x outlet, fume cupboard valve (Broen) nitrogen - N2 1 x outlet pillar (Broen) projection 200 mm drinking water, cold - WPC 1 x outlet, fume cupboard valve (Broen) drinking water, cold - WPC
| 18 | electricity 1 x panel (equipped) 4 x socket 230 V, 16 A 1 x panel (equipped) 4 x socket 230 V, 16 A (inside of workzone)
| 19 | Main ventilator: plastic, acid proof, centrifugal fan. Minimum capacity: 1300 m3/h
| 20 | Training on the operation
| 21 | The system is supplied with operating manual, interconnect schematic, labels for plugs and any necessary instruction for its operation
| 22 | Availability of the spare parts for 8 years.
| 23 | Warranty: two years, from installation and test
| 24 | Delivery time: Within 4 months
| 25 | Response time for failure notification – 48 (forty-eight) hours
| 26 | Maximum period for repair and replacement from the moment of declaring - 30 (thirty) days

## 15.5 Special Fume cupboard1800
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
Laboratory Equipment with Ventilation Units
| 15.5 | Special Fume cupboard 1800 mm with main fan |  | 1 |  | 0
| 1 | The fume cupboard must comply with the requirements of EN 14175. The Bidder must submit with the Bid a Certificate EN 14175 issued by a certified EU institution and a copy of the test protocol as well as the Declaration of Conformity for standard 61010-1 2001.
| 2 | width [mm] 1800 depth [mm] 900 working height [mm] 900 height [mm] 2700
| 3 | top part 1 x Bench-mounted fume cupboard - upper part 1800 mm wide Internal lining: polypropylene Glass pane: Makrolon Side wall left and right: polypropylene
| 4 | The rear wall of the internal work area should be constructed as a "double wall" with built-in channels for heavy fumes that may eventually be collected at the corners of the worktop. The part of the double wall facing the inner work area should be made of polypropylene.
| 5 | The front window sash doors should be designed in such a way that they do not, when fully raised, exceed the maximum height of the fume cupboard 2700 mm. The lower part of the window-door should have horizontal sliding windows.
| 6 | The front door should be connected to the counter via a toothed belt made of plastic materials laced with steel threads. The front window-door frame shall be fitted with security systems that lock the door in the event of a belt break.
| 7 | bottom part 1 x supporting construction for fume cupboard - 1800 mm wide as H- frame. Turned frame to have access to the underbench cabinets through the back of the fume cupboard. with service panels 1 x spine for fume cupboard - 1800 mm wide 1 x 1,62 m2 panel, Solid grade laminate (cover of front side of bottom part)
| 8 | The supporting metal frames of the tables should be connected by 4 legs with cross bars. The supporting frames should be made of steel profiles (60/25/2 mm), with a load capacity of 200 kg / m². All welds should be polished, smooth, with no bumps or openings. The surface should be coated with epoxy-coated powder, the thickness of the layer should be about 80 µm. The feet should have height adjustable feet
| 9 | Interior lighting should be a minimum of 400 lux. The lamp should be energy saver type. The lamp base should be designed in such a way that it is at the same time a safety pressure relief for the shock wave in the event of an explosion in the work area.
| 10 | Functional monitoring and Air flow control of fume cupboards in accordance with EN 14175 Monitoring air outlet volume based on measurement of pressure difference. Calculation of the amount of exhaust air based on the position of the front window. The green LED on the display should illuminate during normal operation. The red LED on the display should indicate a malfunction and should be accompanied by an audible signal. The visual alarm signal disappears on its own when the fault no longer exists. Raising the front door-window above the safe zone (500 mm) must be accompanied by the activation of an alarm and a red LED on the display. The control panel with visual and audible warning signals should be integrated into the aluminum profile - the edge of the fume cupboard at eye level.
| 11 | extraction manifold the manifold should be made of polypropylene with 2 inlets 200 mm and one outlet 250 mm. The VAV system damper should be connected to the manifold and close pipeline when fume cupboard is off.
| 12 | Underbench solvents disposal unit - access from the rear side of fume cupboard 2 x fume cupboard underbench unit - FWF 90, for 2 container changing system width 600 mm depth 595 mm height 635 mm For disposal of combustible hazardous substances in working rooms according to EN 14 470-1 * 1 hinged door * fastening device with thermo-control * ventilation connections in rear panel * bushing for waste disposal line in the rear panel * interior equipment: 1 floor tray - powder coated sheet steel for disposal of solvents (2 funnel changing system)
| 13 | 1 x electric panel (double) to control 1 waste disposal container, in explosion-proof execution for flammable liquids, with electric level indicator: visual + acoustic
| 14 | 2 x funnel tube with spring cover (worktop installation), with removable sieve for disposal of flammable liquids of electrically conductive polyethylene
| 15 | 2 x PE-HD container 5 l, with PTZ+UN-approval for chemical waste
| 16 | worktop 1 x polypropylene worktop 1750 x 708 x 26/33 mm with raised edge on all sides 1 x sink of polypropylene 400 X 400 X 250 mm
| 17 | mechanical services 1 x outlet, fume cupboard valve (Broen) vacuum - V 1 x outlet nozzle 90°, coated with Accofal (Broen) vacuum - V 1 x connection piece – supply spray connections (for „spraying guns“: https://cleanroomworld.com/cleanroom-equipment/nitrogen-spray-guns/90-degree-PTFE-FEP-quarter-inch ; https://cleanroomworld.com/cleanroom-equipment/di-water-spray-guns/PTFE-blow-off-gun-with-8-feet-of-PTFE-tubing-and-quarter-inch-inlet-thread 2 x Spray guns (as described above) 1 x outlet, fume cupboard valve (Broen) nitrogen - N2 1 x outlet, fume cupboard valve (Broen) nitrogen - N2 1 x outlet, fume cupboard valve (Broen) drinking water, cold - WPC 1 x outlet pillar (Broen) projection 200 mm drinking water, cold - WPC 1 x Accofal coating for cold water tap on stand
| 18 | electricity 1 x panel (equipped) 4 x socket 230 V, 16 A 1 x panel (equipped) 4 x socket 230 V, 16 A (inside of workzone)
| 19 | Fume cupboards No: 2 and No:3 will stand side-by-side and should have open space in the wall between them. Open space should have dimension width 300 mm x height 400 mm, placed above worktop zone.
| 20 | Main ventilator: plastic, acid proof, centrifugal fan. Minimum capacity: 1300 m3/h
| 21 | Training on the operation
| 22 | The system is supplied with operating manual, interconnect schematic, labels for plugs and any necessary instruction for its operation
| 23 | Availability of the spare parts for 8 years.
| 24 | Warranty: two years, from installation and test
| 25 | Delivery time: Within 4 months
| 26 | Response time for failure notification – 48 (forty-eight) hours
| 27 | Maximum period for repair and replacement from the moment of declaring - 30 (thirty) days

## 15.6 Special Fume cupboard 1500
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
Laboratory Equipment with Ventilation Units
| 15.6 | Special Fume cupboard 1500 mm with main fan |  | 1 |  | 0
| 1 | The fume cupboard must comply with the requirements of EN 14175. The Bidder must submit with the Bid a Certificate EN 14175 issued by a certified EU institution and a copy of the test protocol as well as the Declaration of Conformity for standard 61010-1 2001.
| 2 | width [mm] 1500 depth [mm] 900 working height [mm] 900 height [mm] 2700
| 3 | top part 1 x Bench-mounted fume cupboard - upper part 1500 mm wide Internal lining: polypropylene Glass pane: Makrolon Side wall left and right: polypropylene
| 4 | The rear wall of the internal work area should be constructed as a "double wall" with built-in channels for heavy fumes that may eventually be collected at the corners of the worktop. The part of the double wall facing the inner work area should be made of polypropylene.
| 5 | The front window sash doors should be designed in such a way that they do not, when fully raised, exceed the maximum height of the fume cupboard 2700 mm. The lower part of the window-door should have horizontal sliding windows.
| 6 | The front door should be connected to the counter via a toothed belt made of plastic materials laced with steel threads. The front window-door frame shall be fitted with security systems that lock the door in the event of a belt break.
| 7 | bottom part 1 x supporting construction for fume cupboard - 1500 mm wide as H- frame. Turned frame to have access to the underbench cabinets through the back of the fume cupboard. with service panels 1 x spine for fume cupboard - 1500 mm wide 1 x 1,35 m2 panel, Solid grade laminate (cover of front side of bottom part)
| 8 | The supporting metal frames of the tables should be connected by 4 legs with cross bars. The supporting frames should be made of steel profiles (60/25/2 mm), with a load capacity of 200 kg / m². All welds should be polished, smooth, with no bumps or openings. The surface should be coated with epoxy-coated powder, the thickness of the layer should be about 80 µm. The feet should have height adjustable feet
| 9 | Interior lighting should be a minimum of 400 lux. The lamp should be energy saver type. The lamp base should be designed in such a way that it is at the same time a safety pressure relief for the shock wave in the event of an explosion in the work area.
| 10 | Functional monitoring and Air flow control of fume cupboards in accordance with EN 14175 Monitoring air outlet volume based on measurement of pressure difference. Calculation of the amount of exhaust air based on the position of the front window. The green LED on the display should illuminate during normal operation. The red LED on the display should indicate a malfunction and should be accompanied by an audible signal. The visual alarm signal disappears on its own when the fault no longer exists. Raising the front door-window above the safe zone (500 mm) must be accompanied by the activation of an alarm and a red LED on the display. The control panel with visual and audible warning signals should be integrated into the aluminum profile - the edge of the fume cupboard at eye level.
| 11 | extraction manifold the manifold should be made of polypropylene with 2 inlets 200 mm and one outlet 250 mm. The VAV system damper should be connected to the manifold and close pipeline when fume cupboard is off.
| 12 | acid/alkali units 1 x fume cupboard underbench unit on plinth slide-in - 600 mm wide handle element, aluminium for acids and alkalines 1 hinged door, 2 extendable shelves with polypropylene tray Depth 550mm Extraction - underbench exhaust 1 x fume cupboard underbench unit on plinth slide-in - 845 mm wide handle element, aluminium for acids and alkalines 2 hinged doors, 2 extendable shelves with polypropylene tray Depth 550mm Extraction - underbench exhaust
| 13 | worktop 1 x polypropylene worktop 1450 x 708 x 26/33 mm with raised edge on all sides
| 14 | mechanical services 1 x outlet, fume cupboard valve (Broen) vacuum - V 1 x outlet nozzle 90°, coated with Accofal (Broen) vacuum - V 1 x outlet, fume cupboard valve (Broen) nitrogen - N2
| 15 | electricity 1 x panel (equipped) 4 x socket 230 V, 16 A 1 x panel (equipped) 4 x socket 230 V, 16 A (inside of workzone)
| 16 | Fume cupboards No: 2 and No:3 will stand side-by-side and should have open space in the wall between them. Open space should have dimension width 300 mm x height 400 mm, placed above worktop zone.
| 17 | Main ventilator: plastic, acid proof, centrifugal fan. Minimum capacity: 1300 m3/h
| 18 | Training on the operation
| 19 | The system is supplied with operating manual, interconnect schematic, labels for plugs and any necessary instruction for its operation
| 20 | Availability of the spare parts for 8 years.
| 21 | Warranty: two years, from installation and test
| 22 | Delivery time: Within 4 months
| 23 | Response time for failure notification – 48 (forty-eight) hours
| 24 | Maximum period for repair and replacement from the moment of declaring - 30 (thirty) days

## 15.7 Fume cupboard 1500 mm CR
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
Laboratory Equipment with Ventilation Units
| 15.7 | Fume cupboard 1500 mm for Clean room zone with main fan |  | 1 |  | 0
| 1 | The fume cupboard must comply with the requirements of EN 14175. The Bidder must submit with the Bid a Certificate EN 14175 issued by a certified EU institution and a copy of the test protocol as well as the Declaration of Conformity for standard 61010-1 2001.
| 2 | width [mm] 1500 depth [mm] 900 working height [mm] 900 height [mm] 2700
| 3 | top part 1 x Bench-mounted fume cupboard - upper part 1500 mm wide Internal lining: melamine Glass pane: tempered safety glass Side wall left and right: Internal lining melamine
| 4 | The rear wall of the internal work area should be constructed as a "double wall" with built-in channels for heavy fumes that may eventually be collected at the corners of the worktop. The part of the double wall facing the inner work area should be made of Trespa, 5 mm thick, and the part of the wall that forms the structure of the fume cupboard should be made of hard-pressed chipboard coated with melamine resin coated.
| 5 | The front window sash doors should be designed in such a way that they do not, when fully raised, exceed the maximum height of the fume cupboard 2700 mm. The lower part of the window-door should have horizontal sliding windows.
| 6 | The front door should be connected to the counter via a toothed belt made of plastic materials laced with steel threads. The front window-door frame shall be fitted with security systems that lock the door in the event of a belt break.
| 7 | bottom part 1 x supporting construction for fume cupboard - 1500mm wide as H- frame. Turned frame to have access to the underbench cabinets through the back of the fume cupboard. with service panels 1 x spine for fume cupboard - 1500mm wide 1 x 1,35 m2 panel, 19 mm particle board (cover of front side of bottom part)
| 8 | The supporting metal frames of the tables should be connected by 4 legs with cross bars. The supporting frames should be made of steel profiles (60/25/2 mm), with a load capacity of 200 kg / m². All welds should be polished, smooth, with no bumps or openings. The surface should be coated with epoxy-coated powder, the thickness of the layer should be about 80 µm. The feet should have height adjustable feet
| 9 | Interior lighting should be a minimum of 400 lux. The lamp should be energy saver type. The lamp base should be designed in such a way that it is at the same time a safety pressure relief for the shock wave in the event of an explosion in the work area.
| 10 | Functional monitoring and Air flow control of fume cupboards in accordance with EN 14175 Monitoring air outlet volume based on measurement of pressure difference. Calculation of the amount of exhaust air based on the position of the front window. The green LED on the display should illuminate during normal operation. The red LED on the display should indicate a malfunction and should be accompanied by an audible signal. The visual alarm signal disappears on its own when the fault no longer exists. Raising the front door-window above the safe zone (500 mm) must be accompanied by the activation of an alarm and a red LED on the display. The control panel with visual and audible warning signals should be integrated into the aluminum profile - the edge of the fume cupboard at eye level.
| 11 | extraction manifold the manifold should be made of polypropylene with 2 inlets 200 mm and one outlet 250 mm. The VAV system damper should be connected to the manifold and close pipeline when fume cupboard is off.
| 12 | Underbench solvents disposal unit - access from the rear side of fume cupboard 1 x fume cupboard underbench unit - FWF 90, for 2 container changing system width 600 mm depth 595 mm height 635 mm For disposal of combustible hazardous substances in working rooms according to EN 14 470-1 * 1 hinged door * fastening device with thermo-control * ventilation connections in rear panel * bushing for waste disposal line in the rear panel * interior equipment: 1 floor tray - powder coated sheet steel for disposal of solvents (2 funnel changing system)
| 13 | 1 x electric panel (single) to control 1 waste disposal container, in explosion-proof execution for flammable liquids, with electric level indicator: visual + acoustic
| 14 | 1 x funnel tube with spring cover (worktop installation), with removable sieve for disposal of flammable liquids of electrically conductive polyethylene
| 15 | 2 x PE-HD container 5 l, with PTZ+UN-approval for chemical waste
| 16 | worktop 1 x stoneware worktop 1450 x 708 x 26/33 mm with raised edge on all sides 1 x sink of stoneware 400 X 400 X 250 mm
| 17 | mechanical services 1 x connection piece – supply spray connections (for „spraying guns“: https://cleanroomworld.com/cleanroom-equipment/nitrogen-spray-guns/90-degree-PTFE-FEP-quarter-inch ; https://cleanroomworld.com/cleanroom-equipment/di-water-spray-guns/PTFE-blow-off-gun-with-8-feet-of-PTFE-tubing-and-quarter-inch-inlet-thread 2 x Spray guns (as described above) 1 x outlet, fume cupboard valve (Broen) demineralized water - WDC 1 x outlet, fume cupboard valve (Broen) nitrogen - N2 water mixer 1 x outlet pillar (Broen) projection 200 mm drinking water, cold - WPC 1 x outlet, fume cupboard valve (Broen) drinking water, cold - WPC
| 18 | electricity 1 x panel (equipped) 4 x socket 230 V, 16 A 1 x panel (equipped) 4 x socket 230 V, 16 A (inside of workzone)
| 19 | Main ventilator: plastic, acid proof, centrifugal fan. Minimum capacity: 1300 m3/h
| 20 | Training on the operation
| 21 | The system is supplied with operating manual, interconnect schematic, labels for plugs and any necessary instruction for its operation
| 22 | Availability of the spare parts for 8 years.
| 23 | Warranty: two years, from installation and test
| 24 | Delivery time: Within 4 months
| 25 | Response time for failure notification – 48 (forty-eight) hours
| 26 | Maximum period for repair and replacement from the moment of declaring - 30 (thirty) days