## 8. Phenotyping Laboratory
Bidder: |  | Offered price:
LOT08 - Phenotyping Laboratory
ID | Equipment name | QTY | Model, Manufacturer and Country of Origin | Unit price | Total Price per line item
Phenotyping Laboratory
8.1 | Phenotyping Laboratory Equipment | 1 |  |  | 0
|  |  | Total Price per Lot |  | 0

## 8.1 Phenotyping Laboratory Equi
Bidder:
Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
Phenotyping Laboratory
| 8.1 | Phenotyping Laboratory Equipment |  | 1 |  | 0
| 1 | Imaging Cabinets with Semi-Automated Plant DeliveryLight-insulated imaging box equipped conveyor roller transportation system and automated doors for convenient placing of samples. Plants are taken automatically for imaging after starting experimental protocol. Includes QR code reader for sample identification. Box construction is made from durable aluminum extrusion frame. Maximum plant height (including pot): 120cm and maximum plant width: 50cm. Number and types of provided trays: 40/tbd |  | 1
| 2 | High-Sensitivity Kinetic Chlorophyll Fluorescence Imaging UnitCustom - designed PAM chlorophyll fluorescence measuring system. A system with fluorescence cameras to image 35 cm x 35 cm area with 1.4MP resolution. Supplied with multiple light sources: • Red-orange 620 nm LED-based light for pulse-modulated short duration flashes for accurate measurement of minimal fluorescence (Fo value) and additional actinic illumination • Cool-white 5700K LED-based light providing saturating pulse for maximal fluorescence (Fm value) determination and standard actinic illumination • Far-Red 735 nm additional LED for determination of Fo'. Measured parameters: Fo, Fm, Fv, Fo', Fm', Fv', Ft. Calculated parameters: Fv/Fm, Fv'/Fm', PhiPSII , NPQ, qN, qP, Rfd and many others used for monitoring of abiotic and biotic stresses that affect photosynthetic mechanisms (e.g. drought, heat, and light stresses, nutrient deficiencies, heavy metal toxicities, fungal infections etc.). Degree of stress, and recovery from stress, are quantified over time. |  | 1
| 3 | RGB Structural Imaging Unit RGB Imaging station with 2 cameras using next generation CMOS chip with 12.3MP resolution. One camera mounted above plants for top view and one mounted on side. Includes suitable LED light source for each camera, robotics for height adjustment and turning platform for precise image acquisition from different angles. System provides measurements of morphometric parameters including: Area; Perimeter; Eccentricity; Roundness; Plant Height; Plant Width; Slenderness of leaves; Solidity/Compactness and many others. The RGB Imaging Station also enables: • Color segmentation for plant fitness evaluation • Sequential computed relative leaf growth rates • Comparison of leaf area differences during whole experiment • Greening index • Contribution of individual colors after color segmentation and characterization |  | 1
| 4 | 3D Laser Scanning Unit - laser triangulation scanner used in the systems is designed for precise structural plant phenotyping. Scanner is mounted on vertical linear robotic guide for smooth and precise point cloud acquisition and rotating stage to generate accurate 3D model with mm resolution in all directions. Point cloud is automatically triangulated, followed by automatic segmentation on individual leaves and stem. Possibility to coregister with RGB texture or parametric images from other imaging units. Automatic analysis calculates important morphological parameters including: • Digital biomass • Leaf/Plant area • Projected area • Leaf angle • Compactness • Plant/Stem height • Number of leaves |  | 1
| 5 | High-Resolution Thermal Imaging Unit - Scientific-grade LWIR imaging unit with calibrated thermal camera with uncooled microbolometer detector. Includes PT1000 temperature reference sensors, programmable LED light source suitable for temperature dynamics studies synchronized with user-defined imaging sequence via experimental protocol (i.e. stomatal conductance research). • Resolution: 1024x768 pixels • Measurement accuracy: ±1°C or ±1% • Dynamic range: 16 bit • Temperature sensitivity (NETD @30°C): 20mK |  | 1
| 6 | VNIR Hyperspectral Imaging Unit Hyperspectral system with ultra-high spectral resolution designed for detection and quantification of pigments, metabolites and calculation of vegetation indices. Includes automated radiometric calibration, suitable light source covering whole spectral range and robotics for image acquisition. • Spectral range: 380-900nm • Spectral resolution: 0.8nm (FWHM) • Spatial bands: 1000 • Spectral bands: 1920 Automatic calculation of vegetation indices (NDVI, PRI, MCARI….) and possibility to define new indices by user. |  | 1
| 7 | SWIR Hyperspectral Imaging Unit Hyperspectral system with ultra-high spectral resolution designed for quantification of water content and other compounds (cellulose, lignin) in plants. Includes automated radiometric calibration, suitable light source covering whole spectral range • Spectral range: 1000-2200nm • Spectral resolution: 4nm (FWHM) • Spatial bands: 256 • Spectral bands: 320 Automatic calculation of water content and possibility to define new indices by user. |  | 1
| 8 | Server service - moderates connections among all system components, provides interface to the main database, executes experimental setups by client application, runs GSM and mail service and log service. This service is running all the time the system is powered on. Client application - connects to the server, displays actual system state, provides GUI for experiment setup and control. A scheduler for client side, designed for visualization, planning and management to the individual experiments for multiple registered users. It is used to access and manage current and planned experiments and design experimental protocols. There can be multiple clients connected to the server, with different privileges assigned based on a built-in authentication mechanism. |  | 1
| 9 | Data analyzer application for "Big Data" handling - provides tools for data browsing and analysis. Multiple clients can be connected to the database, with different privileges assigned based on a built-in authentication mechanism. |  | 1
| 10 | Mini Application server with DB and UPS |  | 1
| 11 | Programmable logic controller unit with sensors |  | 1
| 12 | Installation and training |  | 1
| 13 | Availability of the spare parts for minimum 10 years. Consumables on stock. |  | 1
| 14 | Warranty: minimum 24 Months
| 15 | Delivery time: Within 6 months
| 16 | Response time for failure notification – response time 1 working day, resolution of an issue that can be done remotely within 3 working days, onsite intervention within 7 days or after mutual agreement.
| 17 | Maximum period for repair and replacement from the moment of declaring - 24 days