## Price Schedule
Bidder: |  | Offered Price:
LOT 1 | MEDICAL FURNITURE AND EQUIPMENT
ID | Equipment name | QTY | Technical Specification Offered(model) | Unit price | Total price per item
1 | Intensive care hospital bed | 34 |  |  | 0
2 | Hospital bed | 48 |  |  | 0
3 | Bedside cabinet without table (prepared for table) | 94 |  |  | 0
4 | Dry sterilizer | 3 |  |  | 0
5 | Two step foot stool | 51 |  |  | 0
6 | IV stand | 90 |  |  | 0
7 | Cabinet for drugs | 19 |  |  | 0
8 | Mobile overbed table | 78 |  |  | 0
9 | Examination table with electrical hight adjustment | 28 |  |  | 0
10 | Three section screen | 36 |  |  | 0
11 | Trolley for therapy | 6 |  |  | 0
12 | Trolley for ECG | 12 |  |  | 0
13 | Strecher for patient transport | 10 |  |  | 0
14 | Wheelchair for transportation | 24 |  |  | 0
15 | Daily Care Aid-Toilet type Princess | 8 |  |  | 0
16 | INOX trolley | 8 |  |  | 0
17 | Blood sample chair electrical | 9 |  |  | 0
18 | X-ray film viewer | 1 |  |  | 0
19 | O2 Therapy | 100 |  |  | 0
20 | Electrical suction device | 18 |  |  | 0
21 | Double stainless steel (clean - not clean) bowls | 20 |  |  | 0
22 | Tray for instruments without lid | 20 |  |  | 0
23 | Strecher for patient care | 4 |  |  | 0
24 | Trolley for ultrasound | 10 |  |  | 0
25 | Kidney shape bowl (large) | 20 |  |  | 0
26 | Kidney shape bowl (small) | 20 |  |  | 0
27 | Test tube holder (bigger) | 20 |  |  | 0
28 | Test tube holder (smaller) | 20 |  |  | 0
29 | Operating table for General surgery and laparascopy | 1 |  |  | 0
30 | Examination and operation light, mobile version | 2 |  |  | 0
|  |  | Total Price |  | 0

## TechnicalSpec.
BIDDER:
LOT 1 |  | MEDICAL FURNITURE AND EQUIPMENT |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical specification requested | Technical specification | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Intensive care hospital bed |  | 34 |  | 0
| 1 | Total bed dimensions: 1000mm±20mm x 2150mm ±50mm
| 2 | Safe working load (SWL) min.250kg
| 3 | Bed constructed by robust and durable materials
| 4 | Quadruple bed lying plate, Polypropylene segments , beckrest X ray permeable
| 5 | 4 adaptive holes for patient lifting pole or infusion stand
| 6 | Bed frame is made of steel profiles painted with powder - colored polyester resistent to desinfectant
| 7 | Integrated bed extension minimum 280 mm
| 8 | Double castors Ø150mm, with central locking system. One wheel conductive
| 9 | Bed with at least four two-way electric motors for starting: height, inclination of the back, inclination of the slope of the thigh, Trendelenburg and reverse Trendelenburg
| 10 | Regulation of height with electric motor in min. range: 400mm - 800mm
| 11 | Electrical automatic Tr / ATr ≥ -16 °/ +16°
| 12 | Beck rest angle ≥ 70 °.
| 13 | Thigh rest angle ≥ 40 °
| 14 | Automatic backrest lenght compensation min 90mm
| 15 | Automatic exemination (flat and high) positioning of the bed, examination position
| 16 | Nursing control panel with the following bed functions: up / down, shock position, back and thigh angle, cardiac chair position, CPR, Trendelenburg / Reverse Trendelenburg, blocking individual or all functions
| 17 | Easily visible and approachable manual CPR commands
| 18 | Movable headboard and footboard with security lock
| 19 | One piece safety sides,covering at leas 3/4 of mattress base with locking system in the top position.
| 20 | Lead gel battery for starting all electrical functions when the bed is disconnected from the main supply
| 21 | Rails for mounting accessories on the sides, below the mattress base
| 22 | Telescopic infusion stand with four hooks
| 23 | Patient lifting pole with adjustable soft touch grip
| 24 | Hooks for urine bags placed on both sides of the bed
| 25 | 4 deflection rollers at the corners of the bed
| 26 | Anti-Decubitus mattress with impermeable cover, following dimensions and features of the mattress base, height minimum 16 cm, made of diferent foam density 40-50 kg/m3
2 |  | Hospital bed |  | 48 |  | 0
| 1 | External bed dimensions 950mm±20mm x 2100mm ±50mm
| 2 | Safe working load (SWL) min. 220kg
| 3 | Movable headboard and footboard
| 7 | Steel mesh matress base, made of minimum 2 segments that are easily removable for cleaning and disinfection
| 8 | Bed frame is made of steel profiles painted with powder-colored polyester resistant to disinfectants
| 9 | Movable safety side rails which covers minimum ¾ of mattress length
| 10 | Castors are minimum Ø125mm, with brake on each castor. Minimum one castor of four is antistatic
| 11 | Backrest support easily adjustable by handles on both sides
| 12 | Backrest segment manually adjustable by gas spring, in minimal range from 0-70°
| 5 | Fixed hight of bed : 50 cm
| 6 | Trapeze for lifting a patient with adjustable grip
| 7 | Infusion stand with at least 4 hooks made of stainless or chromed steel
| 5 | Bumpers at the 4 corners of the bed
| 7 | Anti-Decubitus mattress with impermeable cover, following dimensions and features of the mattress base, height minimum 16 cm, made of diferent foam density 40-50 kg/m3
| 8 | Mattress, height minimum 16 cm, made of diferent foam density 40-50 kg/m3
| 9 | Mattress, height minimum 16 cm, made of diferent foam density 40-50 kg/m3
3 |  | Bedside cabinet without table (prepared for table) |  | 94 |  | 0
| 1 | Extrenal dimensions: 550mm ± 20mm x 450mm ± 20mm x 900mm ± 20mm
| 2 | Structure is made of quality steel with durable coating
| 3 | All the materials used for the manufacture of the cabinet are resistant to disinfectants
| 4 | 4 double castors, two with brakes
| 5 | Upper board and drawers fronts are made of HPL
| 6 | Upper drawer must have installed handle
| 7 | The lower large compartment opens up as a drawer, which in the interior has a plastic insert with bins for minimum 2 bottles and two separate compartments for personal belongs. The plastic insert can be easily removed.
| 8 | At the bottom of the cabinets is a shoe disposal section
| 9 | On sideways of the bedside cabinet is holder for eating table.
4 |  | Dry sterilizer |  | 3 |  | 0
| 1 | Dry sterilizer for sterilization up to 200 ° C
| 2 | Sterilizer working volume 60 liters ± 5%
| 3 | Workspace material - stainless steel, Workspace material thickness ≠ 1 mm
| 4 | Manual door opening with safety closure
| 5 | Double glass sealed in silicone rubber on the door to observe the working space without opening the door
| 6 | Workspace insulated with 50-70 mm thick mineral wool
| 7 | Exhaust humid air through a 50 mm diameter adjustable ventilation opening
| 8 | Two height-adjustable perforated stainless steel shelves in the workspace
| 9 | Fan with air turbine for mixing hot air
| 10 | Microprocessor temperature control controller with two LED displays in two colors to show the current temperature, set temperature and sterilizer operation time at set temperature
| 11 | Outer shell of the sterilizer made of pickled sheet metal ≠ 1.5 mm thick, plasticized
| 12 | Pt-100 L-300 temperature reading probe
5 |  | Two step foot stool |  | 51 |  | 0
| 1 | Structure made of chromed steel tubes
| 2 | Can be disassembled
| 3 | Steps are covered with anti-slip material, in black colour
| 4 | Four adjustable rubber made legs for fine levelling on uneven surfaces
6 |  | IV stand |  | 90 |  | 0
| 1 | Made of stainless steel tubes, Can be disassembled
| 2 | Five star base made od plastic on 5 wheels
| 3 | Height adjustable from 160 mm up to 200 mm
| 4 | Two plastic hooks
7 |  | Cabinet for drugs |  | 19 |  | 0
| 1 | Made of steel profiles protected with powder coating finish
| 2 | Two segments in one corpus
| 3 | Upper part is made of glass with glass door, with 3 height-adjusable steel shelves, Lower part is made of sheet steel, with 1 height-adjusable steel shelf
| 4 | Both parts have a built-in cetralčited locks
| 5 | Dimensions – double winged/ 80x40x180 – width, depth, height
8 |  | Mobile overbed table |  | 78 |  | 0
| 1 | Robust main frame on 4 (four) double 50 mm wheels, two with brakes
| 2 | Top surface made of HPL
| 3 | Dimensions of table (minimum): 40 cm x 90 cm
| 4 | The table top can be tilted, one small part of table remains horizontal
| 5 | Height adjustable table with gas spring assistance in range from 70 cm up to 100 cm
| 6 | Load capacity is minimum 15 kg
9 |  | Examination table with electrical hight adjustment |  | 28 |  | 0
| 1 | Strong basic steel frame covered with durable ABS plastic cover
| 2 | Retractable castors with activation lever on both sides of the table
| 3 | Four height adjustable feet
| 4 | Two section lying surface with gas spring adjustable back section, by levers on both sides of table
| 5 | Seamless padding 6 cm thick, made of high quality imitation leather
| 6 | Height adjustment by electric motor in min. range of 50cm up to 90 cm
| 7 | Foot switch for height adjustment function
| 8 | Back section adjustment in range from -25° up to + 50°
| 9 | Safe working load: 200 kg
| 10 | Dimensions of lying surface: 205 cm x 65 cm
| 11 | paper roll holder
10 |  | Three section screen |  | 36 |  | 0
| 1 | Structure made of chromed steel tubes
| 2 | Curtains are made of washable and self-extinguishing plastic
| 3 | On the wheels
| 4 | Dimensions: 140 x 170 mm
11 |  | Trolley for therapy |  | 6 |  | 0
| 1 | Structure made of powder-coated steel, finished with a, can be washed and disinfectedcrylic polyurethane enamels resistant to scratches, stains and corrosion
| 2 | Worktop made of self-extinguishing ABS with raised edges on three sides
| 3 | Drawers are made of powder-coated steel and run on telescopic soft-closing slides allowing their complete extraction
| 4 | Drawers are provided with ergonomic and anti-trauma handles
| 5 | Centralized key lock
| 6 | Drawer configuration: 1 drawer - height about 100 mm, 2 drawers - height about 150 mm, 2 drawers - height about 200 mm
| 7 | Ergonomic push handle on a side
| 8 | Double shock absorbent bumper to prevent the trolley from turning over
| 9 | 4 thermoformed swivel castors, Ø 125 mm, mounted on bearings; 2 are lockable
| 10 | Knee-operated 10L waste bin
| 11 | Dimensions about 650x600x1100 h mm
12 |  | Trolley for ECG |  | 12 |  | 0
| 1 | Uprights made of anodized aluminum
| 2 | Base and shelves completely made of epoxy painted steel; shelves are height adjustable. It is equipped with 3 shelves
| 3 | Base on 4 swivel castors, diam. 100 mm, with corner bumpers; 2 are lockable
| 4 | Shelves dim.: min500x400 mm
| 5 | Upright height: 750 mm ±50mm
13 |  | Strecher for patient transport |  | 10 |  | 0
| 1 | Four-piece plate from multi-layer laminate, permeable to X rays
| 2 | Construction of stretchers made of welded steel profiles, painted with powder paint based on polyester
| 3 | The structure must be protected by a casting of molded technical plastics, with a ledge for oxygen bottle and a place for storing the patient's stuff
| 4 | Change the height hydraulically over the foot control
| 5 | Lying plate height min./max .: 500mm / 800mm ± 20mm
| 6 | External dimensions of the strecher 850 x 2200mm mm ± 50mm
| 7 | Dual castors Ø 150 mm closed type with central locking system of all wheels via controls at each wheel
| 8 | The possibility of blocking one point per direction for easier steering
| 9 | Stainless steel rail mounting bracket for accessories (pumps, monitor, etc.) at the head and foot of the patient
| 10 | Fully folding safety fences, height of at least 350mm, with automatic locking system in the top position
| 11 | Handles for pushing made of stainless steel or chrome steel
| 12 | Four sleeves for mounting an infusion stand
| 13 | Telescopic stand for infusion with four hooks
| 14 | Backrest inclination adjustable with gas spring up to 70 °, ± 5 °
| 15 | Adjustable thigh inclination up to 45 °, ± 5 °
| 16 | Trendelenburg and inverted Trendelenburg 12 ° with a manual command at the head and front
| 17 | Cassette for X Ray film under the backrest
| 18 | Wall deflection rollers on all four corners of the stretcher, and two rollers on the sides, at the foot end.
| 19 | Permissible load: ≥ 250 kg
| 20 | Covering cold foam mattress with a thickness of at least 100 mm. Specific foam density ≥55kg / m³
| 21 | Mattress cover made of non-impermeable material, resistable to desinfectants
14 |  | Wheelchair for transportation |  | 24 |  | 0
| 1 | Seat-fabric with lining
| 2 | Frame- Steel-Folding
| 3 | Armrest- Folding
| 4 | Footrests- Detachable, adjustable, length, rotating
| 5 | Rear wheels- 24" Solid with radial wheels, Front wheels- 8" x 2" Solid (200x50mm)
| 6 | Min. load capacity 125kg
15 |  | Daily Care Aid-Toilet type Princess |  | 8 |  | 0
| 1 | Brakes on Rear Wheels
| 2 | Folding Frame
| 3 | Seat width 45 cm, seat depth 40cm
| 4 | Height from floor to seat adjustable 45-56 cm
| 5 | Load capacity 100 kg
16 |  | INOX trolley |  | 8 |  | 0
| 1 | Tubular structure completely made of AISI 304 stainless steel
| 2 | 2 fixed reinforced AISI 304 stainless steel shelves, thickness 0.9 mm, with raised edges.
| 3 | 2 tubular stainless steel push handles
| 4 | Base on 4 solid rubber 125 mm castors, 2 lockable;
| 5 | 4 PVC non-marking rubber bumpers.
| 6 | Dim.1000 x 600 x 950 h mm
| 7 | Shelves dimensions.: 900 x 500
17 |  | Blood sample chair electrical |  | 9 |  | 0
| 1 | Chair surface consists of 3 sections (back, seat and leg section)
| 2 | Base of the chair and other metal construction has to be covered with cover from ABS and PMMA material, to ensure easy disinfection and cleaning, shaped without corners
| 3 | Back section and leg section has to be synchronized and adjustable by lockable gas spring, to ensure the chair to be quickly set into Trendelenburg position
| 4 | Chair has to be equipped with special, red colored, shock lever, for easy and fast adjusting of the patient into Trendelenburg position. The lever is pushed by foot, so that hands stay free to help the patient
| 5 | Chair has to have maximum allowed working load of 300 kg
| 6 | Chair has to be equipped with individually lockable antistatic castors with diameter 125 mm
| 7 | Arm rests have to be height adjustable, length adjustable (along the sitting part) and can be also totally removed
| 8 | Upholstery of the chair has to be minimum 8 cm thick with seamless endings and simulated leather, which has to be resistant to blood urine, disinfectants, salt water, UV rays, microorganisms and mold
18 |  | X-ray film viewer |  | 1 |  | 0
| 1 | Dimensions: 800 mm x 430 mm
| 2 | Wall mounted
| 3 | Made of enamelled steel sheet
| 4 | Fluorescent cold white light
| 5 | White monochromatic diffuser
19 |  | O2 Therapy |  | 100 |  | 0
| 1 | O2 flowmeter, highly precise, with flow range 2 – 16 lit/min
| 2 | Conventional principle of operation, with ball as flow value indicator
| 3 | Flowmeter should be robust and tested to withstand pressures up to 80 bar which ensures safety for medical stuff and patient. Flowmeter made of aluminium, with glass flow indicator protected with pressure dome
| 4 | Flowmeter can be connnected directly to supply unit. It has long connector so it can be connected over medical rail.
| 5 | Flowmeters are compatibile with breathing gas humidifier, that is used for inhalation and humidifies gas to relative humidity of aprox. 50-70%. Humidifier is connected directly to flowmeter (M34 x 1,5 thread)
| 6 | Humidifier is autoclavable on temperatures up to 134°C, capacity 220ml
| 7 | Flowmeter is delivered with O2 mask (disposable, with connection tube 2,1m)
20 |  | Electrical suction device |  | 18 |  | 0
| 1 | Suction capacity- 50 l/min +/- 15%
| 2 | Vacuum level range from 10 to 85 kPa
| 3 | Maximum noise level at maximum load (vacuum level)- 37 dBa
| 4 | Three-stage protection mechanism via hydrophobic bacterial filter
| 5 | Autoclavable (134 ° C, at least 15 min) polysuflon suction bottle with integrated rail carrier in the bottle cap
| 6 | Suction bottle volume minimal 5000 ml
| 7 | It has a LowFriction Piston system that consists of a Pyrex glass cylinder and matching graphite piston pistons
| 8 | Device height at least 1000mm ±50mm
| 9 | The device is delivered with suitable transport trolleys and two suction bottles
21 |  | Double stainless steel (clean - not clean) bowls |  | 20 |  | 0
| 1 | Engraved “ Clean, Not clean”
| 2 | Can be fitted one into another
22 |  | Tray for instruments without lid |  | 20 |  | 0
| 1 | Made of medical stainless steel
| 2 | Dimensions: 320 x 240 x 50 mm
23 |  | Strecher for patient care |  | 4 |  | 0
| 1 | Stretcher with manual lifting of headboard without side rails
| 2 | Mechanical lifting of the backrest
| 3 | Construction of steel pipes and profiles
| 4 | Lying surface made from plywood covered with sponge and eco-leather
| 5 | Protection of the metal parts of the structure: Plastification and galvanizing
| 6 | Dimensions: 180x60x75 cm
24 |  | Trolley for ultrasound |  | 10 |  | 0
| 1 | Tubular structure completely made of AISI 304 stainless steel
| 2 | 2 fixed reinforced AISI 304 stainless steel shelves, thickness 0.9 mm, with raised edges.
| 3 | 2 tubular stainless steel push handles
| 4 | Base on 4 solid rubber 125 mm castors, 2 lockable;
| 5 | 4 PVC non-marking rubber bumpers.
| 6 | Dim.1000 x 600 x 950 h mm
| 7 | Shelves dimensions.: 900 x 500
25 |  | Kidney shape bowl (large) |  | 20 |  | 0
| 1 | Made of stainless steel
| 2 | Length 280 mm
26 |  | Kidney shape bowl (small) |  | 20 |  | 0
Made of stainless steel
Length 250 mm
27 |  | Test tube holder (bigger) |  | 20 |  | 0
| 1 | Test tube holder - 96RM
| 2 | Made of PMMA polymer
28 |  | Test tube holder (smaller) |  | 20 |  | 0
| 1 | Test tube holder - 48RM
| 2 | Made of PMMA polymer
29 |  | Operating table for General surgery and laparascopy |  | 1 |  | 0
| 1 | Mobile opearting table with minimum 5 segments (leg section,seat section, back section, upper back section, head section)
| 2 | Electromotorised adjusment of table movements
| 3 | Integrated battery supply, rechargeable batteries
| 4 | Minimal maximum table load 400 kg in central position of column and table top
| 5 | The operating tabletop is fully radiolucent between the struts
| 6 | Possibility to position table in Flex up, and flex down positions minimum 240°/100°
| 7 | Level of protection against water penetration min. IP X4
| 8 | Table control via column and cable remote controler
| 9 | 0 position function on remote controler- motorised levelling of all sections
| 10 | Four (4) rotatable, double-joint, conductive wheels, with Electric brake system which engages four retractable legs
| 11 | Table lock via remote controler
| 12 | Safety class I medical device
| 13 | Table top consists of minimum five segments and possibility to add extension segment on both sides of the table
| 14 | Table width with rails minimum 600mm
| 15 | Table lenght with all segments and extension min. 2000mm
| 16 | Foot clearance min 100mm (distance from floor to table base)
| 17 | U cut of the table base, for closer approach to the patient
| 18 | Adjustment of operating table
| 19 | Possibility of motorised positioning of min. two segment
| 20 | Minimal table height without pads, max. 650±20mm. Electromotorised adjustment
| 21 | Maximal table height without pads, min. 1000±20mmmm. Electromotorised adjustment
| 22 | Trendelenburg positioning min :+30⁰/-30⁰ (±5°) Electromotorised adjustment
| 23 | Lateral tilt min ±25⁰ (±5°) Electromotorised adjustment
| 24 | Longitudinal shift min. 350 mm. (±20mm) Electromotorised adjustment
| 25 | Back section adjustment :+80⁰/ -60⁰ (±5°) Electromotorised adjustment
| 26 | Leg section adjustment +80⁰/-92⁰ (±5°) Electromotorised adjustment
| 27 | Upper back section, with pad
| 28 | Standard accessories for mobile table
| 29 | Leg section ,two part spredable with pad
| 30 | Head section, double joint with pad, adjustable min. +45° / -30°
| 31 | Anaesthesia screen extendable, height adjustable via clamp
| 32 | Clamp radial setting, 2 pcs
| 33 | Infusion pole , 4 hooks, with clamp
| 34 | Body strap rotatable, with locking clamps
| 35 | Arm support with clamp, 2 pcs
| 36 | Leg holder, type GOEPEL, 2 pcs
| 37 | Clamp to attach GOEPEL leg holder, 2 pcs
30 |  | Examination and operation light, mobile version |  | 2 |  | 0
| 1 | LED examination light with support arms-fully rotatable
| 2 | Base of the light has 4 castors, minimum 2 are antistatic
| 3 | LED light has min 80.000 lux illumination level at 1m distance
| 4 | Colopr temperature min 4.500 K
| 5 | Color rendering Index (Ra) min 95
| 6 | Pattern size at 1m -16cm
| 7 | Dimmable from 100%-30%
| 8 | Average service life of the LEDs - min 50,000 h
| 9 | Laminar flow surface (cm2)- 340cm2
| 10 | Package has 3 sterillzable handles