## Price schedule
Bidder:
LOT 3 |  | STERILIZATION |  |  | Date:
Total price:
Line item No. | ID | Equipment name | Technical Specification Offered (model) | QTY | Unit price | Total Price per line item
1 |  | WASHING AND DISINFECTION
| 1.1 | Washer - disinfector, double door pass-through |  | 2 |  | 0
| 1.2 | Washer-disinfector for reprocessing endoscopes, single door |  | 1 |  | 0
| 1.3 | Endoscope drying and storage cabinet |  | 1 |  | 0
| 1.4 | Ultrasonic cleaning unit |  | 1 |  | 0
| 1.5 | Bedpan washer - disinfector |  | 4 |  | 0
| 1.6 | Washer - disinfector for clinical dishes |  | 2 |  | 0
| 1.7 | Disinfector spray gun |  | 1 |  | 0
TOTAL 1: WASHING AND DISINFECTION | [5 empty] | 0
2 |  | STERILIZERS
| 2.1 | Automatic steam strilizer - single door, 1 STU |  | 1 |  | 0
| 2.2 | Automatic steam strilizer - double doors pass-through, 6 STU |  | 1 |  | 0
| 2.3 | Automatic steam-formaldehyde strilizer - double doors pass-through, 6 STU |  | 1 |  | 0
| 2.4 | Plasma strilizer, 2 doors pass-through |  | 1 |  | 0
| 2.5 | Mobile air decontamination unit |  | 1 |  | 0
| 2.6 | Bactericide lamp |  | 4 |  | 0
| 2.7 | Water preparation system for sofened and RO water |  | 1 |  | 0
| 2.8 | System for tracking and tracing of sterile instruments |  | 1 |  | 0
TOTAL 2: STERILIZERS | [5 empty] | 0
3 |  | AUXILIARY EQUIPMENT
| 3.1 | Sterilization box 1 STU / full height |  | 16 |  | 0
| 3.2 | Sterilization box 1 STU / half height |  | 16 |  | 0
| 3.3 | Sterilization box 1/2 STU / full height |  | 16 |  | 0
| 3.4 | Wire basket 1 STU / full height |  | 20 |  | 0
| 3.5 | Wire basket 1 STU / half height |  | 20 |  | 0
| 3.6 | Magnifying lamp |  | 1 |  | 0
| 3.7 | Rotary heat sealers for bonding sterilization foil |  | 2 |  | 0
| 3.8 | Roll holders with cutting device |  | 2 |  | 0
| 3.9 | Troley for sterile material transport |  | 2 |  | 0
| 3.10 | Paper trolley |  | 1 |  | 0
| 3.11 | Wall bracket |  | 3 |  | 0
| 3.12 | General packing table |  | 3 |  | 0
| 3.13 | Linnen inspection an folding table with shelf |  | 1 |  | 0
| 3.14 | Instrument washing sink |  | 1 |  | 0
| 3.15 | Solid shelf system for clean laundry storage |  | 33 |  | 0
| 3.16 | Spray gun |  | 1 |  | 0
TOTAL 3: AUXILIARY EQUIPMENT | [5 empty] | 0
Total STERILIZATION (Total1+Total2+Total3) | [5 empty] | 0

## Specifications
Bidder:
LOT 3 |  | STERILIZATION |  |  | Date:
Total price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | Final unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | WASHING AND DISINFECTION
| 1.1 | Washer - disinfector, double door pass-through |  | 2 |  | 0
| 1 | Pass-through model with hinged glass doors and electrical heaters.
| 2 | Capacity 10 DIN trays, loading on 5 levels
| 3 | Each unit equipped with 1 piece of 5 level wash cart.
| 4 | Rotational spray arms at top and bottom of chamber
| 5 | Independant temperature and validation test port
| 6 | PLC controlled with microprocessor and battery for back-up, possibility to store 20 programs
| 7 | Control panel with LCD screen and control buttons and/ or touch screen
| 8 | Control panel displays all informtions about the cycle and alarms
| 9 | Choosing of programs by buttons or by touch-screen
| 10 | Integrated printer of process parameters at front panel
| 11 | HEPA filter controlled by differential pressure measurement, device has integrated hot air drying
| 12 | Doors made mainly of glass, light inside the chamber
| 13 | Equipped with system for cooling of waste water, max. allowed temperature of waste water 70*C
| 14 | System for control of remaining detergent and flow of detergents on pumps.
| 15 | Dimensions: width max. 700 mm, depth max. 700 mm, height max. 2.100 mm, weight max. 200 kg
| 16 | Standards: EN 15883 (Declaration of Conformity), MDD 93/42/EC, CE marked
| 17 | Device must be equipt with: transport cart (2pc), instrument rack for loading 12 standard instrument trays (1pc),
| 18 | Rack for AN material (1pc), rack for container for sterilization (1pc), 2 level rack (1pc) with basket for OR shoes (2pc)
| 19 | Universal basket for kidney dishes (1pc), MIS rack (1pc) and tray for for MIS inst. That can be taken apart (1pc)
| 1.2 | Washer-disinfector for reprocessing endoscopes, single door |  | 1 |  | 0
| 1 | Free standing model intended for reprocessing flexible and rigid endoscopes
| 2 | The capacity of device min 2 pc flexible endocopes
| 3 | Disinfector has two completely asynchronous chambers for reprocessing 2 flexibile endosopes
| 4 | Front loading, single door (non-passable model)
| 5 | Device has a door whose frame is made of high quality stainless steel AISI 316L, and HST reinforced glass
| 6 | Device has a cycle locking system that prevents the operator from opening the door during the cycle
| 7 | Device provides liquid sterilization of flexible endoscopes
| 8 | Device automatically stops the cycle in case of any anomalies and displays an alarm
| 9 | Leak tests and channels monitoring are executed during the whole washing / disinfection cycle
| 10 | USB port on the front panel for storing cycle data
| 11 | Build-in cycle printer and LED status displey
| 1.3 | Endoscope drying and storage cabinet |  | 1 |  | 0
| 1 | The capacity of the drying and storing cabinet is up to 8 flexible endoscopes
| 2 | The external frame and machine panels are made of AISI 304 stainless steel
| 3 | HST temperate double glass hinged door
| 4 | Filtration system composed of HEPA filter
| 5 | Device is compliant with the EN ISO 16442 directive
| 7 | Visual and audio alarms in the drying cabinet for signaling possible faults to the user
| 8 | Control panel with 3 digit LED display
| 1.4 | Ultrasonic cleaning unit |  | 1 |  | 0
| 1 | Tank volume 20-30 l
| 2 | Ultrasound frequncy 35-40 kHz
| 3 | Total power max. 2 kW
| 1.5 | Bedpan washer - disinfector |  | 4 |  | 0
| 1 | For washing, thermal disinfection and partial drying of material
| 2 | Microprocessor controlled , with control panel at the front of the unit
| 3 | 3 standard programs
| 4 | Chamber volume cca 70 liters, 2 dosing pumps for detergents
| 5 | Installed power max. 5 kW
| 6 | Dimensions: 600x650x1800 (+/- 100) mm
| 7 | Standards: EN ISO 13485, MDD 93/42/EEC (CE), EN 61000-3-2 and EN 50081-2, EN 15583
| 1.6 | Washer - disinfector for clinical dishes |  | 2 |  | 0
| 1 | Machine usede for cleaning and desinfection of clinical dishes
| 2 | Automatic door locking and unlocking
| 3 | 3 available programs: Short, Standard and Intensive
| 4 | Program cycle of maximum 10 min
| 5 | Machine with integrated drain cooling in order to reduce the drain water temperature to 60°C or below.
| 6 | Machine with integrated hot air druying. After washing and disinfection phases, drying phase with HEPA filter that ensure that the goods are quickly and efficiently dried.
| 7 | Reverse osmosis system including stop valve and 10 μm fine filter with active carbon
| 8 | Device must be equipt with: universal rack for clinical dishes (1pc), rack for operating theater boots capacity 6 boots (1pc)
| 1.7 | Disinfector spray gun |  | 1 |  | 0
| 1 | Water and compressed air connection
| 2 | Thermally insulated handle of the spray gun
| 3 | Hose, min. 1 m length, nozzles - 8 pcs
2 |  | STERILIZERS
| 2.1 | Automatic steam strilizer - single door, 1 STU |  | 1 |  | 0
| 1 | Chamber volume min. 60 liters, 1 STU, single door (non-passable model)
| 2 | The device is used for sterilization of various medical material with steam at temperatures of 121-134 ° C.
| 3 | Microprocessor controlled: steam sterilizer, with manual vertically sliding door (one) - non-passable model
| 4 | Steam supply from own built-in steam generator.
| 5 | Chamber volume: min. 60 lit. (1 STU)
| 6 | The chamber door is manually operated, sliding vertically, heat and sound insulated.
| 7 | The door must have a security system to prevent it from opening during the process.
| 8 | Steam generator power: max. 10 kW
| 9 | The steam generator heaters are not in direct contact with water
| 10 | Feed water: municipality water
| 11 | The device has an injector vacuum pump.
| 12 | The control panel is "touch-screen" type, with a minimum diagonal of 8 ".
| 13 | The device itself has a printer that records the parameters of the steril. process.
| 14 | Water preparation system integrated: integrated RO system
| 15 | Appliance dimensions: max. 600x1600x900 mm (width x height x depth)
| 16 | Standards: EN ISO 13485, MDD 93/42 EEC, PED 97/23 EC,(SRPS) EN 285
| 2.2 | Automatic steam strilizer - double doors pass-through, 6 STU |  | 1 |  | 0
| 1 | Temperature range is 121-135°C (or wider, including this range).
| 2 | The device is free-standing unit with, or without side panels, positioned between 2 walls.
| 3 | Chamber capacity: 6 STU, 450-500 lit
| 4 | Device has 2 doors which is vertically sliding (up-down), doors movement is automatic, achieved by pneumatic cylinder.
| 5 | The sterilizer is equipped with integrated steam generator, installed power: 50-60 kW.
| 6 | The process is completely controlled by micro-processor control system (PLC).
| 7 | The user interface (operating panel) is in the form of color touch-screen, diagonal is minimum 10” on both side of sterilizer
| 8 | The external communication is achieved over integrated RS232, or Ethernet port.
| 9 | The user identification by using some “contact-less” identification system
| 10 | Pressure and temperature sensors control the sterilization process.
| 11 | The unit is equipped with a printer 2” or bigger.
| 12 | The sterilizer is equipped with USB connection for recording process cycles in PDF file
| 13 | Sterilizer support structure must be made of stainless steel
| 14 | Operating Performance (EN285, full textile load): Process time: max. 35 minutes per cycle with full capacity
| 15 | Installation Height (bottom to top): max. 210 cm
| 16 | Installation Depth (front to back): max. 140 cm
| 17 | Chamber loading height (bottom internal chamber surface): 700-800 mm
| 18 | Standards: EN ISO 13485, MDD 93/42 EEC, PED 97/23 EC,(SRPS) EN 285,
| 19 | Device must be equipt with: 1 loading cart and 2 transport wagons
| 2.3 | Automatic steam-formaldehyde strilizer - double doors pass-through, 6 STU |  | 1 |  | 0
| 1 | Temperature range is 121-135°C (or wider, including this range).
| 2 | Sterilizer performs sterilization by means of dry-saturated steam. Pre-vacuum, sterilization and post-vacuum drying are main cycle phases.
| 3 | In addition, sterilizer has the option for sterilization of heat-sensitive materials by formaldehyde in temperature range 55-80°C (or wider, including this range).
| 4 | Formaldehyde is in the liquid stage inside the bottle or similar container, which is positioned in the compartment within the sterilizer before the cycle start.
| 5 | At the end of the cycle the device will suck out and remove all remaining formaldehyde out of the bottle, so it can safely be disposed.
| 6 | The device is free-standing unit with, or without side panels, positioned between 2 walls.
| 7 | Chamber capacity: 6 STU, 450-500 lit
| 8 | Device has 2 doors which is vertically sliding (up-down), doors movement is automatic, achieved by pneumatic cylinder.
| 9 | The sterilizer is equipped with integrated steam generator, installed power: 50-60 kW.
| 10 | The process is completely controlled by micro-processor control system (PLC).
| 11 | The user interface (operating panel) is in the form of color touch-screen, diagonal is minimum 10” on both side of sterilizer
| 12 | The external communication is achieved over integrated RS232, or Ethernet port.
| 13 | The user identification by using some “contact-less” identification system
| 14 | Pressure and temperature sensors control the sterilization process.
| 15 | The unit is equipped with a printer 2” or bigger.
| 16 | The sterilizer is equipped with USB connection for recording process cycles in PDF file
| 19 | Operating Performance (EN285, full textile load): Process time: max. 35 minutes per cycle with full capacity
| 20 | Installation Height (bottom to top): max. 210 cm
| 21 | Installation Depth (front to back): max. 140 cm
| 22 | Chamber loading height (bottom internal chamber surface): 700-800 mm
| 23 | Standards: EN ISO 13485, MDD 93/42 EEC, PED 97/23 EC,(SRPS) EN 285
| 24 | Device must be equipt with: 1 loading cart and 2 transport wagons
| 2.4 | Plasma strilizer, 2 doors pass-through |  | 1 |  | 0
| 1 | Unit for sterilisation of thermo sensible medical elements made of metal or other materials
| 2 | Sterilisation temperature 50-60*C
| 3 | Automatic vertical sliding doors, with pneumatic or electro-mechanic drive
| 4 | Chamber volume usable: 100-140 liters, horizontal chamber with square cross section with rounded edges
| 5 | In chamber plasma generation
| 6 | Integrated HEPA filter, hydrophobicfor air filtration of the air which is introduced into the chamber during a cycle for pressure control inside the chamber
| 7 | Integrated catalyser for the decomposition of H2O2, placed between the sterilization chamber and the vacuum pump in order to prevent harmful emissions and for the protection of the vacuum pump
| 8 | Unit has colour touch-screen with min. 7" diagonal, placed at the front of the unit, above the chamber
| 9 | Integrated thermal printer placed beside/above the door of the chambe
| 10 | Safety system for opening of doors during cycle and preventing closing of doors in case of any obstacles
| 11 | Device is equipt with 3 program for sterilization
| 12 | Standard program for the steril. of instruments, flexible and rigid lumens with steril. time of max. 50 min
| 13 | Fast program for surface sterilisation of smaller quant. of material (3-4 kg) with steril. time of max. 35 min
| 14 | Dimensions: width max. 700 mm, depth max. 900 mm, height max. 2.000 mm
| 15 | Standards applied: MDD 93/42/EC, ISO 14937
| 2.5 | Mobile air decontamination unit |  | 1 |  | 0
| 1 | Mobile unit for disinfection of rooms and material
| 2 | Disinfection by means of hydrogenperoxyde fog
| 3 | Destruction range: bactericidal, virucidal and fungicidal (including spores)
| 4 | Reduction rate of micfroroganisms:
| 4.1 | bacteria and mycobacteria >= 5 log,
| 4.2 | viruses, yeast and funghi >= 4 log,
| 4.3 | spores >= 3 log
| 5 | Touch-screen controls, USB port for data download, integral printer
| 6 | Automatic connection to unit for ventilation process after performed decontamination
| 7 | Fast removal of biocide after automated surface disinfection procedure
| 8 | Performance validated to standard NF-T 72-281
| 2.6 | Bactericide lamp |  | 4 |  | 0
| 1 | Bactericide device with air flow
| 2 | With swich and working hours counter
| 3 | UV source radiation 2x55W
| 2.7 | Water preparation system for sofened and RO water |  | 1 |  | 0
| 1 | Unit for softening of tap water, type duplex
| 2 | Capacity of system should be as per sterilizers and washer disinfectors consumption
| 3 | Time and volume controlled
| 4 | Unit for demineralisation of softened water type RO
| 5 | Quality and quantities of treated water appropriate for use with steam sterilisers and washing machines
| 6 | Consists of housing with RO membrane, all necessary automation and tank with pump for distribution of permeate
| 2.8 | System for tracking and tracing of sterile instruments |  | 1 |  | 0
| 1 | The system completely covers the instruments traceability in CSSD.
| 2 | One fuctionality is that all of the washers and sterilizers are connected to the system
| 3 | Possibility to monitor in real time and record the cycles data.
| 4 | With supporting hardware solution: workstations, printers, scaners etc.
3 |  | AUXILIARY EQUIPMENT
| 3.1 | Sterilization box 1 STU / full height |  | 16 |  | 0
| 1 | Dimensions: 600x300x300 mm
| 2 | Material: aluminium
| 3.2 | Sterilization box 1 STU / half height |  | 16 |  | 0
| 1 | Dimensions: 600x300x150 mm
| 2 | Material: aluminium
| 3.3 | Sterilization box 1/2 STU / full height |  | 16 |  | 0
| 1 | Dimensions: 300x300x300 mm
| 2 | Material: aluminium
| 3.4 | Wire basket 1 STU / full height |  | 20 |  | 0
| 1 | Dimensions: 600x300x300 mm
| 2 | Material: stainless steel
| 3.5 | Wire basket 1 STU / half height |  | 20 |  | 0
| 1 | Dimensions: 600x300x150 mm
| 2 | Material: stainless steel
| 3.6 | Magnifying lamp |  | 1 |  | 0
| 1 | Inspection lamp with integrated magnifier for professional use
| 3.7 | Rotary heat sealers for bonding sterilization foil |  | 2 |  | 0
| 1 | Microprocessor controlled automatic sealing device to be used for sealing of sterilisation reels
| 2 | Control of critical parameters according to ISO 11607-2
| 3 | User interface by means of a colour touch-screen with min. 4" diagonal
| 4 | Possibility to record process parameters with 3 programs for quick access
| 5 | Automatic stand-by, when not in use, safety system for overheating
| 6 | Slow-start function for reels with bigger dimensions
| 7 | Equipped with Ethernet and USB connections
| 8 | Working temperature: 100-200*C
| 9 | Sealing speed: 8-10 m/min
| 10 | Dimensions max. 600x300x200 mm
| 11 | Standards: ISO 13485, LV 2006/95/EC, EMC 2004/108/EC
| 3.8 | Roll holders with cutting device |  | 2 |  | 0
| 1 | 4 carrying and positioning bars for rolls
| 2 | For storing and cutting of sterilisation rolls
| 3 | Spare set of cutting knives
| 4 | Dimensions 120x35x20 cm (+/- 5 cm)
| 3.9 | Troley for sterile material transport |  | 2 |  | 0
| 1 | Closed troley with 2 shelves
| 2 | Material stainless steel; handles, wheels made of other appropriate materials
| 3 | Inside the trolley there are non-fixed L-profiles which allow the positioning of STU
| 4 | With secure door locking without ortifices in the trolley
| 5 | Capacity: 6 STU
| 3.1 | Paper trolley |  | 1 |  | 0
| 1 | 4 carrying and positioning bars adjustable for 4 different diameters of rolls
| 2 | Mobile, with 4 wheels, where of at least 2 wheels with brakes
| 3 | Dimensions 140 (+/-10 cm) x 70 (+/- 10 cm) x 100 cm (+/- 10 cm)
| 3.11 | Wall bracket |  | 3 |  | 0
| 1 | Dimensions: 1800x700x850 mm ±5%
| 2 | Material: stainless steel, AISI 304
| 3.12 | General packing table |  | 3 |  | 0
| 1 | Material: stainless steel AISI 304 or better
| 2 | Dimensions: 1.200x750 mm ±5%
| 3 | Height adjustable 750-1200 mm ±5%
| 4 | LED light below lower shelf, 3 power connections 220 V, 50 Hz
| 5 | 2 shelves above main working surface.
| 3.13 | Linnen inspection and folding table |  | 1 |  | 0
| 1 | Table for inspection and folding of linen
| 2 | Integrated partially transparent plastic part in working surface with lamp for illumination installed below working surface
Lamp 220 V, 50 Hz installed below transparent part of the working surface
| 3 | Dimensions: 2.000x1.250x900 mm ±5%
| 3.14 | Instrument washing sink |  | 1 |  | 0
| 1 | Instrument washing sink with two basin
| 2 | Two combined taps for cold/hot water
| 3 | Two sets of drainers
| 4 | Dimensions: 1800x700x850 mm ±5%
| 5 | Lower shelf included.
| 3.15 | Solid shelf system for clean laundry storage |  | 33 |  | 0
| 1 | material: stainless steel, AISI 304
| 2 | non-perforated shelves, 5 pcs
| 3 | shelf dimension 1200x600 mm (+/- 50 mm)
| 4 | shelf rack height 1800 - 2000 mm
| 3.16 | Spray gun |  | 1 |  | 0
| 1 | Water and compressed air connection
| 2 | Thermally insulated handle of the spray gun
| 3 | Hose, min. 1 m length, nozzles - 8 pcs