## Price Schedule
Bidder: |  |  |  | Offered price (currency): | 0
LOT 10 | MEDICAL FURNITURE
ID | Equipment name (Technical Specification Requested) | QTY | Technical Specification Offered (model) | Unit price | Total price per line item
1 | Strecher for patient transportation | 11 |  |  | 0
2 | Mobile trolley for anesthesia | 6 |  |  | 0
3 | Intensive care bed 1 | 20 |  |  | 0
4 | Intensive care bed 2 | 5
5 | Mobile trolley for intensive care | 10 |  |  | 0
6 | Mobile trolley for emergency | 4 |  |  | 0
7 | Semi-intensive care bed | 12 |  |  | 0
8 | Bedside cabinet with over-bed table | 12 |  |  | 0
9 | Hospital bed with electrical functions | 71 |  |  | 0
10 | Bedside cabinet with over-bed table | 30 |  |  | 0
11 | Bedside cabinet | 41 |  |  | 0
12 | Electro-hydraulic examination bed | 20 |  |  | 0
13 | Mobile screen, 3-elements | 40 |  |  | 0
14 | Telescopic infusion holder | 50 |  |  | 0
15 | Mobile examination lamp | 25 |  |  | 0
16 | Column scale | 8 |  |  | 0
17 | Chair for blood sampling | 2 |  |  | 0
18 | Hydraulic instrument table | 30 |  |  | 0
19 | Wheelchair | 10 |  |  | 0
|  |  |  | Total Price per Lot: | 0

## MedicalFurniture
Bidder:
LOT 10 |  | MEDICAL FURNITURE |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Strecher for patient transportation |  | 11 |  | 0
| 1 | Metal construction coated with impact-resistant epoxy paint
| 2 | Hydraulic - height adjustment using foot pedals - according to the system of two columns adjustable in height or equivalent
| 3 | Height adjustment range minimum height: 55-90 ±2 cm and larger range
| 4 | Height and trendeleburg height adjustment / reverse trendeleburg, using foot pedals
| 5 | Backrest adjustment: 0-85 ° and larger range
| 6 | Trendelenburg 17 ° / Cloudy Trendelenburg 17 °
| 7 | Mattress platform, two-piece, made of wire mesh, dimensions: 200 cm x 65 cm with a tolerance of +/- 2 cm
| 8 | Telescopic infusion carrier with 4 hooks
| 9 | Space for placing an oxygen bottle on a metal base
| 10 | The control handles on the patient's head and legs can be removed
| 11 | Side, metal sides that can be easily and quickly lowered,
| 12 | 4-point diameter minimum 200 mm, antistatic, with independent brakes on all four wheels.
| 13 | Mattress for stretchers, 10 cm thick with a deviation of +/- 1 cm, dim. 198 x 65 x 10 cm
| 14 | Dimensions: 205 cm x 80 cm with a resolution of +/- 5 cm
2 |  | Mobile trolley for anesthesia |  | 6 |  | 0
| 1 | Dimensions: 95 x 55 x 160 cm ± 5 cm
| 2 | Aluminum base, top plate and ABS plastic sides
| 3 | Easy to clean and disinfect
| 4 | Spacious upper work surface, limited and rounded edge
| 5 | 4 wheels 125 mm, antistatic with front wheel brakes
| 6 | 2 plastic drawers height 70 mm
| 7 | 3 plastic drawers 145 mm high
| 8 | Side drawer on the top panel made of ABS pull-out with laptop stand
| 9 | Side drawer on the top panel made of ABS pull-out with space for storing accessories with a lid for a laptop stand
| 10 | Locking system for side drawers
| 11 | Door with locking system for drawers on the front side
| 12 | Bracket on the top plate
| 13 | Stainless steel instrument bowl with 1 plastic compartment
| 14 | Large stainless steel instrument bowl with 3 plastic compartments
| 15 | Bracket with two hooks for hanging the infusion
| 16 | Ampoule opener
| 17 | Wire basket on the side of the cart
| 18 | Three cassettes on the side of the cart, plastic, opening at an angle
| 19 | Possibility of combining drawers by dimensions
| 20 | Possibility to choose the color of the drawer
3 |  | Intensive care bed |  | 20 |  | 0
| 1 | Medical bed for intesive care with electric functions
| 2 | Base frame of welded and epoxy-coated steel pipes
| 3 | Frame provided with proper housings for fitting patient lifting pole and I.V. poles at the 4 corners
| 4 | Mattress platform, 4-section, 3 of which jointed and 1 fixed centrally, made of polypropylene (PP), easy and quick maintenance and disinfection , backrest made of HPL - radiolucent
| 5 | Contol unit for medical stuff, placed on foot-end, with inhibition of the electric controls on the safety sides
| 6 | Contol unit for medical stuff with following functions : high-low height adjustment, backrest adjustment, knee-break adjustment, auto-contour function, cardiac chair position, trendlenburg/reverse trendelenburg positions, examination positions, cardiac massage position, anti-shock position
| 7 | Drawer for store of control unit for medical stuff, placed below of foot-end
| 8 | Break assist moving system for easy moving the bed on 10% incline
| 9 | System for control the forward/backward movement of the bed
| 10 | The bed must have anti-crushing sensor that ensures the operators safety during backwards maneuvers
| 11 | 4 sector safety side with integrated electric control
| 12 | Bilateral emergency CPR control
| 13 | 2-color led indication system
| 14 | Translation of min. 130 mm
| 15 | Head and Foot ends, made of polypropilene, easily removable
| 16 | Head and Foot ends equipped with locking system
| 17 | Angle indicators on the safety sides
| 18 | Legrest with manual lifting
| 19 | 4 antistatic, Ø 150 mm twin-wheel casters with central locking
| 20 | Integrated bed lengthener : min. 300 mm
| 21 | Breaking sound alarm
| 22 | Telescopic I.V. pole with 4 hooks, made of stainless steel
| 23 | Patient lifting pole with adjustable handle
| 24 | Overall dimensions : 219x100 cm ± 3 cm
| 25 | Mattress platform height : 39-79 cm ± 2 cm
| 26 | Backrest section : 0°- 65°
| 27 | Knee-break section : 0°- 30°
| 28 | Leg section : -10°/+10°
| 29 | Trendelenburg/Reverse trendelenburg : +16°/-16°
| 30 | Safe working load : min. 300 Kg
4 |  | Intensive care bed 2 |  | 5 |  | 0
| 1 | Medical bed for intesive care with electric functions
| 2 | Base frame of welded and epoxy-coated steel pipes
| 3 | Frame provided with proper housings for fitting patient lifting pole and I.V. poles at the 4 corners
| 4 | Mattress platform, 4-section, 3 of which jointed and 1 fixed centrally, made of polypropylene (PP), easy and quick maintenance and disinfection , backrest made of HPL - radiolucent
| 5 | Contol unit for medical stuff, placed on foot-end, with inhibition of the electric controls on the safety sides
| 6 | Contol unit for medical stuff with following functions : high-low height adjustment, backrest adjustment, knee-break adjustment, auto-contour function, cardiac chair position, trendlenburg/reverse trendelenburg positions, examination positions, cardiac massage position, anti-shock position
| 7 | Drawer for store of control unit for medical stuff, placed below of foot-end
| 8 | Break assist moving system for easy moving the bed on 10% incline
| 9 | System for control the forward/backward movement of the bed
| 10 | Integrated weighing system with high level of precision independently from the mattress platform position
| 11 | The bed must have anti-crushing sensor that ensures the operators safety during backwards maneuvers
| 12 | 4 sector safety side with integrated electric control
| 13 | Bilateral emergency CPR control
| 14 | 2-color led indication system
| 15 | Translation of min. 130 mm
| 16 | Head and Foot ends, made of polypropilene, easily removable
| 17 | Head and Foot ends equipped with locking system
| 18 | Angle indicators on the safety sides
| 19 | Legrest with manual lifting
| 20 | 4 antistatic, Ø 150 mm twin-wheel casters with automatic self-brake system
| 21 | Integrated bed lengthener : min. 300 mm
| 22 | Breaking sound alarm
| 23 | Telescopic I.V. pole with 4 hooks, made of stainless steel
| 24 | Patient lifting pole with adjustable handle
| 25 | Overall dimensions : 219x100 cm ± 3 cm
| 26 | Mattress platform height : 39-79 cm ± 2 cm
| 27 | Backrest section : 0°- 65°
| 28 | Knee-break section : 0°- 30°
| 29 | Leg section : -10°/+10°
| 30 | Trendelenburg/Reverse trendelenburg : +16°/-16°
| 31 | Safe working load : min. 300 Kg
| 32 | Antidecubitus mattress with compressor
5 |  | Mobile trolley for intensive care |  | 10 |  | 0
| 1 | Dimensions: 100 x 60 x 160 cm±5cm
| 2 | Aluminum base, top plate and ABS plastic sides
| 3 | Easy to clean and disinfect
| 4 | Spacious upper work surface, limited and rounded edge
| 5 | 4 wheels 125 mm, antistatic with front wheel brakes
| 6 | 2 plastic drawers height 70 mm
| 7 | 3 plastic drawers 145 mm high
| 8 | Container for small waste
| 9 | Support for box gloves
| 10 | Locking system for side drawers
| 11 | Door with locking system for drawers on the front side
| 12 | Bracket on the top plate
| 13 | Stainless steel instrument bowl with 1 plastic partition
| 14 | Large stainless steel instrument bowl with 3 plastic compartments
| 15 | Bracket with two hooks for hanging the infusion
| 16 | Ampoule opener
| 17 | Wire basket on the side of the cart
| 18 | Three cassettes on the side of the cart, plastic, opening at an angle
| 19 | Possibility of combining drawers by dimensions
| 20 | Possibility to choose the color of the drawer
6 |  | Mobile trolley for emergency |  | 4 |  | 0
| 1 | Dimensions: 90 x 70 x 160 cm±5cm
| 2 | Trolley base made of aluminum, top plate and sides made of ABS plastic, which allows easy cleaning and maintenance
| 3 | Spacious upper work surface, limited by rounded edges
| 4 | 4 wheels 125 mm, antistatic, with front wheel brakes
| 5 | 2 plastic drawers 70 mm high, of appropriate dimensions with medicine cabinet
| 6 | 3 plastic drawers 145 mm high, of appropriate dimensions with medicine cabinet
| 7 | Side drawer on the top panel made of pull-out ABS plastic with laptop stand
| 8 | Side drawer on the top panel made of pull-out ABS plastic with space for storing accessories with a lid
| 9 | Locking system for side drawers
| 10 | Door with locking system for drawers on the front page
| 11 | Telescopic infusion bracket with 4 hooks, height adjustable
| 12 | Oxygen bottle holder with 2 safety straps for fastening
| 13 | Mesh bracket for suction pump with 2 safety straps for fastening
| 14 | Stainless steel defibrillator bracket, with 360-degree rotation system, 2 safety defibrillator straps
| 15 | Plexiglas CPR plate, located at the rear of the stroller
| 16 | Possibility of combining drawers by dimensions
| 17 | Possibility to choose colors for drawers
7 |  | Semi-intensive care bed |  | 12 |  | 0
| 1 | Medical bed for intesive care with electric functions
| 2 | Base frame of welded and epoxy-coated steel pipes
| 3 | Frame provided with proper housings for fitting patient lifting pole and I.V. poles at the 4 corners
| 4 | Mattress platform, 4-section, 3 of which jointed and 1 fixed centrally, made of polypropylene (PP), easy and quick maintenance and disinfection
| 5 | Contol unit for medical stuff, placed on foot-end, with inhibition of the electric controls on the safety sides
| 6 | Contol unit for medical stuff with following functions : high-low height adjustment, backrest adjustment, knee-break adjustment, auto-contour function, cardiac chair position, trendlenburg/reverse trendelenburg positions, examination positions, cardiac massage position, anti-shock position
| 7 | Drawer for store of control unit for medical stuff, placed below of foot-end
| 8 | Break assist moving system for easy moving the bed on 10% incline
| 9 | System for control the forward/backward movement of the bed
| 10 | The bed must have anti-crushing sensor that ensures the operators safety during backwards maneuvers
| 11 | 4 sector safety side with integrated electric control
| 12 | Bilateral emergency CPR control
| 13 | 2-color led indication system
| 14 | Translation of min. 130 mm
| 15 | Head and Foot ends, made of polypropilene, easily removable
| 16 | Head and Foot ends equipped with locking system
| 17 | Angle indicators on the safety sides
| 18 | Legrest with manual lifting
| 19 | 4 antistatic, Ø 150 mm twin-wheel casters with central locking
| 20 | Integrated bed lengthener : min. 300 mm
| 21 | Breaking sound alarm
| 22 | Telescopic I.V. pole with 4 hooks, made of stainless steel
| 23 | Patient lifting pole with adjustable handle
| 24 | Overall dimensions : 219x100 cm ± 3 cm
| 25 | Mattress platform height : 39-79 cm ± 2 cm
| 26 | Backrest section : 0°- 65°
| 27 | Knee-break section : 0°- 30°
| 28 | Leg section : -10°/+10°
| 29 | Trendelenburg/Reverse trendelenburg : +16°/-16°
| 30 | Safe working load : min. 300 Kg
8 |  | Bedside cabinet with over-bed table |  | 12 |  | 0
| 1 | Bedside cabinet with over-bed table
| 2 | Structure made of polyethylene with rounded edges
| 3 | Material and colour must be in accordance with material and colour of the medical beds
| 4 | Dimensions: 60 x 45 x 85 cm±2cm
| 5 | 4 double castors, 2 front castors with brake-system
| 6 | Upper worktop with fluid-retaining edges, resistant on fluids and disinfectant
| 7 | Extractable monobloc drawer with key-lock
| 8 | Inside of drawer with 3 plastic partition, easily be removed and easy for cleaning
| 9 | Lower door with integrated handle
| 10 | Lower panel of cabinet perforated – for ventilation
| 11 | Intermediate open space to allow the introduction of over-bed table
| 12 | Over-bed table with 4 swiveling castors, 2 front castors with brake-system
| 13 | Bottle-holder with two places, made of ABS
| 14 | Anti-scratch and stratified laminate top, with fluid retain edges, colour in accordance with colour of cabinet,
| 15 | Laminate top height adjustable
| 16 | Laminate top have lateral shift and have a possibility of setting in position for feeding and reading
| 17 | Removable vanity box with integrated mirror
| 18 | Towels holder
| 19 | Bedside cabinet with Over-bed table is easy for cleaning and disinfection
9 |  | Hospital bed with electrical functions |  | 71 |  | 0
| 1 | Medical bed for intesive care with electric functions
| 2 | Base frame of welded and epoxy-coated steel pipes
| 3 | Frame provided with proper housings for fitting patient lifting pole and I.V. poles at the 4 corners
| 4 | Mattress platform, 4-section, 3 of which jointed and 1 fixed centrally, made of polypropylene (PP), easy and quick maintenance and disinfection
| 5 | Contol unit for medical stuff, placed on foot-end, with inhibition of the electric controls on the safety sides
| 6 | Contol unit for medical stuff with following functions : high-low height adjustment, backrest adjustment, knee-break adjustment, auto-contour function, cardiac chair position, trendlenburg/reverse trendelenburg positions, examination positions, cardiac massage position, anti-shock position
| 7 | Drawer for store of control unit for medical stuff, placed below of foot-end
| 8 | Break assist moving system for easy moving the bed on 10% incline
| 9 | System for control the forward/backward movement of the bed
| 10 | The bed must have anti-crushing sensor that ensures the operators safety during backwards maneuvers
| 11 | 4 sector safety side with integrated electric control
| 12 | Bilateral emergency CPR control
| 13 | 2-color led indication system
| 14 | Translation of min. 130 mm
| 15 | Head and Foot ends, made of polypropilene, easily removable
| 16 | Head and Foot ends equipped with locking system
| 17 | Angle indicators on the safety sides
| 18 | Legrest with manual lifting
| 19 | 4 antistatic, Ø 150 mm twin-wheel casters with central locking
| 20 | Integrated bed lengthener : min. 300 mm
| 21 | Breaking sound alarm
| 22 | Telescopic I.V. pole with 4 hooks, made of stainless steel
| 23 | Patient lifting pole with adjustable handle
| 24 | Overall dimensions : 219x100 cm ± 3 cm
| 25 | Mattress platform height : 39-79 cm ± 2 cm
| 26 | Backrest section : 0°- 65°
| 27 | Knee-break section : 0°- 30°
| 28 | Leg section : -10°/+10°
| 29 | Trendelenburg/Reverse trendelenburg : +16°/-16°
| 30 | Safe working load : min. 300 Kg
10 |  | Bedside cabinet with over-bed table |  | 30 |  | 0
| 1 | Bedside cabinet with over-bed table
| 2 | Structure made of polyethylene with rounded edges
| 3 | Material and colour must be in accordance with material and colour of the medical beds
| 4 | Dimensions: 60 x 45 x 85 cm±2cm
| 5 | 4 double castors, 2 front castors with brake-system
| 6 | Upper worktop with fluid-retaining edges, resistant on fluids and disinfectant
| 7 | Extractable monobloc drawer with key-lock
| 8 | Inside of drawer with 3 plastic partition, easily be removed and easy for cleaning
| 9 | Lower door with integrated handle
| 10 | Lower panel of cabinet perforated – for ventilation
| 11 | Intermediate open space to allow the introduction of over-bed table
| 12 | Over-bed table with 4 swiveling castors, 2 front castors with brake-system
| 13 | Bottle-holder with two places, made of ABS
| 14 | Anti-scratch and stratified laminate top, with fluid retain edges, colour in accordance with colour of cabinet,
| 15 | Laminate top height adjustable
| 16 | Laminate top have lateral shift and have a possibility of setting in position for feeding and reading
| 17 | Removable vanity box with integrated mirror
| 18 | Towels holder
| 19 | Bedside cabinet with Over-bed table is easy for cleaning and desinfection
11 |  | Bedside cabinet |  | 41 |  | 0
| 1 | Bedside cabinet
| 2 | Structure made of polyethylene with rounded edges
| 3 | Material and colour must be in accordance with material and colour of the medical beds
| 4 | Dimensions: 60 x 45 x 85 cm±2cm
| 5 | 4 double castors, 2 front castors with brake-system
| 6 | Upper worktop with fluid-retaining edges, resistant on fluids and disinfectant
| 7 | Extractable monobloc drawer with key-lock
| 8 | Inside of drawer with 3 plastic partition, easily be removed and easy for cleaning
| 9 | Lower door with integrated handle
| 10 | Lower panel of cabinet perforated – for ventilation
| 11 | Intermediate open space to allow the introduction of over-bed table
| 12 | Bedside cabinet is easy for cleaning and desinfection
12 |  | Electro-hydraulic examination bed |  | 20 |  | 0
| 1 | Bed base made of non-ferrous metal
| 2 | Two-piece, electrically adjustable height using the control unit
| 3 | Adjustable height from 60-90cm
| 4 | The back is adjusted hydraulically by means of levers located on the sides of the bed
| 5 | Back and seat part made of fire resistant material, easy to wash and maintain
| 6 | Paper roll holder
| 7 | 4 wheels, with central brake
| 8 | Bed capacity: 200 Kg
| 9 | Dimensions: 190 x 65x 60/90cm ±2cm
13 |  | Mobil screen, 3-elements |  | 40 |  | 0
| 1 | Mobil screen, 3-elements
| 2 | Structure made of epoxy-coated steel
| 3 | Light blue washable and fire-resistant PVC curtains
| 4 | Central clothes hook with locking function
| 5 | Base with central feet and two pairs of wheels, Ø 50 mm, to facilitate moving
| 6 | Dimensions : 165 cm x 170 cm
14 |  | Movable telescopic infusion holder |  | 50 |  | 0
| 1 | Infusion holder base made of chromed metal
| 2 | Possibility of height adjustment
| 3 | 4 stainless steel hooks, load capacity 5 kg per hook
| 4 | 5 wheels Ø 50 mm, of which 2 wheels have brakes
15 |  | Mobile examination lamp |  | 25 |  | 0
| 1 | Color temperature 4,500 K
| 2 | Color rendering index> Ra 95
| 3 | Ilumination at 100 cm: 40,000 lx
| 4 | Field diameter at 100 cm < 24 cm
| 5 | The lamp base is stable and with wheels with a brake
| 6 | An on / off switch is present on the lamp head
| 7 | Lighthead power consumption 9 W
| 8 | Service life> 60.000 h
16 |  | Column scale |  | 8 |  | 0
| 1 | Digital Column scale with height measurement rod
| 2 | Functions: BMI, TARE, HOLD
| 3 | Weighing platform made of steel
| 4 | Measurement rod made of aluminium
| 5 | Easily movable with transport wheels
| 6 | Capacity up to 200 kg
17 |  | Chair for blood sampling |  | 2 |  | 0
| 1 | Blood sampling chair with fixed height
| 2 | Back rest, seat part and leg rest
| 3 | Adjustable backrest
| 4 | ABS covered sub-frame and outer sides of chair
| 5 | 4 wheels with brake.
| 6 | Special ergonomics arm pads for positiong of hand
| 7 | Safe working load 200 kg
| 8 | Fireproof padding and upholstery
18 |  | Hydraulic instrument table |  | 30 |  | 0
| 1 | The whole table is made of stainless steel, with 3 wheels with a diameter of Ø 50 mm
| 2 | Working plate size 700 mm x 500 mm,
| 3 | Working plate table can be moved -90 ° to + 90 °
| 4 | Possibility to adjust the height of the hydraulic table with the foot pedal from 1000 mm to 1300 mm.
| 5 | The table is mobile on 3 wheels, with the possibility of locking the wheels.
| 6 | Table load maximum 30 kg.
19 |  | Wheelchair |  | 10 |  | 0
| 1 | Seat chair for patient transport
| 2 | Detachable arms
| 3 | Seat width 40 cm±5%
| 4 | Width of whole chair 50cm ±5%
| 5 | Height of seat 55 cm
| 6 | Castors with brakes of min 10cm
| 7 | Weight of chair 10kg
| 8 | Load of 100kg