## Price Schedule
Bidder: |  |  |  | Offered price (currency): | 0
LOT 3 | STERILIZATION
ID | Equipment name (Technical Specification Requested) | QTY | Technical Specification Offered (model) | Unit price | Total price per line item
1 | Steam Sterilizer 8STU | 3 |  |  | 0
2 | Steam Sterilizer 6 STU | 2 |  |  | 0
3 | Washer disinfector double door | 2 |  |  | 0
4 | Washer disinfector single chamber 10 DIN | 1 |  |  | 0
5 | Ultrasonic washer with compatible basket | 5 |  |  | 0
6 | Low temperature sterilizer double door | 4 |  |  | 0
7 | Roll holders with cutting device | 4 |  |  | 0
8 | Trolley for transport sterile material | 4 |  |  | 0
9 | Height adjustable packing table | 4 |  |  | 0
10 | Linen inspection table | 1 |  |  | 0
11 | Paper trolley | 2 |  |  | 0
12 | Magnifying lamp | 4 |  |  | 0
13 | Spray gun | 2 |  |  | 0
14 | Working table with 1 sink, lower shelf and tap, 200x65x90 cm | 1 |  |  | 0
15 | Working table with 2 sinks, lower shelf and tap, 200x65x90 cm | 2 |  |  | 0
16 | Unit for washing of trolleys | 1 |  |  | 0
17 | Cutting device for sterilisation rolls, 120 cm | 4 |  |  | 0
18 | Sealing device for sterilisation rolls | 2 |  |  | 0
19 | Sterilisation container, 1 STU, 58x28x26 cm | 10 |  |  | 0
20 | Sterilisation container, 1/2 STU, 58x28x13.5 cm | 10 |  |  | 0
21 | Sterilisation container, 1/2 STU, 28.5x28x26 cm | 10 |  |  | 0
22 | Wire basket for sterilisation, 1 STU, 58x28x26 cm | 60 |  |  | 0
23 | Wire basket for sterilisation, 1/2 STU, 58x28x13.5 cm | 60 |  |  | 0
24 | Washer disinfector for substerile rooms | 2 |  |  | 0
|  |  |  | Total Price per Lot: | 0

## STERILIZATION
Bidder:
LOT 3 |  | STERILIZATION |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Steam Sterilizer 8STU |  | 3 |  | 0
| 1 | CE Marked steam sterilizer with trolleys and loading cars
| 2 | Double door unit, chamber size 550 - 600 L
| 3 | Vertical sliding pneumatic doors
| 4 | Integral electric steam generator
| 5 | Stainless steel piping
| 6 | Fully jacketed chamber
| 7 | Chamber material AISI 316L
| 8 | Large touch screens on both side (10" or larger)
| 9 | Integrated thermal printer
| 10 | Door obstruction sensor for user safety
| 11 | Doubles sensors for temperature and for pressure monitoring
| 12 | Hydrophobic air filter 0,2um
| 13 | Safety stop button on both sides
| 14 | Ethernet connection
| 15 | Square design of chamber without sharp angles for easy cleaning
| 16 | Dimensions not larger then: height 205+/-5cm, width 100+/-5cm, cm, depth 200cm
| 17 | Applied standards: ISO 9001, ISO 13485, EN 285:2015, EN 61010, 93/42/EEC, 2014/68/EU,
| 18 | Power supply 400V, 50Hz
| 19 | Test cycles: Bowie-Dick test, Leak test
| 20 | Pre-programmed warm up cycle to prepare chamber and jacket
| 21 | Standard working temperature for sterilization 121C and 134C
| 22 | Internal racks for 8 STU chamber
| 23 | Loading carts - 2pcs
| 24 | Service and maintenance from front and side
2 |  | Steam Sterilizer 6 STU |  | 2 |  | 0
| 1 | CE Marked steam sterilizer with trolleys and loading cars
| 2 | Singe door unit, chamber size approx 450 L
| 3 | Vertical sliding pneumatic doors
| 4 | Integral electric steam generator
| 5 | Stainless steel piping
| 6 | Fully jacketed chamber
| 7 | Chamber material AISI 316L
| 8 | Large touch screens on both side (10" or larger)
| 9 | Integrated thermal printer
| 10 | Door obstruction sensor for user safety
| 11 | Doubles sensors for temperature and for pressure monitoring
| 12 | Hydrophobic air filter 0,2um
| 13 | Safety stop button on both sides
| 14 | Ethernet connection
| 15 | Square design of chamber without sharp angles for easy cleaning
| 16 | Dimensions not larger then: hight 205cm+/-5cm, width 105cm+/-5cm, depth 200cm
| 17 | Applied standards: ISO 9001, ISO 13485, EN 285:2015, EN 61010, 93/42/EEC, 2014/68/EU,
| 18 | Power supply 400V, 50Hz
| 19 | Test cycles: Bowie-Dick test, Leak test
| 20 | Preprogrammed warm up cycle to prepare chamber and jacket
| 21 | Standard working temperature for sterilization 121C and 134C
| 22 | Internal racks for 6 STU chamber
| 23 | Loading carts - 1 pc
| 24 | Service and maintenance from front and side
3 |  | Washer disinfector double door |  | 2 |  | 0
| 1 | Washer disinfector with pneumatic vertical sliding doors, double door configuration
| 2 | Washing chamber, washing arms, pre-heating tanks and water filters made of high quality stainless steel AISI 316 L
| 3 | Frame and external cabinet made in stainless steel AISI 304
| 4 | Doors made from tempered glass
| 5 | Control panels are touch screen displays on both sides of washer, 20cm of larger
| 6 | 8 factory loaded and more then 18 custom cycles
| 7 | Supervisor can adjust parameters within each cycle ( number of prewashing or washing phases..)
| 8 | Fast cycles: prewashing, ensimatic washing, washing, rinsing and thermal desinfection and drying cycle time reduced to 30 minutes or shorter
| 9 | 2 washing chamber spray arms for efficient water and air distribution inside the chamber, top and bottom
| 10 | Automatic connection of racks to bottom spray manifold
| 11 | LED light in washing chamber
| 12 | Electric heating coil
| 13 | Removable stainless steel filter in a chamber sump for residuals
| 14 | Water level sensor in a sump
| 15 | Supply connections from the top (electricity, water, air)
| 16 | Doors with obstruction sensors that protect user. Only one door can be opened at the moment
| 17 | Maximum dimesions: width 110cm, dept 110cm, height 210cm
| 18 | Applied standards: ISO 15883-1:2006, 93/42/EEC, 2007/47/EC, 2006/42/EC, IEC 61010-1 2010, EN 61326-1 2013, IEC 61010-2-040:2005
| 19 | Number of pumps - minimum 5
| 20 | Pumps are positioned out of washer, closer to detergent containers
| 21 | USB port on the front panel to download historical cycle data and for software upgrade
| 22 | Chamber volume ~360 lit
| 23 | Capacity up to 15 DIN
| 24 | Safety stop button on both sides
| 25 | Audio and visual alarms
| 26 | Integrated printer
| 27 | Report memory for 250 or more cycles
| 28 | Regular and preventive maintenance reminder
| 29 | System for cooling down drainaged water to 60C or less
| 30 | Service and maintenance from front side
| 31 | Transfer carts - pcs 2
| 32 | 5 level wash rack
| 33 | 4 level wash rack
| 34 | 2 level wash rack
| 35 | 3 level wash rack for MIS instruments
| 36 | Anaesthesia/Respiratory rack
4 |  | Washer disinfector single chamber 10 DIN |  | 1 |  | 0
| 1 | Washer desinfector with pneumatic verticaly sliding doors, single door configuration
| 2 | Washing chamber, washing arms, sump made of high quality statinless steel AISI 304 or better
| 3 | Frame and external panels made of stainless steel AISI 304L or better
| 4 | Door made from tempered glass
| 5 | Control panels are touch screen displays on both side of washers
| 6 | 8 factory loaded programs and more then 5 custom cycles
| 7 | Supervisor can adjust parameters within each cycle
| 8 | Cycle length (prewashing, ensimatic washing, washing, rinsing and thermal desinfection and drying) 30 min or faster (190 or more trays per 8hrs shift)
| 9 | 2 washing chamber spray arms for efficient water and air distribution inside the chamber, top and bottom
| 10 | Automatic connection of racks to bottom spray manifold
| 11 | LED light in washing chamber
| 12 | Water consumption up to 145 l per cycle
| 13 | Electric heating coil
| 14 | Removable stainless steel filter in a sump
| 15 | Supply connections from the top of the machine, except the drain (electric, steam, compressed air, water)
| 16 | Door with obstruction sensors that protect user.
| 17 | Water level safety sensor for chamber sump water level
| 18 | Maximum dimensions :width 110cm, depth 80cm, height 210cm
| 19 | Applied standards ISO 15883-1:2006, ISO 15883-2:2006, IEC 61010-1 2010, IEC 61326-1:2005, IEC 61010-2-040:2005, 93/42/EEC, 2006/42/EC
| 20 | Number of pumps - 5 minimum
| 21 | Pumps are positioned out of washer, closer to detergent to department with containers, up to 30m distance far
| 22 | USB port on the front panel to download historical data and software upgrade
| 23 | Chamber volume ~ 300l
| 24 | Capacity up to 10 DIN
| 25 | Safety stop button on both sides
| 26 | Audio and visual alarms
| 27 | Integrated printer
| 28 | Control system back-up all cycle memory for up to 10 years
| 29 | Regular and preventive maintenance reminder
| 30 | System for cooling down drainaged water to 60C or lower
| 31 | Service and maintenance from front side
| 32 | Power supply 400V , 50Hz
| 33 | Transfer carts - pc 1
| 34 | 5 level wash rack
| 35 | 4 level wash rack
| 36 | 3 level wash rack
| 37 | 2 level wash rack
5 |  | Ultrasonic washer with compatible basket |  | 5 |  | 0
| 1 | Total tank capacity from 35 to 45 L
| 2 | Automatic fll & detergent process
| 3 | Fast cycles
| 4 | Lumen flash for 10 or more cannulated instruments
| 5 | Weight less then 40kg
| 6 | Electric power less than 2,5kW
| 7 | Electrical supply 50Hz, 230V
| 8 | Touch screen control pannel
| 9 | USB port connection
6 |  | Low temperature sterilizer double door |  | 4 |  | 0
| 1 | Low temperature sterilizer double door - pass through
| 2 | Low-temperature hydrogen peroxide sterilizer designed for sterilisation thermo-sensitive instruments and materials
| 3 | Chamber is made of high grade AISI 316L stainless steel or aluminum
| 4 | Chamber volume is 110-145l
| 5 | Useful volume 95% minimum
| 6 | Rectangular chamber without sharp angles for easy cleaning
| 7 | Two shelves in chamber
| 8 | H2O2 is packed inside of cartridges - for minimum 15 cycles
| 9 | Chamber H202 concentration maximum 59%
| 10 | Lenght of the cycles - from 16 to 55 min
| 11 | Electrical Power Supply: 220V/~/50Hz
| 12 | Sterilization temperature up to 52C
| 13 | Sterilizer has casters for repositioning
| 14 | Touch screen control display on both sides, 10" or larger
| 15 | Integrated printer
| 16 | Audio and visual alarms
| 17 | Ethernet connection
| 18 | Power suply 50Hz, 400V
| 19 | Softwear for sterilisation process tracking
| 20 | Biological Incubator
7 |  | Roll holders with cutting device |  | 4 |  | 0
| 1 | Housing made of stainless steel AISI 304 except knife and roller on which rolls slide
| 2 | Desk model that can be attached to the wall
| 3 | Capacity (roll width) min. 900 mm
8 |  | Trolley for transport sterile material |  | 4 |  | 0
| 1 | Trolley for transport of sterile containers and baskets for 6 sterilization unit
| 2 | 2 pushing handles
| 3 | Two doors on hinges, lockable
| 4 | Door seal around the whole door and frame
9 |  | Height adjustable packing table |  | 4 |  | 0
| 1 | Table for packing sterilization materials
| 2 | Electrically adjustable tabletop height by push buttons
| 3 | Possibility of presetting tabletop heights
| 4 | Tabletop made of AISI 304 stainless steel or better
| 5 | Columns for shelves made of stainless steel or aluminum
| 6 | One drawer on one side under the worktop
| 7 | 2 shelves above the table top, placed between vertical holders
| 8 | Power supply 220V-50Hz
| 9 | Dimensions min. 1800x700x800-1200 mm (WxWxH)
| 6 | ALIMS Registration is not necessary
10 |  | Linen inspection table |  | 1 |  | 0
| 1 | Work table with all-clear illumintation part for easy inspection of linen.
| 2 | The working table for textile control should have neon or LED lighting, for an uninterrupted process of control
| 3 | The worktop made of HPL or better.
| 4 | Frame and shelf made of AISI304 stainless steel or better.
| 5 | Dimensions min. 2000x1200x900 mm (WxDxH)
| 6 | ALIMS Registration is not necessary
11 |  | Paper trolley |  | 2 |  | 0
| 1 | Made of stainless steel AISI 304 or better
| 2 | Min. 4 frames for paper holding with camps for holding paper in place
| 3 | Solid bottom shelf made of stainless steel AISI304 or better
| 4 | 4 castors, 2 of which have brakes
| 6 | ALIMS Registration is not necessary
12 |  | Magnifying lamp |  | 4 |  | 0
| 1 | Stain free quality magnifiers
| 2 | Effortless adjustment without swinging movements
| 3 | Strong shadow free light
| 4 | Lighting with fluorescent lamp or LED, max 30W
| 5 | Distortion-free
| 6 | 4-diopter magnifier min. Ø120 mm
| 7 | Two arms of total length min 800mm
| 8 | Adjustment angle arms: 140°/130°, Rotation 360°
| 9 | El. power 230V/50Hz
| 10 | ALIMS Registration is not necessary
13 |  | Spray gun |  | 2 |  | 0
| 1 | The device is intended for washing and cleaning, with the possibility of connection to hot water or compressed air (the same device has both possibilities)
| 2 | Adjustable water or air flow
| 3 | Min. 8 attachments for washing different elements
| 4 | Wall-moounted holder for attachments
| 5 | Connection hose – min. 2 m
| 6 | Spray gun made of stainless steel, aluminum, and similar
| 7 | ALIMS Registration is not necessary
14 |  | Working table with 1 sink, lower shelf and tap, 200x65x90 cm |  | 1 |  | 0
| 1 | material: stainless steel, AISI 304
| 2 | 1 sink, dimension 50x40x20 cm
| 3 | combined tap for cold/ hot water
| 4 | lower shelf at 30 cm height from floor
| 5 | dimensions 200x65x90 cm (+/- 5 cm)
| 6 | ALIMS Registration is not necessary
15 |  | Working table with 2 sinks, lower shelf and tap, 200x65x90 cm |  | 2 |  | 0
| 1 | material: stainless steel, AISI 304
| 2 | 2 sinks, dimension 50x40x20 cm
| 3 | combined tap for cold/ hot water
| 4 | lower shelf at 30 cm height from floor
| 5 | dimensions 200x65x90 cm (+/- 5 cm)
| 6 | ALIMS Registration is not necessary
16 |  | Unit for washing of trolleys |  | 1 |  | 0
| 1 | unit for washing of trolleys and other equipment
| 2 | tank for detergents
| 3 | dosing system for detergents
| 4 | hose, min. 10 m length
| 5 | adjustable nozzle
17 |  | Cutting device for sterilisation rolls |  | 4 |  | 0
| 1 | material: stainless steel, AISI 304
| 2 | for storing and cutting of sterilisation rolls
| 3 | 4 carrying and positioning bars for rolls
| 4 | dimensions 120x35x20 cm (+/- 5 cm)
| 5 | ALIMS Registration is not necessary
18 |  | Sealing device for sterilisation rolls |  | 2 |  | 0
| 1 | Microprocessor controlled automatic sealing device to be used for sealing of sterilisation reels
| 2 | possibility to validate sealing machine
| 3 | control of critical parameters according to ISO 11607-2
| 4 | equipped with Ethernet and USB connections
| 5 | possibility to connect device to external printer
| 6 | user interface by means of a colour touch-screen
| 7 | possibility to record process parameters with 3 programs for quick access
| 8 | automatic stand-by, when not in use
| 9 | safety system for overheating
| 10 | working temperature: 100-200*C
| 11 | electrical connection 220 V, 50 Hz
| 12 | dimensions max. 600x300x200 mm
| 13 | Standards: ISO 13485, LV 2006/95/EC, EMC 2004/108/EC
19 |  | Sterilisation container, 1 STU, 58x28x26 cm |  | 10 |  | 0
| 1 | for storing and sterilisation of sterile and non-sterile material
| 2 | consisting of bottom and lid, perforated or non-perforated
| 3 | material aluminum
| 4 | with sterile filter
| 5 | with silicone gasket between bottom and lid
| 6 | with handles
| 7 | dimensions 580 (+/-15) x 280 (+/-15) x 260 (+/-15) mm
20 |  | Sterilisation container, 1/2 STU, 58x28x13.5 cm |  | 10 |  | 0
| 1 | for storing and sterilisation of sterile and non-sterile material
| 2 | consisting of bottom and lid, perforated or non-perforated
| 3 | material aluminum
| 4 | with sterile filter
| 5 | with silicone gasket between bottom and lid
| 6 | with handles
| 7 | dimensions 580 (+/-15) x 280 (+/-15) x 135 (+/-15) mm
21 |  | Sterilisation container, 1/2 STU, 28.5x28x26 cm |  | 10 |  | 0
| 1 | for storing and sterilisation of sterile and non-sterile material
| 2 | consisting of bottom and lid, perforated or non-perforated
| 3 | material aluminum
| 4 | with sterile filter
| 5 | with silicone gasket between bottom and lid
| 6 | with handles
| 7 | dimensions 285 (+/-15) x 280 (+/-15) x 260 (+/-15) mm
22 |  | Wire basket for sterilisation, 1 STU, 58x28x26 cm |  | 60 |  | 0
| 1 | Baskets are intended for storage and transportation of non-sterile and sterile material
| 2 | Baskets are made of wire mesh;raster prevents standard elements from falling out.
| 3 | Square shape
| 4 | Made of stainless steel AISI 304
| 5 | Type: 1 STU,
| 6 | Dimensions: approximately 600x300x300 mm (LxWxH)
23 |  | Wire basket for sterilisation, 1/2 STU, 58x28x13.5 cm |  | 60 |  | 0
| 1 | Baskets are intended for storage and transportation of non-sterile and sterile material
| 2 | Baskets are made of wire mesh;raster prevents standard elements from falling out.
| 3 | Square shape
| 4 | Made of stainless steel AISI 304
| 5 | Type: 1/2STU
24 |  | Washer disinfector for substerile rooms |  | 2 |  | 0
| 1 | Washing machine for washing and disinfection of medical elements
| 2 | It us used for washing and disinfection of medical elements at temperatures of 90°C
| 3 | Floor model, with PLC controller
| 4 | Pump with flow of min. 350 lit/min.
| 5 | Draining pump with min. 40 lit/min.
| 6 | Heating by electrical heaters
| 7 | 2 pumps for washing chemical detergents
| 8 | Final rinse with high purified water
| 9 | Thermal disinfection above 90*C
| 10 | HEPA filter for drying
| 11 | Chamber with 2 or 3 levels for washing, 4 DIN baskets (480x250x50 mm).
| 12 | Chamber volume: 140-180 lit.
| 13 | Model with 1 door
| 14 | Sprinklers at the top and bottom of the chamber
| 15 | Complies with EN 15883
| 16 | With accessories:
| 17 | Standard 2- level trolley for 4 DIN baskets
| 18 | Anesthesia trolley
| 19 | 4 DIN baskets
| 20 | Dimensions: 60x40x90 (140) cm
| 21 | Chamber material: AISI 316L
| 22 | Door material: thermostable glass door