## Price Schedule
Bidder: |  |  |  | Offered price: | 0
LOT 8 | DIAGNOSTIC
ID | Equipment name | QTY | Technical Specification Offered (model) | Estimated value per unit | Estimated value total
1 | Transcranial doppler device | 1 |  |  | 0
2 | Device for the diagnosis of peripheral vascular disease | 1 |  |  | 0
3 | ECG Device 1 with trolley | 27 |  |  | 0
4 | ECG device - tablet with trolley | 3 |  |  | 0
5 | Defibrillator for hospital use - advanced | 6 |  |  | 0
6 | Defibrillator for hospital use - standard | 10 |  |  | 0
7 | Defibrillator for hospital use with internal defibrillation | 4 |  |  | 0
8 | Cardio-pulmonary exercise testing | 1 |  |  | 0
9 | ECG and Blood pressure holter system | 1 |  |  | 0
10 | PC based spirometry device with ultrasonic sensor | 1 |  |  | 0
11 | Aneroid Blood Pressure Measurement Device | 40 |  |  | 0
12 | Stetoscope | 70 |  |  | 0
13 | BLE Pulse oxymeter | 30 |  |  | 0
|  |  |  | Total Price per Lot: | 0

## Diagnostic
Bidder:
LOT 8 |  | DIAGNOSTIC |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Transcranial doppler device |  | 1 |  | 0
| 1 | Bilateral 2 channel device
| 2 | M Mode and Power M mode
| 3 | Software for bilateral monitoring
| 4 | Software for emboli detection and vasomotoric reactivity
| 5 | Software for automatic tracking of doppler signal and depth selection
| 6 | Multigating
| 7 | Probes:
| 8 | 2Mhz (PW), 4Mhz (CW/PW), 8Mhz (CW/PW),
| 9 | Headset with fixation for two 2Mhz probes for functional patient monitoring
| 10 | Recording and playback of recorded spectar, M mode and audio
| 11 | Integrated touchscreen monitor min 15 inch
| 12 | Integrated HDD, USB and LAN, Dicom connectivity
2 |  | Device for the diagnosis of peripheral vascular disease |  | 1 |  | 0
| 1 | Must have following physiological measurements to diagnose peripheral arterial disease (PAD):
| 2 | 4 & 8 MHz Doppler
| 3 | VRT
| 4 | Bilateral PPG
| 5 | Bilateral PVR
| 6 | Doppler pressure
| 7 | Simultaneous PPG pressures - 4
| 8 | Doppler ABI
| 9 | PPG ABI-4
| 10 | PPG TBI-4
| 11 | PWV-4
| 12 | Must have following components:
| 13 | Leg/arm cuffs, Latex-free
| 14 | 4 and 8 MHz CW Doppler probes
| 15 | PPG sensors for all measured pressures
| 16 | Dedicated system trolley with facilities for the storage of the cuffs, sensors, probes and printer
| 17 | Device must be offered with color laser printer for printing of full-page colour reports
| 18 | Possibility for export reports in PDF mode
| 19 | Embedded PC with 7 inch display
3 |  | ECG Device 1 with trolley |  | 27 |  | 0
| 1 | Must support simultaneous 12-channel recordings in standard, right precordial, left posterior, standard C4r and Nehb lead configurations
| 2 | Must have an integrated color display, diagonal min. 8", resolution min. 1024x768
| 3 | Must display all 12-leads on the integrated display
| 4 | Must have an integrated, sealed, rubber or silicone, alphanumeric keyboard with function keys
| 5 | Must have an automatic and manual mode
| 6 | Must have manually selectable QRS amplitude signal sensitivity 5-10-20 mm/mV
| 7 | Must have an ECG sampling frequency min. 30 kHz
| 8 | Must have a frequency bandwidth range min. 0 - 250 Hz
| 9 | Must have an integrated pacemaker signal detection function with a signal detection frequency > 20 kHz
| 10 | Must be protected from defibrillation, rated to at least 5000 VDC
| 11 | Must have an integrated high resolution thermal printer for Z-folded paper, width min. 210 mm
| 12 | Must support printing 12-channel ECG recordings with printing speeds 12.5 - 25 - 50 mm/s adjustable in manual mode
| 13 | Must have an integrated print preview function
| 14 | Must support exporting ECG recordings in PDF format to a USB stick
| 15 | Must have lead quality check (color coded or equivalent) and lead reversal detection functionality
| 16 | Must have an integrated power supply
| 17 | Must have an integrated battery supply with min. 7 h operating time
| 18 | Must have an integrated software for ECG signal interpretation
| 19 | Must have an integrated software for ECG signal interpretation for athletes according to "Seattle criteria"
| 20 | Must support QT correction formulas at least Bazett, Fredericia, Fremingham and Hodges
| 21 | Must be upgradable with software for locating obstructed coronary arteries, risk assessment and therapy recommendations for acute chest pain patients
| 22 | Must have an integrated resting ECG recording function with at least 5 min of continuous recording
| 23 | Must have an integrated memory for at least 350 ECG recordings
| 24 | Must have support for bi-directional communication with hospital information systems using DICOM or XML standards over integrated communication module(s)
| 25 | Must have integrated communication modules, at least WiFi and LAN
| 26 | Must have the complete software translated in Serbian
| 27 | Must be delivered with appropriate standard operating accessories and an ECG trolley with a basket for necessary accessories
4 |  | ECG device - tablet with trolley |  | 3 |  | 0
| 1 | Must support simultaneous 12-channel recordings in standard, right precordial, left posterior, standard C4r and Nehb lead configurations
| 2 | Must have an integrated multi-touch color display, diagonal min. 8", resolution min. 1024x768
| 3 | Must display all 12-leads on the integrated display
| 4 | Must enable reviewing of ECG recordings on the integrated display at any time
| 5 | Must enable saving, transmitting and printing of ECG recordings at any time
| 6 | Must have an automatic and manual mode
| 7 | Must have manually selectable QRS amplitude signal sensitivity 5-10-20 mm/mV
| 8 | Must have an ECG sampling frequency min. 30 kHz
| 9 | Must have a frequency bandwidth range min. 0 - 250 Hz
| 10 | Must have an integrated pacemaker signal detection function with a signal detection frequency > 20 kHz
| 11 | Must be protected from defibrillation, rated to at least 5000 VDC
| 12 | Must have an integrated high resolution thermal printer for Z-folded paper, width min. 110 mm
| 13 | Must support printing 6-channel ECG recordings with printing speeds 12.5 - 25 - 50 mm/s adjustable in manual mode
| 14 | Must have an integrated print preview function
| 15 | Must support exporting ECG recordings in PDF format to a USB stick
| 16 | Must have lead quality check (color coded or equivalent) and lead reversal detection functionality
| 17 | Must have an integrated battery supply with min. 4 h operating time
| 18 | Must have an integrated software for ECG signal interpretation
| 19 | Must have an integrated software for ECG signal interpretation for athletes according to "Seattle criteria"
| 20 | Must support QT correction formulas at least Bazett, Fredericia, Fremingham and Hodges
| 21 | Must be upgradable with software for locating obstructed coronary arteries, risk assessment and therapy recommendations for acute chest pain patients
| 22 | Must have an integrated resting ECG recording function with at least 3 min of continuous recording
| 23 | Must have an integrated memory for at least 350 ECG recordings
| 24 | Must have support for bi-directional communication with hospital information systems using DICOM or XML standards over integrated communication module(s)
| 25 | Must have integrated communication modules, at least WiFi and LAN
| 26 | Must have the complete software translated in Serbian
| 27 | Must be delivered with appropriate standard operating accessories and an ECG trolley with a basket for necessary accessories
5 |  | Defibrillator for hospital use - advanced |  | 6 |  | 0
| 1 | Must have an integrated touch color display, diagonal min. 7", readable in all conditions, protected by tempered glass
| 2 | Must have multiple user-selectable display configurations containing at least 6-channel ECG signal with lead-off detection, current heart rate (HR), vital parameters and number of delivered shocks
| 3 | Must have an integrated power supply
| 4 | Must have an integrated battery supply with a capacity of min. 250 shocks with max. energy or min. 8 h of ECG monitoring
| 5 | Must have an indicator to display the charging-discharging state of the battery
| 6 | Must support bi-phasic defibrillation with max. defibrillation energy 200 J +/- 5% selectable in steps of 1 J ranging from 1 J to 15 J
| 7 | Must have paddles for defibrillation with energy level selection, and charge, shock and print buttons
| 8 | Must have a patient's impedance check function selector and indicator
| 9 | Must support screenshot function in all operating modes
| 10 | Must have an integrated thermal printer, paper width min. 80 mm
| 11 | Must support printing 3-channel ECG recordings
| 12 | Must have an integrated memory for at least 24 h of ECG recordings, patient and defibrillation data, impedance curves, registered events, CPR feedback and screenshots
| 13 | Must have integrated external pacemaker funcionality with Demand and Fixed modes
| 14 | Must have integrated pulse-oximetry (SpO2) measurement functionality over Masimo or equivalent sensor
| 15 | Must have integrated non-invasive blood pressure (NIBP) measurement functionality
| 16 | Must be upgradable with EtCO2 measurement functionality
| 17 | Must be upgradable with internal defibrillation functionality
| 18 | Must have integrated support for connecting reusable CPR sensors with frequency, depth and chest compression feedback support
| 19 | Must support automatic self-testing functionality with report printing and automatic submission of preformed self-tests to the information sistem over integrated communication module(s)
| 20 | Must have integrated communication modules, at least WiFi and USB
| 21 | Must have a gross weight (including battery and paddles) of 6 kg +/- 1 kg
| 22 | Must be delivered with appropriate standard operating accessories including paddles for defibrillation, ECG patient cable, SpO2 kit, NIBP kit, defibrillation gel, adult pacing/defibrillation pads
6 |  | Defibrillator for hospital use - standard |  | 10 |  | 0
| 1 | Must have an integrated touch color display, diagonal min. 7", readable in all conditions, protected by tempered glass
| 2 | Must have multiple user-selectable display configurations containing at least 6-channel ECG signal with lead-off detection, current heart rate (HR), vital parameters and number of delivered shocks
| 3 | Must have an integrated power supply
| 4 | Must have an integrated battery supply with a capacity of min. 250 shocks with max. energy or min. 8 h of ECG monitoring
| 5 | Must have an indicator to display the charging-discharging state of the battery
| 6 | Must support bi-phasic defibrillation with max. defibrillation energy 200 J +/- 5%
| 7 | Must have paddles for defibrillation with energy level selection, and charge, shock and print buttons
| 8 | Must have a patient's impedance check function selector and indicator
| 9 | Must support screenshot function in all operating modes
| 10 | Must have an integrated thermal printer, paper width min. 80 mm
| 11 | Must support printing 3-channel ECG recordings
| 12 | Must have an integrated memory for at least 24 h of ECG recordings, patient and defibrillation data, impedance curves, registered events, CPR feedback and screenshots
| 13 | Must have integrated external pacemaker funcionality with Demand and Fixed modes
| 14 | Must be upgradable with pulse-oximetry (SpO2), non-invasive blood pressure (NIBP) and EtCO2 measurement functionality
| 15 | Must have integrated support for connecting reusable CPR sensors with frequency, depth and chest compression feedback support
| 16 | Must support automatic self-testing functionality with report printing and automatic submission of preformed self-tests to the information sistem over integrated communication module(s)
| 17 | Must have integrated communication modules, at least WiFi and USB
| 18 | Must have a gross weight (including battery and paddles) of 6 kg +/- 1 kg
| 19 | Must be delivered with appropriate standard operating accessories including paddles for defibrillation, ECG patient cable, defibrillation gel, adult pacing/defibrillation pads
7 |  | Defibrillator for hospital use with internal defibrillation |  | 4 |  | 0
| 1 | Must have an integrated touch color display, diagonal min. 7", readable in all conditions, protected by tempered glass
| 2 | Must have multiple user-selectable display configurations containing at least 6-channel ECG signal with lead-off detection, current heart rate (HR), vital parameters and number of delivered shocks
| 3 | Must have an integrated power supply
| 4 | Must have an integrated battery supply with a capacity of min. 250 shocks with max. energy or min. 8 h of ECG monitoring
| 5 | Must have an indicator to display the charging-discharging state of the battery
| 6 | Must support bi-phasic defibrillation with max. defibrillation energy 200 J +/- 5% selectable in steps of 1 J ranging from 1 J to 15 J
| 7 | Must have paddles for defibrillation with energy level selection, and charge, shock and print buttons
| 8 | Must have a patient's impedance check function selector and indicator
| 9 | Must support screenshot function in all operating modes
| 10 | Must have an integrated thermal printer, paper width min. 80 mm
| 11 | Must support printing 3-channel ECG recordings
| 12 | Must have an integrated memory for at least 24 h of ECG recordings, patient and defibrillation data, impedance curves, registered events, CPR feedback and screenshots
| 13 | Must have an integrated internal defibrillation mode with appropriate electrodes
| 14 | Must have integrated external pacemaker funcionality with Demand and Fixed modes
| 15 | Must be upgradable with pulse-oximetry (SpO2), non-invasive blood pressure (NIBP) and EtCO2 measurement functionality
| 16 | Must have integrated support for connecting reusable CPR sensors with frequency, depth and chest compression feedback support
| 17 | Must support automatic self-testing functionality with report printing and automatic submission of preformed self-tests to the information sistem over integrated communication module(s)
| 18 | Must have integrated communication modules, at least WiFi and USB
| 19 | Must have a gross weight (including battery and paddles) of 6 kg +/- 1 kg
| 20 | Must be delivered with appropriate standard operating accessories including paddles for defibrillation, 2x adult spoon electrodes for internal defibrillation with connecting cable, ECG patient cable, defibrillation gel, adult pacing/defibrillation pads
8 |  | Cardio-pulmonary exercise testing |  | 1 |  | 0
| 1 | Cardio-pulmonary diagnostic station
| 2 | Must be able to record and edit 12-channel resting and stress ECG recordings in standard, Nehb, Cabrera, Frank, right precordial, left posterior lead configurations
| 3 | Must have an alphanumeric keyboard
| 4 | Must be able to display 12-channel ECG recordings
| 5 | Must have a sampling rate at least 8 kHz
| 6 | Must have a sampling resolution at least 24 bit
| 7 | Must have a frequency range at least 0.05 - 300 Hz
| 8 | Must have a condensation insensitive pneumotacho sensor
| 9 | Must have an electro-chemical O2 analyser
| 10 | Must have an ultrasonic CO2 analyser
| 11 | Must have an integrated software for ECG signal interpretation
| 12 | Must have an integrated software for ECG signal interpretation for athletes according to "Seattle criteria"
| 13 | Must have an integrated software for high-frequency QRS analysis to detect depolarization abnormalities
| 14 | Must have an online oxygen consumption analysis method in stress (breath-by-breath or equivalent)
| 15 | Must be upgradable with 3D vector cardiography software
| 16 | Must have the ability to compare ECG strips
| 17 | Must be able to display previous ECG recordings during a test
| 18 | Must be able to simultaneously display cardio-pulmonary parameters and graphics during exercise testing
| 19 | Must be able to automatically detect AT and RCP
| 20 | Must be able to display animations for better cooperation with patients
| 21 | Must have integrated communication interfaces, at least LAN and USB
| 22 | Must have QRS trigger signal output
| 23 | Must be based on Microsoft Windows
| 24 | Must have two full HD color monitors with a diagonal min. 23" in a side-by-side configuration
| 25 | Must have an integrated thermal printer for A4 thermal paper
| 26 | Must be able to print 12-channel ECG recordings
| 27 | Must be compatible with external laser printers
| 28 | Must have an integrated rechargeable battery for uninterupted functioning
| 29 | Must be delivered with a color laser printer, ergo-spiro masks (XS, S, M, L) with hairnets, ruler for patient mask size measurement, reduction valve, calibration gas bottle, trolley with external printer shelf, isolated transformer and drawers for accessories
| 30 | Treadmil
| 31 | Must enable constant load for patient
| 32 | Must fullfill medical device risk class classification min. IIb
| 33 | Must be compatible with the cardio-pulmonary diagnostic station
| 34 | Must have a non-slip belt surface with dimensions min. 150 x 50 cm
| 35 | Must support a speed range min. 0 - 22 km/h
| 36 | Must support a slope range min. 0 - 25%
| 37 | Must permit a maximum patient weight of at least 200 kg
| 38 | Bycicle for stress echo ergometry
| 39 | Must fullfill medical device risk class classification min. IIa
| 40 | Must contain an integrated automatic blood pressure measurement system
| 41 | Must have an access port for ultrasound cardiac transducers
| 42 | Must be horizontally and laterally adjustable, at least 0 - 45 degrees
| 43 | Must be height adjustable with separate head rest position adjustment
| 44 | Must be equiped with a remote control for patient adjustment
| 45 | Must enable a load range of at least 20 - 800 W
| 46 | Must permit a maximum patient weight of at least 160 kg
| 47 | Must have an integrated graphical display
| 48 | Must be compatible with the cardio-pulmonary diagnostic station
| 49 | Must be delivered with appropriate standard operating accessories including a cable for connection to the cardiology diagnostic workstation
9 |  | ECG and Blood pressure holter system |  | 1 |  | 0
| 1 | Must be delivered with a PC system with associated software, high resolution ECG holter recorder and ambulatory blood pressure recorder
| 4 | Holter analysis software |  | 5
| 5 | Must include software modules for high resolution ECG and blood pressure holter analysis
| 6 | Must have a common database for ECG and blood pressure holter recordings
| 7 | Must support user-defined screen layout, workflow and report configuration
| 8 | High resolution ECG holter analysis module |  | 3
| 9 | Must have at least 99% beat detection accuracy
| 10 | Must have an integrated software for ST segment analysis
| 11 | Must have an integrated software for QT segment analysis with QT detection formulas, at least Bazzet, Fridericia, Pfeufer and Sagie
| 12 | Must have an integrated software for pacemaker signal detection
| 14 | Must have an integrated software for advanced heart rate variability analysis with time and frequency domain and graphical result representation for 24 h recordings
| 15 | Must have an integrated software for atrium abnormality detection based on the true P-wave detection
| 16 | Must have an integrated software for apnoea screening with sinchronised presentation of SpO2 and ECG signal/waveform and values
| 17 | Blood pressure holter analysis module |  | 2
| 18 | Must be able to display graphical and numerical representations of the measured blood pressure values (systolic, diastolic and mean)
| 19 | Must enable adding comments to recording segments or exclude certain measurements
| 20 | Must enable trend analysis and histogram display
| 21 | Must support user-defined threshold definition for custom analysis
| 22 | Must have an integrated software for masked hypertension recognition
| 23 | Must have an integrated software for white-coat hypertension recognition
| 24 | PC system |  | 5
| 25 | Must be based on Microsoft Windows
| 26 | Must have two full HD color monitors with a diagonal min. 21"
| 27 | Must include a laser printer
| 28 | High resolution ECG holter recorder (3-channel) |  | 25
| 29 | Must support high resolution 3-channel ECG holter recording
| 30 | Must have an integrated dual power supply with internal rechargeable (at least 7 days of continuous recording) and external AA or AAA type batteries with a joined capacity of at least 14 days of continuous recording
| 31 | Must support recording to memory card
| 32 | Must have an integrated ECG display for at least 1-channel ECG preview
| 33 | Must have a sampling rate at least 30 kHz
| 34 | Must have a sampling resolution at least 15 bit
| 35 | Must support true P-wave detection
| 36 | Must be compatible with 5 and 7-wire patient cables with automatic cable detection
| 37 | Must have an integrated software for respiration detection for apnoea screening
| 38 | Must support SpO2 measurement and ECG waveform synchronization
| 39 | Must be able to detect patient movement and check the ECG signal quality via the PC workstation
| 40 | Must have an integrated BT wireless communication interface to connect to wireless SpO2 sensors
| 41 | Must have an integrated voice recording function
| 42 | Must have a gross weight without external batteries of at most 130 g
| 43 | Must be delivered with appropriate standard operating accessories
| 44 | High resolution ECG holter recorder (12-channel) with SpO2 sensor |  | 5
| 45 | Must support high resolution 12-channel ECG holter recording capable of continuously recording for at least 72 h
| 46 | Must have an integrated color ECG display for at least 3-channel ECG preview
| 47 | Must have a sampling rate at least 8 kHz
| 48 | Must have a sampling resolution at least 12 bit
| 49 | Must be able to detect pacemaker signals
| 50 | Must support true P-wave detection
| 51 | Must be able to automatically test the correctness of the connected patient cable
| 52 | Must be able to detect patient movement
| 53 | Must have an integrated BT wireless communication interface to connect to wireless SpO2 sensors
| 54 | Must have an integrated voice recording function
| 55 | Must have an integrated event registration button
| 56 | Must use at most 1 AA or AAA type battery
| 57 | Must have a weight of at most 200 g
| 58 | Must be compatible with the holter analysis software
| 59 | Must be delivered with appropriate standard operating accessories including a 10-wire ECG patient cable and compatible SpO2 sensor kit with wireless connectivity
| 60 | Ambulatory blood pressure holter recorder |  | 15
| 61 | Must have an integrated color display
| 62 | Must support user-selectable recording duration, 24 h and 48 h
| 63 | Must support auscultatory and oscillometric measurement methods
| 64 | Must have a pressure measurement range of at least 30 - 300 mmHg
| 65 | Must support user-defined time intervals for blood pressure measurement with a range of at least 5 - 120 minutes
| 66 | Must support automatic and user-adjustable cuff deflation speeds douring measurements with a range of at least 2 - 9 mmHg/s
| 67 | Must support user-defined pressure measurement protocols
| 68 | Must have an integrated voice recording function with at least 30 seconds of voice recording
| 69 | Must use rechargeable AA or AAA type batteries
| 70 | Must be delivered with appropriate standard operating accessories
10 |  | PC based spirometry device with ultrasonic sensor |  | 1 |  | 0
| 1 | Must support the following lung function tests:
| 2 | Slow spirometry
| 3 | Forced spirometry (flow/volume)
| 4 | Bronchodilator monitoring/testing
| 5 | Must support flow and volume measurement and should be equipped with an ultrasonic sensor, without the need for (re)calibration
| 6 | Must have a reusable ultrasonic sensor resistant to humidity
| 7 | Must have a sampling rate min. 2000 Hz
| 8 | Must have a volume measurement range min. 0 - 20 L
| 9 | Must have a volume measurement accuracy max. +/- 2%
| 10 | Must have a flow measurement range min. 0 - 20 L/s
| 11 | Must have a flow measurement accuracy max. +/- 2%
| 12 | Must be equipped for on-line temperature measurement and BTPS factor correction during measurement
| 13 | The PC software must support the following:
| 14 | Most recent Microsoft Windows operating system
| 15 | Conducting complete spirometry tests
| 16 | Display of the spirometry and flow/volume curve
| 17 | Automatic measurement and display of measured units
| 18 | Patient database with examination history
| 19 | Pre/post-tests with monitoring and comparison of previous test series
| 20 | Display of animations for better cooperation with patients (should be adapted)
| 21 | User defined reports with a wide range of graphical analyzes and 3D diagrams
| 22 | Must be delivered with an appropriate PC system with operating system, laser printer, spirometry software, standard accessory set, 1 pc. reusable ultrasonic sensor, additional 200 pcs. single use ultrasonic sensors and 200 pcs. single use antibacterial filters
11 |  | Aneroid Blood Pressure Measurement Device |  | 40 |  | 0
| 1 | Must be for clinical use
| 2 | Must be a single tube system
| 3 | Must have a scale diameter min. 60 mm with scaling min. 0 - 300 mmHg, with metal ring
| 4 | Must contain a beryllium-copper diaphragm
| 5 | Must be shock resistant and made in accordance with ISO 81060-1
| 6 | Must have an inflation bulb with palm rest and metal valve
| 7 | Must be resilient up to at least 50.000 uses
| 8 | Must have a one-piece cuff (without bladder)
| 9 | Must have a biocompatible cuff
| 10 | Must be have an easily cleanable cuff
12 |  | Stethoscope |  | 70 |  | 0
| 1 | Must be a double-tube system
| 2 | Must have a stainless steel collapsible head piece
| 3 | Must have a double-head chest piece, with anti-chill
| 4 | Must have a high acoustic performance diaphragm, diameter min. 42 mm, frequency range min. 20 - 2000 Hz
| 5 | Must be tested in accordance with ISO 266
| 6 | Must have a total length at least 70 cm
13 |  | BLE Pulse oxymeter |  | 30 |  | 0
| 1 | Must have a SpO2 measurement range 0 - 100%
| 2 | Must have an SpO2 measurement error at most 3% when in motion and 2% without motion
| 3 | Must have a pulse measurement range at least 25 - 240 bpm
| 4 | Must have a pulse measurement error at most 5 bpm when in motion and 3 bpm without motion
| 5 | Must have a perfusion index measurement range at least 0.02 - 20 %
| 6 | Must have a color display with a viewing angle at least 150 degrees with rotation
| 7 | Must display numeric values for SpO2 and pulse, a plethysmography curve and battery capacity
| 8 | Must be powered by AAA batteries with at least 15 hours of operating time
| 9 | Must be at least IP23 certified (ingress protection)
| 10 | Must have a weight without batteries of at most 60 g
| 11 | Must support wireless communication via Bluetooth LE