## Price Schedule
Bidder: |  |  | Offered price (currency): | 0
LOT 1 | OP DEPARTMENT
ID | Equipment name (Technical Specification Requested) | QTY | Unit price | Total price per line item
1 | Anesthesia machine with integrated gas monitoring, high range | 4 |  | 0
2 | Anesthesia machine with integrated gas monitoring, mid-range | 4 |  | 0
3 | Set of furniture for the operating room | 4 |  | 0
4 | Suction pump | 4 |  | 0
5 | HF generator | 4 |  | 0
6 | Operating table with universal table top | 3 |  | 0
7 | Intraaortic balloon pump | 2 |  | 0
8 | ECMO unit | 2 |  | 0
9 | 4K Video endoscopic tower for cardiac surgery | 2 |  | 0
10 | Mobile x-ray device | 1 |  | 0
11 | Ventilator system, high-range | 13 |  | 0
12 | Ventilator system, mid-range | 12 |  | 0
13 | Portable suction unit | 9 |  | 0
14 | Transport respirator | 3 |  | 0
15 | Videolaringoscope - rigid | 6 |  | 0
16 | Videolaringoscope - flexibile | 2 |  | 0
17 | Laringoscopy set | 15 |  | 0
18 | Set of equipment for difficult intubation | 5 |  | 0
19 | Self inflating rescusitation system | 50 |  | 0
20 | Monitor for hemodinamic parametars with SG catheter and oximetry central chateter | 5 |  | 0
21 | Unit for dialsys and hemodiafiltration | 1 |  | 0
22 | Videoendoscopic tower with gastroscope and colonoscope | 1 |  | 0
23 | Mobile Fiberoptic LED Bronchoscope | 4 |  | 0
24 | Surgical light with satelites, camera and two monitors | 3 |  | 0
|  |  | Total Price per Lot: | 0
0

## OP.BLOCK
Bidder:
LOT 1 |  | OP Department |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Anesthesia machine with integrated gas monitoring, high range |  | 4 |  | 0
| 1 | The device should be for working with children and adult patients
| 2 | The device must have a closed - circular system of anesthesia
| 3 | The device should have the possibility of low-flow or minimal flow anesthesia
| 4 | The device must have an aspirator (suction module) and an auxiliary module O₂
| 5 | The device must have a system for draining excess anesthetic gases
| 6 | The device should be able to operate autonomously (without mains power supply) min. 90 minutes
| 7 | The device should be able to accurately measure and monitor the inhaled anesthetic consumed
| 8 | Fresh gas delivery:
| 9 | Electronic control of the delivery of a mixture of oxygen and N2O or oxygen and air
| 10 | Electronic control of carrier change between air and N2O
| 11 | Electronic regulation of O2 concentration in nitric oxide-oxygen mixture from at least 28 Vol% to 100Vol%
| 12 | Electronic regulation of O2 concentration in air-oxygen mixture from at least 21 Vol% to 100Vol%
| 13 | The device must have a button for immediate delivery of O2 at least 35 l / min
| 14 | O2 safety flow for emergency ventilation adjustable via a special system from 0 to 10 L / min
| 15 | The device must be able to display the mixture of gases and gas concentrations on the screen
| 16 | The device must be able to install two vaporizers
| 17 | The vaporizer must have electronic injection
| 18 | The device must be able to charge the vaporizer during the operation of the device
| 19 | One Sevoflurane vaporizer is supplied with the device
| 20 | Respiratory system:
| 21 | The device must have the possibility of a quick one-step transition from manual to automatic ventilation and vice versa
| 22 | The manual ventilation system must have a relief APL valve
| 23 | The device must have a CO2 absorber canister with a volume of 0.7 liters or more
| 24 | The device must be able to bypass the CO2 absorber during operation
| 25 | Ventilator:
| 26 | Suitable for all patients (from neonatal to adults)
| 27 | Ventilation modes:
| 28 | Volume controlled ventilation (VC)
| 29 | Pressure controlled ventilation (PC)
| 30 | Press Support (PS)
| 31 | Pressure volume controlled (PRVC)
| 32 | Synchronized intermittent mandatory ventilation (SIMV)
| 33 | Manual emergency ventilation possible in case of power failure (external and battery)
| 34 | Patient parameters are automatically adjusted depending on the patient's age
| 35 | Basic parameters that the device should provide:
| 36 | Fresh gas flow 0.3 l / min to 20 l / min
| 37 | Respiratory volume from 20 ml to 2000 ml
| 38 | Respiratory rate 4-100 per min
| 39 | Inspiratory pause 0 to 30% or 0 - 1.5 sec
| 40 | Ratio I: E min. 4: 1 to 1:10
| 41 | PEEP: 0-50 cmH2O
| 42 | Adjustable trigger
| 43 | Monitoring of anesthesia machines
| 44 | The device should have integrated monitoring of the following sizes:
| 45 | Pressure
| 46 | Inhalation and exhalation volume
| 47 | Minute volume inhalation and exhalation
| 48 | Ability to display flow-volume and volume-pressure loops
| 49 | The device must have an integrated gas analyzer and display on the monitor:
| 50 | CO2 concentration inhalation and exhalation
| 51 | O2 inhalation and exhalation concentration
| 52 | H2O inhalation and exhalation concentration
| 53 | Concentration of anesthetic gas inhalation and exhalation
| 54 | MAC calculation
| 55 | Automatic detection of gases in the mixture
| 56 | All ventilation and gas monitoring data are compared and processed on one central screen
| 57 | The screen of the device should have a minimum of 15 inches in color, touch screen
| 58 | Display of 6 color curves at the same time, where optional curves can be displayed: pressure, flow, volume,anesthetic, CO2, O2
| 59 | Automated complete check-in upon switching on the appliance
| 60 | Automated compliance and leak test
| 61 | Automated vaporizer test
| 62 | Possibility to adjust the brightness of the screen
| 63 | Storing the following parameters (graphic and numeric trends):
| 64 | Nitric oxide concentration
| 65 | Oxygen concentration
| 66 | CO2 concentration
| 67 | Anesthetic gas concentration
| 68 | Minute volume
| 69 | Lung compliance
| 70 | Alarms
| 71 | Alarm of minute exhalation volume
| 72 | System leak alarm
| 73 | High airway pressure alarm
| 74 | High continuous APL pressure alarm (in manual mode)
| 75 | PEEP alarm
| 76 | Respiration rate alarm
| 77 | Apnea alarm
| 78 | O2 inhalation and exhalation concentration alarm
| 79 | CO2 inhalation and exhalation concentration alarm
| 80 | Inhalation and exhalation anesthetic concentration alarm
| 81 | Respiratory N2O concentration alarm
| 82 | Technical malfunction alarm (technical alarm)
| 83 | Existence of alarm priority (ie different alarm levels)
2 |  | Anesthesia machine with integrated gas monitoring, mid-range |  | 4 |  | 0
| 2.1 | The device should be for working with children and adult patients
| 2.2 | The device must have a closed - circular system of anesthesia
| 2.3 | The device should have the possibility of low-flow or minimal flow anesthesia
| 2.4 | The device must have an aspirator (suction module)
| 2.5 | The device must have a system for draining excess anesthetic gases
| 2.6 | The device should be able to operate autonomously (without mains power supply) min. 90 minutes
| 2.7 | The device should be able to accurately measure and monitor the inhaled anesthetic consumed
| 2.8 | Fresh gas delivery:
| 2.9 | Electronic control of the delivery of a mixture of oxygen and N2O or oxygen and air
| 2.1 | Electronic control of carrier change between air and N2O
| 2.11 | Electronic regulation of O2 concentration in nitric oxide-oxygen mixture from at least 28 Vol% to 100Vol%
| 2.12 | Electronic regulation of O2 concentration in air-oxygen mixture from at least 21 Vol% to 100Vol%
| 2.13 | The device must have a button for immediate delivery of O2 at least 56 l / min
| 2.14 | O2 safety flow for emergency ventilation adjustable via a special system from 0 to 10 L / min
| 2.15 | The device must be able to display the mixture of gases and gas concentrations on the screen
| 2.16 | The device must be able to install minimum one vaporizers
| 2.17 | The vaporizer must have electronic injection
| 2.18 | The device must be able to charge the vaporizer during the operation of the device
| 2.19 | One Sevoflurane vaporizer is supplied with the device
| 2.2 | Respiratory system:
| 2.21 | The device must have the possibility of a quick one-step transition from manual to automatic ventilation and vice versa
| 2.22 | The manual ventilation system must have a relief APL valve
| 2.23 | The device must have a CO2 absorber canister with a volume of 0.7 liters or more
| 2.24 | The device must be able to bypass the CO2 absorber during operation
| 2.25 | Ventilator:
| 2.26 | Suitable for all patients (from neonatal to adults)
| 2.27 | Ventilation modes:
| 2.28 | Volume controlled ventilation (VC)
| 2.29 | Patient parameters are automatically adjusted depending on the patient's age
| 2.3 | Basic parameters that the device should provide:
| 2.31 | Fresh gas flow 0.1 l / min to 20 l / min
| 2.32 | Respiratory volume from 20 ml to 2000 ml
| 2.33 | Respiratory rate 4-100 per min
| 2.34 | Inspiratory pause 0 to 30% or 0 - 1.5 sec
| 2.35 | Ratio I:E : min. 1:8 to 4:1
| 2.36 | PEEP: 0-50 cmH2O
| 2.37 | Inspiratory flow: min. 200L/min
| 2.38 | Inspiratory minute volume: 0.3-60/min
| 2.39 | Adjustable trigger
| 2.4 | Respiratory monitoring of anesthesia machines
| 2.41 | Inspiratory and expiratory minute volume
| 2.42 | Inspiratory and expiratory tidal volume
| 2.43 | Pressure
| 2.44 | Lung characteristics
| 2.45 | Ability to display flow-volume and volume-pressure loops
| 2.46 | The device must have an integrated gas analyzer and display on the monitor:
| 2.47 | CO2 concentration inhalation and exhalation
| 2.48 | O2 inhalation and exhalation concentration
| 2.49 | H2O inhalation and exhalation concentration
| 2.5 | Concentration of anesthetic gas inhalation and exhalation
| 2.51 | MAC calculation
| 2.52 | Automatic detection of gases in the mixture
| 2.53 | All ventilation and gas monitoring data are compared and processed on one central screen
| 2.54 | The screen of the device should have a minimum of 15 inches in color, touch screen
| 2.55 | Display of 6 color curves at the same time, where optional curves can be displayed: pressure, flow, volume,anesthetic, CO2, O2
| 2.56 | Automated complete check-in upon switching on the appliance
| 2.57 | Automated compliance and leak test
| 2.58 | Automated vaporizer test
| 2.59 | Possibility to adjust the brightness of the screen
| 2.6 | Alarms
| 2.61 | Expiratory Minute Volume alarm : High and Low
| 2.62 | System leak alarm
| 2.63 | High airway pressure alarm
| 2.64 | High continuous APL pressure alarm
| 2.65 | PEEP alarm
| 2.66 | Respiration rate alarm
| 2.67 | Apnea alarm
| 2.68 | O2 inhalation and exhalation concentration alarm
| 2.69 | CO2 inhalation and exhalation concentration alarm
| 2.7 | Inhalation and exhalation anesthetic concentration alarm
| 2.71 | Respiratory N2O concentration alarm
| 2.72 | Technical malfunction alarm (technical alarm)
3 |  | Set of furniture for the operating room |  | 4 |  | 0
| 1 | Hydraulic instrument table: |  | 2
| 2 | The whole table is made of stainless steel, with 3 wheels with a diameter of Ø 50 mm
| 3 | Working plate size 700 mm x 500 mm,
| 4 | Working plate table can be moved -90 ° to + 90 °
| 5 | Possibility to adjust the height of the hydraulic table with the foot pedal from 1000 mm to 1300 mm.
| 6 | The table is mobile on 3 wheels, with the possibility of locking the wheels.
| 7 | Table load maximum 30 kg.
| 8 | Round chair with black upholstered seat: |  | 1
| 9 | The chair is mobile, hydraulically adjustable in height from 600 mm to 850 mm ±20mm
| 10 | The seat is at least 350 mm wide and is upholstered. Upholstery thickness minimum 100 mm
| 11 | The maximum load of the chair is up to 200 kg.
| 12 | Transport table for instruments and materials: |  | 2
| 13 | The table is made of stainless steel with 2 plates measuring 1000 x 650 mm ±50mm Plates must be able to move to the desired height
| 14 | The table is mobile with four wheels and the possibility of locking the wheels.
| 15 | The maximum load of each plate is up to 20 kg. Table height 1000 mm.
| 16 | Stair with 2 treads: |  | 1
| 17 | The stair is made of stainless steel, and the top layer is coated with a special rubber that prevents
| 18 | slipping.
| 19 | The maximum load is 100 kg.
| 20 | Dimensions: 600 mm x 650 mm x 200/440 mm.
| 21 | Stainless steel mobile holder for 2 waste bins: |  | 1
| 22 | The carrier is mobile with 5 wheels with a diameter of Ø 50 mm, the volume of the container is 5.5 liters.
| 23 | The maximum weight of waste in each container is 4 kg.
| 24 | Beam base diameter 635 mm, height 825 mm
| 25 | Mobile holder for 1 waste bag: |  | 1
| 26 | The carrier is mobile on 4 wheels with a diameter of Ø 75 mm, made of stainless steel
| 27 | Bracket dimensions 350 mm x 450 mm x 905 mm.
4 |  | Suction pump |  | 4 |  | 0
| 1 | Suction capacity nominal +/– 5 % 40 l/min
| 2 | Vacuum 0–90 kPa, (0–680 mmHg)
| 3 | Suction jar 4 l with screw cover |  | 2
| 4 | Support ring for 4 l jar
| 5 | Bottle against overflow
| 6 | Silent run with maximum noise ratio of 39,4 dB
| 7 | PVC suction circuit with stop valve
| 8 | Input Voltage 230V /50–60 Hz
| 9 | Weight (without accessories) 20 kg±1%
| 10 | Mobile stand
| 11 | MSF microbiological suction filter |  | 2
| 12 | Vacuum gauge in kPa, mmHg and Bar, colour coded
5 |  | HF generator |  | 4 |  | 0
| 1 | Microprocessor controlled electrosurgical unit
| 2 | Minimum 20 million measurements per second as a function of real-time current control in relation to the applied power to the tissue
| 3 | Output power minimum 400 W when cutting
| 4 | Output power minimum 200 W during coagulation
| 5 | Wi-Fi connection for service mode or operating room integration system
| 6 | Equipotency connection
| 7 | User color TFT or LCD touch screen with a diagonal of at least 26 cm and to have a precise and clear display of each activation and program.
| 8 | Possibility to set a minimum of 300 programs
| 9 | Number of monopolar cutting modes minimum 5 modes
| 10 | Number of monopolar coagulation modes minimum 5 modes
| 11 | A mode that enables the simultaneous operation of two positive monopolar electrodes on two operating fields of one patient
| 12 | Number of bipolar cutting modes minimum 2 modes
| 13 | Number of bipolar coagulation modes minimum 2 mod
| 14 | Number of thermofusion modes 1 mode and more
| 15 | Number of terminals for monopolar and bipolar active instruments minimum 4 terminals plus 1 terminal for neutral electrode
| 16 | The connectors on the device are easily interchangeable and have the possibility of combining the settings of the connectors according to the choice of the end user
| 17 | Possibility to select the connector to which it can be used for both mopolar or bipolar instrument
| 18 | That the HF generator can be upgraded with: Argon generator, Water knife, smoke evacuator or irrigation pump
| 19 | COMPOSITION OF THE ELECTROSURGICAL UNIT AND NECESSARY ACCESSORIES:
| 20 | HF generator with appropriate monitor and special MONOPOLAR software for
| 21 | simultaneous coagulation on two operative fields
| 22 | Foot pedal
| 23 | Neutral self-adhesive electrode with equipotential additional protective ring, 50/1
| 24 | Neutral electrode cable
| 25 | Positive electrode handle with keys and cable
| 26 | Set of monopolar electrodes
| 27 | Bipolar forceps
| 28 | Cable for Bipolar forceps
| 29 | Set of monopolar electrodes spatula 3x24 mm, L 45 mm, package 5/1
| 30 | Set of monopolar electrodes spatula 3x24 mm, L 60 mm, package 5/1
| 31 | Set of monopolar electrodes spatula 2.3x21 mm, L 180 mm
| 32 | Set of monopolar needle electrodes 0.8x20 mm, curved, package 5/1
| 33 | Set of monopolar bead electrodes Ø 4 mm, L 35 mm, package 5/1
| 34 | Elongation of monopolar electrodes 100 mm
6 |  | System Operating table with universal table top |  | 3 |  | 0
| 1 | Operating table with electro-hydraulic control
| 2 | Trolley insertion angles should be at least 90°, for easy transfer even in crowded and tight spaces
| 3 | Universal table top width 530 minimum
| 4 | Universal table top length - minimum 730 without extensions /maximum 2.300 mm with all extensions
| 5 | Height range - 1.100 mm or higher - 650 mm or lower
| 6 | Loading capacity - 350 kg without limitations
| 7 | Articulated patient weight including accessories - up to 250 kg
| 8 | Motorized Trandelenburd/Reverse Trendelenburg +45°/-45°
| 9 | Motorized lateral tilt +28°/-28° minimum
| 10 | Lower back +80°/-30°
| 11 | Upper back section inclination should be automatic and min 65degrees for UP and 35degrees for down
| 12 | Leg plate +90°/-90°
| 13 | Right/left leg sections +90°/-90° individually
| 14 | Longitudinal shift minimum 370 mm
| 15 | O-position with one control
| 16 | 3 motorized articulations to improve fine-tuning of patient positioning
| 17 | Floor lifting trolley to move the base in OR when there is table top on trolley between 2 procedures
| 18 | Battery type and battery working time per charge must be indicated (Battery charger must be included)
| 19 | The electronic sensors permits to avoid the collisions between the motorised sections and between the motorized sections and the floor
| 20 | Components
| 21 | Transportable column
| 22 | Floor lifting trolley to move the base in OR when there is table top on trolley
| 23 | Table top universal with 80mm pad
| 24 | Transport trolley - 1 pc
| 25 | Remote control - wired and wireless with charger and cable
| 26 | Double joint headrest with 80mm pad
| 27 | Extension back plate with 80mm pad
| 28 | Double-hinged armboard - 2 pcs
| 29 | 4-piece leg rest with 80mm pad
| 30 | Body strap - 2 pcs
| 31 | Infusion stand
| 32 | Flexible anestesiology frame with clamps needed
| 33 | GelPad for headrest
| 34 | GelPad for arms - 2 pcs
| 35 | Heel protectors -2 pcs
7 |  | Intraaortic balloon pump |  | 2 |  | 0
| 1 | Intraaortic balloon pump with fiber-optic technology
| 2 | One button start
| 3 | Automatic lead & trigger selection
| 4 | Automatic inflate & deflate timing
| 5 | User options to fine-tune deflation timing within auto mode
| 6 | Battery time for one battery 90 min
| 7 | Doppler fixed on the IABP to control the arterial perfusion
| 8 | Automatic condensate removal module
| 9 | Print-out of the current data
| 10 | Print-out of alarm history information
| 11 | Print-out of trigger history information
| 12 | Zeroing possible after insertion of the IAB
| 13 | In autofill mode the IABP automatically purge and fill the IABP when local atmospheric pressure decreases or increases
| 14 | Automatic in vivo recalibration every 2 hours or sooner should patient or environmental conditions change
| 15 | Fiber-optic pressure signal transmission to a patient monitor
| 16 | An electrically isolated, low-level pressure output from the pump to send the fiber-optic arterial pressure waveform directly to a patient monitor by simply attaching an appropriate cable
8 |  | ECMO unit |  | 2 |  | 0
| 1 | Compact portable device for ECMO (extracorporeal membrane oxygenation) for use in intensive care units, interventional cardiology departments, operating rooms and emergency centers. Completely certified possibility of using the device for transporting patients by both road and air transport (helicopter, plane).
| 2 | One drive and control device, compact, weight up to 10 kg±1kg with installed software and input connections for measurement and control: 4 external pressures, 3 internal pressures, 2 external temperatures, 2 internal temperatures, venous oxygen saturation, hemoglobin, hematocrit , flow and detection of air bubbles. Level sensor integrated in the device. Power supply of the device 230V / 50-60Hz DC as well as 11-28V AC. Battery support for a minimum of 90 minutes.
| 3 | Connections on the device for: USB external memory for storing data on the operation of the device, connection for alarming the nurse, Ethernet, input for ECG signal from an external drive unit. Class IIb device reference to the IPX4 standard for water resistance (during transport of the patient in the rain, during floods ...). Cable set, included in delivery. Night mode option - backlit control screen.
| 4 | Possibility to use sets with flow from 0.5 to 7L / min, small filling volume (up to 275ml), diffusion membrane with heparin, which has integrated sensors for necessary measurements in transport: venous oxygen saturation, hemoglobin, hematocrit, minimum 3 pressures and temperatures . Operation of the device with one set coated with heparin, certified for operation of 30 days, either at rest or in the transport of the patient.
| 5 | Patient heating device, with Hansa connection and a set of quick-connect power hoses. Maintains the temperature in the extracorporeal circulatory flow in the range from 33 ° C to 39 ° C. Automatic shut-off when the temperature reaches 40 ° C. Measurement tolerance of + 0.1 ° C. A sensor that measures temperature in the range of 7 ° C to 50 ° C. Audible alarm, 230V power supply, 50 / 60Hz, weight less than 10kg. Class IIb medical device.
9 |  | 4K Video endoscopic tower for cardiac surgery |  | 2 |  | 0
| 1 | 4K camera control unit
| 2 | Touch screen display
| 3 | Resolution: min.4096 x 2160 pixel-a
| 4 | USB port on camera controller for directly saving images and videos to USB storage media
| 5 | Softwer for enhanced tissue differentiation
| 6 | Patient data can be entered via a USB keyboard
| 7 | 3G-SDI input (picture in picture)
| 8 | Video output:2 x HDMI 4K, or HD-SDI (4K)
| 9 | 3G-SDI output
| 10 | 4K 3CCD Camera Head
| 11 | Fully autoclavable 134°C
| 12 | 3 x 1/3 inch CMOS image sensor
| 13 | Two (2) individually programmable camera head buttons for remote control
| 14 | C-Mount Zoom Coupler, Focal length ≈ f = 13-29mm.
| 15 | Led Light Source 300W
| 16 | Working life of lamp approx. 30.000h
| 17 | Brightness control 0-100%
| 18 | Fiber optic light cable, diameter min. 5 mm, length min.230 cm
| 19 | 4K Medical Monitor min.31"
| 20 | Diagonal:min 31 inch
| 21 | Resolution-min:4096 x 2160 Pixel
| 22 | VESA 100 standard connector
| 23 | Input signals HDMI, BNC (3G-SDI)
| 24 | 4K UHD Telescope 30°, diameter 10mm, length ≈300 mm, autoclavable
| 25 | 4K UHD Telescope 0°, diameter 10mm, length ≈300 mm, autoclavable
| 26 | Universal video trolley
| 27 | 3 shelves, 2 individually adjustable
| 28 | with integrated isolating transformer
| 29 | Insufflator
| 30 | Touch screen display
| 31 | With smoke evucation
| 32 | Integrated CO2 gas heating function (37°C)
| 33 | Gas flow rate: primary filling 1 l/min., high flow 2 l/min. to 45 l/min. (± 20 %) preselectable in 1 l/min steps
| 34 | Irrigation and aspiration pump
| 35 | Irrigation ≈ 2l/min. (± 10%)
| 36 | Suction ≈ 2 l/min. (± 10%),
| 37 | Tubing set, reusable
10 |  | Mobile radiography digital system with wireless FPD (Flat Panel Detector) |  | 1 |  | 0
| 1 | Power min. 35 kW
| 2 | Voltage in range of min 40-133 kV
| 3 | Maximum current value, min. 450 mA
| 4 | mAs values in range of min 0.4-320 mAs
| 5 | Battery power indicator
| 6 | X-RAY TUBE
| 7 | Anode heat capacity, min. 120 kHU
| 8 | Single focus max 0.8 mm or dual focus, small max 0.8 mm, big min 1.2 mm
| 9 | Manual or automatic collimation
| 10 | Collimator rotation min +/- 90º
| 11 | WIRELESS PORTABLE FLAT PANEL DETECTOR
| 12 | Scintilator material: CsI
| 13 | Detector dimensions min: 35х43 cm
| 14 | Acquisition depth min. 16 bits
| 15 | Pixel size max. 150 μm
| 16 | ACQUISITION WORKSTATION INTEGRATED IN SYSTEM
| 17 | LCD touch screen monitor
| 18 | Image preview in max 3 seconds after exposure
| 19 | Gray scale inversion
| 20 | Zoom
| 21 | Image Rotation
| 22 | L/R markers
| 23 | Annotations
| 24 | DICOM functionality min: Send, Storage Commitment Print, MWM
| 25 | Image archiving on CD/DVD and USB in DICOM format
| 26 | Exposure Index (EI) display
| 27 | MECHANICAL PARAMETERS
| 28 | Motorized back/forward movement on batteries
| 29 | System weight max 400 kg
| 30 | System width max 600 mm
| 31 | System height in transport position, max 150cm±10cm
| 32 | System height in transport position, max 150cm±10cm
11 |  | Ventilator system, high-range |  | 13 |  | 0
| 1 | The device is on its own mobile trolley
| 2 | Power supply: 220V, 50 Hz
| 3 | The device must have a battery for auxiliary power supply for at least 3 hours
| 4 | Primary gas supply: central hospital gas distribution Air / O2 (inlet pressure 2.0-6.0 bar)
| 5 | Auxiliary gas supply: possibility of installing 2 auxiliary cylinders under pressure with a volume of 5 liters
| 6 | Display: minimum 15 "TFT-LCD touchscreen monitor
| 7 | Parameter recording: trend minimum 72 h.
| 8 | Figure 4 curves simultaneously: pressure, flow, volume, CO2
| 9 | Loop display: volume-pressure and flow / volume
| 10 | The device has the possibility of upgrading a mini compressor for the production of dry medical air
| 11 | The device has integrated nebulizer
| 12 | Patient categories:
| 13 | * Adults, Pediatric (3-250 Kg)
| 14 | * Neonatus (0.3-8 Kg)
| 15 | Ventilation modes:
| 16 | Controlled ventilation:
| 17 | Pressure controlled ventilation (PC)
| 18 | Volume controlled ventilation (VC)
| 19 | Pressure volume controlled (PRVC)
| 20 | Supported ventilation:
| 21 | Pressure support/ CPAP (PS / CPAP)
| 22 | Volume support (VS)
| 23 | Combined ventilation:
| 24 | Synchronized Intermittent Mandatory Ventilation SIMV (VC) + PS
| 25 | Synchronized Intermittent Mandatory Ventilation SIMV (PC) + PS
| 26 | Synchronized Intermittent Mandatory Ventilation SIMV (PRVC) + PS
| 27 | Bi-Vent / APRV
| 28 | Non-invasive ventilation:
| 29 | NIV PC
| 30 | NIV PS
| 31 | Nasal CPAP
| 32 | Optional ventilation modes:
| 33 | Neurally Adjusted Ventilatory Assist NAVA
| 34 | Optional non-invasive ventilation modes:
| 35 | NIV Neurally Adjusted Ventilatory Assist NAVA
| 36 | Basic parameters :
| 37 | Inspiratory respiratory volume: 2-4000 ml
| 38 | Inspiratory minute volume: 0.5-60 ml
| 39 | PEEP: 0-50 cmH2O
| 40 | O2 concentration: 21-100%
| 41 | Respiratory rate: 4-150 in min
| 42 | Ratio I:E : min. 1:10-4:1
| 43 | Alarms:
| 44 | Airway pressure
| 45 | Exhalation minute volume: upper and lower limit
| 46 | Apnea alarm
| 47 | Respiratory frequency
| 48 | Expiratory pressure: upper and lower limit
| 49 | High continious pressure
| 50 | O2 concentration
| 51 | NECESSARY ACCESSORIES
| 52 | O2 connection hose, NIST standard, 3 m |  | 13
| 53 | Air connection hose, NIST standard, 3 m |  | 13
| 54 | Hose positioning bracket |  | 13
| 55 | Expiratory cassette |  | 13
| 56 | Expiratory filters, 60/1 |  | 13
| 57 | Set of breathing system for adults |  | 13
12 |  | Ventilator system, mid-range |  | 12 |  | 0
| 1 | The device is on its own mobile trolley
| 2 | Power supply: 220V, 50 Hz
| 3 | The device must have a battery for auxiliary power supply for at least 3 hours
| 4 | Primary gas supply: central hospital gas distribution Air / O2 (inlet pressure 2.0-6.0 bar)
| 5 | Auxiliary gas supply: possibility of installing 2 auxiliary cylinders under pressure with a volume of 5 liters
| 6 | Display: minimum 12 "TFT-LCD monitor
| 7 | Parameter recording: trend minimum 24 h.
| 8 | Figure 4 curves simultaneously: pressure, flow, volume, CO2
| 9 | Loop display: volume-pressure and flow / volume
| 10 | If necessary, the device can be removed from the mobile wheelchair and used as a transport respirator (for intrahospital transport, transport by air by helicopter and transport by ambulance)
| 11 | The device has the ability to upgrade the set for use in the field of magnetic resonance imaging
| 12 | The device has the possibility of upgrading a mini compressor for the production of dry medical air
| 13 | Patient categories:
| 14 | * Adults (10-250 Kg)
| 15 | * Neonatus, Pediatric (0.5-30 Kg)
| 16 | Ventilation modes:
| 17 | Controlled ventilation:
| 18 | Pressure controlled ventilation (PC)
| 19 | Volume controlled ventilation (VC)
| 20 | Pressure volume controlled (PRVC)
| 21 | Supported ventilation:
| 22 | Press supported / CPAP (PS / CPAP)
| 23 | Volume supported (VS)
| 24 | Combined ventilation:
| 25 | Synchronized Intermittent Mandatory Ventilation SIMV (VC) + PS
| 26 | Synchronized Intermittent Mandatory Ventilation SIMV (PC) + PS
| 27 | Synchronized Intermittent Mandatory Ventilation SIMV (PRVC) + PS
| 28 | Bi-Vent / APRV
| 29 | Non-invasive ventilation:
| 30 | NIV PC
| 31 | NIV PS
| 32 | Nasal CPAP
| 33 | Optional ventilation modes:
| 34 | Neurally Adjusted Ventilatory Assist NAVA
| 35 | Optional non-invasive ventilation modes:
| 36 | NIV Neurally Adjusted Ventilatory Assist NAVA
| 37 | Basic parameters :
| 38 | Inspiratory respiratory volume: 2-4000 ml
| 39 | Inspiratory minute volume: 0.3-60 ml
| 40 | PEEP: 0-50 cmH2O
| 41 | O2 concentration: 21-100%
| 42 | Respiratory rate: 4-150 in min
| 43 | Ratio I: E min. 1: 10-4: 1
| 44 | Alarms:
| 45 | Airway pressure
| 46 | Exhalation minute volume: upper and lower limit
| 47 | Apnea alarm
| 48 | Respiratory frequency
| 49 | Expiratory pressure: upper and lower limit
| 50 | High permanent pressure
| 51 | O2 concentration
| 52 | NECESSARY ACCESSORIES
| 53 | O2 connection hose, NIST standard, 3 m |  | 12
| 54 | Air connection hose, NIST standard, 3 m |  | 12
| 55 | Hose positioning bracket |  | 12
| 56 | Expiratory cassette |  | 12
| 57 | Expiratory filters, 60/1 |  | 12
| 58 | Set of breathing system for adults |  | 12
13 |  | Portable suction unit |  | 9 |  | 0
| 1 | Suction capacity minimum 40 liters per minute
| 2 | Vacuum from 0-90 kPa with the possibility of regulation
| 3 | Suction container of 2 liters with lid and holder
| 4 | Safety container
| 5 | Microbiological filter (2 pieces)
| 6 | The noise level must not exceed 40 decibels / submit a certificate as proofissued by a domestic accredited institution
| 7 | Power supply 230V / 50-60 Hz
14 |  | Transport respirator |  | 3 |  | 0
| 1 | Transport ventilator for ventilation therapy of adults and children.
| 2 | LCD screen min 5”
| 3 | AC and battery power
| 4 | Internal battery provides at least 4 hours of work. Possibility to upgrade with second battery to provide at least 10 hours of work.
| 5 | O2 gas supply pressure 250-600 kPa. Flow at least 180 L/min
| 6 | Ventilation modes: A/C-V , A/C-P, SIMV-V, SIMV-P, CPAP, Manual
| 7 | Additional functions: Apnea backup ventilation , manual breath
| 8 | Minimal range of tidal volume: 0 - 2000 ml
| 9 | Minimal range of breathing frequency: 1 - 120 bpm
| 10 | Minimal range of inspiratory pressure: 5 cm H2O to 50 cm H2O
15 |  | Videolaringoscope - rigid |  | 6 |  | 0
| 1 | One handle for neonates, pediatric, adult, obese pateints, no need different handle for different blades
| 2 | Touch Screen Operation, quick and convenient
| 3 | Built-in Training Video, quick learning
| 4 | 3.5 inch LCD Monitor, large displaying
| 5 | HD Camera, good resolution
| 6 | Real-time anti-fog function, no foggy troubles
| 7 | Indoor/Outdoor mode, suitable different environment
| 8 | LED Light Source
| 9 | Photography,Video Record, Sound Record, rhelpful for teaching and sharing
| 10 | USB/HDMI/WIFI, easy to output data
| 11 | Portable aluminum case containing a laryngoscope and all necessary accessories
16 |  | Videolaringoscope - flexibile |  | 2 |  | 0
| 1 | 3.5 "(inch) touch screen
| 2 | 4 hours of operation without charging the battery in the handle
| 3 | USB cable
| 4 | Working range: 3-50 mm
| 5 | Diameter of insertion tube: 5,2 mm
| 6 | Diameter of working channel: 2,2 mm
| 7 | Working length: 600 mm
| 8 | Bending capacity: up 150 ° and down 130 °
| 9 | LED light integrated
| 10 | AC adapter for charging the battery
| 11 | Photography,Video Record, Sound Record, rhelpful for teaching and sharing
| 12 | USB/HDMI/WIFI, easy to output data
| 13 | Portable aluminum case containing a laryngoscope and all necessary accessories
| 14 | laryngoscope tightness tester, working channel cleaning brush, biopsy gripper and user manual
17 |  | Laryngoscope set |  | 15 |  | 0
| 1 | Laringoscopes set includes: * Macintosh Fiber Optic Blade 3 134x14 * Macintosh Fiber Optic Blade 4 155x15 * Macintosh Fiber Optic Blade 5 177x15 * Fiber Optic Handle * Table charger * case
| 2 | Macintosh Fiber Optic Blades: * Chrome-plated, stainless, high-grade steel * Blade resistant to the oxidation and / or corrosion damage * Blades have no screwed joints, no external fiber bundles and no openings that could contaminate * The clean shape without any edges or corners * Compatible with all handles to the ISO 7376 (green standard)
| 3 | Laringoscope Handle * Compatible with all laryngoscope blades to the ISO 7376 (green standard) * Colour Temperature: 5,000 K
| 4 | Table Charger
18 |  | Set of equipment for difficult intubation |  | 5 |  | 0
| 1 | Aspiration catheter CH 14 (5mm) x 1
| 2 | Aspiration catheter CH 10 (3.3mm) x 1
| 3 | Macintosh laryngoscope set with spatulas 2,3,4 x 1
| 4 | - With a reliable 2.7 V vacuum light source for good illumination of the test area
| 5 | - Bulb that is quickly removed on the front side of the spatula
| 6 | - Can be autoclaved at 134 ° C for 5 minutes
| 7 | - Stainless steel spatula that is easy to clean
| 8 | - Innovative spatula design for easy intubation and good visibility
| 9 | - All blades are compatible with handles in accordance with ISO 7376
| 10 | - Supplied with spatulas (number 2 length 102mm, distal tip width 11.3 mm, number 3 length 130mm, distal tip width 13.7 mm, number 4 length 155mm, distal tip width 13.7 mm)
| 11 | - The rounded edges of the Macintosh blade reduce the risk of injury
| 12 | - Solid chromed metal handle
| 13 | - Carved surface of the handle for hygienic use and secure grip in the hand
| 14 | - Solid metal contact for reliable lighting activation.
| 15 | - Easy battery change at the bottom of the handle
| 16 | - laryngoscope handles are standardized according to ISO 7376
| 17 | - Durable metal contact to activate the lighting
| 18 | - Macintosh spatulas with innovative design for optimal visibility of epiglottis and vocal cords, easy endotracheal tube insertion
| 19 | Forceps for language x 1
| 20 | MAGILL forceps x 1
| 21 | Endotracheal tube 6 mm x 1
| 22 | Endotracheal tube 7 mm x 1
| 23 | Endotracheal tube 8 mm x 1
| 24 | Endotracheal tube 9 mm x 1
| 25 | Endotracheal tube 10 mm x 1
| 26 | GUDEL AIRWAY 60 MM-0 x 1
| 27 | GUDEL AIRWAY 70 MM-1 x 1
| 28 | GUDEL AIRWAY 80 MM-2 x 1
| 29 | GUDEL AIRWAY 90MM-3 x 1
| 30 | GUDEL AIRWAY 100 MM-4 x 1
| 31 | Aspiration catheter CH 12 (4mm) x 1
| 32 | Equipment bag x 1
| 33 | - Dimensions 65x35xh 35 cm
| 34 | - External pockets 5
| 35 | - Internal modular partitions 6
| 36 | - Made of solid waterproof polyester material 600D
| 37 | - Equipped with double yellow reflective strips.
| 38 | - Color: red
| 39 | - Waterproof rubber bottom to prevent water infiltration
| 40 | - Can be worn on the shoulder or arm.
19 |  | Self inflating rescusitation system |  | 50 |  | 0
| 1 | Recomended patient weight: over 50 kg
| 2 | Volumen of bag: 1,5 l for adult
| 3 | Conection tube for oxygen
| 4 | Maximal oxygen contrecation 99%
| 5 | Settings : 15 BPM and flow 15l/min
| 6 | Pressure realise valve : one way valve
| 7 | Detachable bag
20 |  | Monitor for hemodynamic parametars with SG catheter and oximetry central chateter |  | 5 |  | 0
| 1 | Hemodinamic monitor, with following features:
| 2 | Continuous measurement of cardiac output (CO),
| 3 | Continuous measurement of SVR, SV, SVV, PPV, MAP, SYS, DIA, HR, dP/dT, prediction index of hypotension
| 4 | Indexed values for measured parameters,
| 5 | All measures are obtained by analysing the arterial waveform
| 6 | Arterial waveform analysing is obtained using the minimally invasive sensor, through arterial approach and using monitor algorithm,
| 7 | monitor allows graphical image of the arterial waveform
| 8 | possibility of continuous measurement of the oxygenated central venous blood (ScvO₂),
| 9 | monitor allows graphical image of the hypotension index (HPI),
| 10 | monitor measures O2 distribution (DO2),
| 11 | Invasive measurement of hemodinamic parameters, with Swan Ganz cathethers
| 12 | Mixed Venous Oxygen Saturation (SvO2)
| 13 | Continuous Cardiac Output (CCO, CCI)
| 14 | Continuous Right Ventricular End Diastolic Volume (RVEDV, RVEDVI)
| 15 | Continuous Right Ventricular Ejection Fraction (RVEF)
| 16 | Systemic Vascular Resistance (SVR, SVRI) - Stroke Volume (SV, SVI)
| 17 | Right Ventricular Stroke Work Index
| 18 | Pulmonary Vascular Resistance
| 19 | Central Venous Oxygen Saturation (ScvO2)
| 20 | Simultaneously display of four differently coloured parameters in graphical and numerical data (parameters can be chosen),
| 21 | Continuous display of measured parameters in tabelar view,
| 22 | Continuous analyse of the measured parameters in the inerval of 5, 10, 15, 20 and 30 min – displayed in %
| 23 | Continuous measurement of tissue and cerebral oxymetry with four sensors,
| 24 | Racking of the measured parameters for the last 72h,
| 25 | Display : min 12" with resolution LCD
| 26 | Easy to use - touch screen,
| 27 | Monitor allows export of data via USB port, JPG, in PDF formatand tabelar form (EXCEL)
| 28 | Monitor has back up battery
| 29 | Monitor is modular and it is on stand
21 |  | Unit for dialsys and hemodiafiltration |  | 1 |  | 0
| 1 | General data
| 2 | Nominal voltage: 230 V (option: 120/240 V)
| 3 | Nominal frequency: 50 Hz / 60 Hz
| 4 | Nominal current (max.): max. 11 A for 230 V or 16 A for 120 V
| 5 | Weight (empty): approx. 85 kg in the basic design
| 6 | Water intake
| 7 | Pressure range: 0.5 to 6 bar
| 8 | Temperature range: +10 to +30 °C
| 9 | Concentrate supply: Canister/central supply/bicarbonate cartridge
| 10 | Pressure range: 0 to +1 bar
| 11 | Dialysis fluid system
| 12 | Temperature working range: selectable between +33 and +40 °C
| 13 | Conductivity processing: conductivity regulated
| 14 | Working range: - conductivity of bicarbonate 2 to 4 mS/cm
| 15 | or 4 to 7 mS/cm
| 16 | - total conductivity 12.5 to 16 mS/cm
| 17 | Measurement tolerance: ±0.2 mS/cm
| 18 | Flow: 300 to 800 ml/min
| 19 | Tolerance: ±5%
| 20 | Blood leak detector: optical, color-specific
| 21 | Alarm limit value: > 0.50 ml/min (HCT 45%) > 0.35 ml/min (HCT 25%)
| 22 | Ultrafiltration: - volume regulated through the balance chamber,
| 23 | ultrafiltration through ultrafiltration pump
| 24 | sequential ultrafiltration (Bergström)
| 25 | Working range: 0 to +4000 ml/h
| 26 | Measurement tolerance: 0.2 ml per chamber cycle, UF-pump tolerance <1%
| 27 | Degassing device: mechanical, through regulated degassing low
| 28 | pressure
| 29 | Extracorporeal circulation
| 30 | Blood pump: 2-roller pump
| 31 | Transportation rate: 50 to 600 ml/min
| 32 | Transportation tolerance: < 10% with pressure up to -150 mmHg
| 33 | Heparin pump: injection pump for 10-, 20-, and 30-ml syringes
| 34 | Transportation rate: 0.1 to 10 ml/h
| 35 | Transportation tolerance: < 10%
| 36 | Safety air detector: ultrasound measurement in the tube
| 37 | Protection system: ultrasound detector, automatic cyclical testing
| 38 | during the entire operating phase
| 39 | Pressure measurement at the
| 40 | arterial feed into the dialyzer
| 41 | (PBE) working area: 0 to +700 mmHg
| 42 | Measurement tolerance: ±10 mmHg
| 43 | Arterial input pressure measurement
| 44 | (PA) working area: -400 to +400 mmHg
| 45 | Measurement tolerance: ±10 mmHg
| 46 | Venous backflow pressure measurement
| 47 | (PV) working area: 20 to +390 mmHg
| 48 | Measurement tolerance: ±10 mmHg
| 49 | Disinfection
| 50 | Chemical disinfection/cleaning: automatic running of program with minimal
| 51 | free rinse time; various disinfection means may
| 52 | be installed
| 53 | Thermal disinfection/cleaning: automatic running of program at approx. 85 °C
| 54 | Central thermal/chemical for cleaning/rinsing the feed connection
| 55 | cleaning/rinsing: for the device with simultaneous chemical or
| 56 | thermal disinfection
| 57 | System for reverse osmosis capacity 100l/h
22 |  | Videoendoscopic tower with gastroscope and colonoscope |  | 1 |  | 0
| 1 | HD video processor of the latest generation of manufacturers
| 2 | Compatible with LED light source
| 3 | Digital Output - minimum 1 DVI output resolution 1920 x 1080px, minimum 1 DVI output resolution 1280 x 1024px,
| 4 | Compatible with generation endoscopes with CMOS chip (camera) on top - distal part of endoscope
| 5 | Compatible with older generation endoscopes with CCD chip (camera) on top - distal endoscope
| 6 | Control all CPU functions via the front panel
| 7 | Ability to connect and control the processor via the keyboard
| 8 | Connect to the light source with a special cable on the back of the processor and light source
| 9 | Internal memory to store a minimum of 20,000 photos in JPEG format
| 10 | External memory connector for storing USB photos
| 11 | Possibility to connect a digital printer, digital DVD recorder
| 12 | Connecting the latest generation endoscope wirelessly without electrical connection
| 13 | Possibility of memorization:
| 14 | * 45 patients
| 15 | * 20 clinical procedures
| 16 | * 20 attending physicians
| 17 | DICOM interface incorporated in the processor
| 18 | Power supply: 100-240V; 50-60Hz; 0.8-0.5A
| 19 | Possibility of chromoendoscopy at the moment before image processing, in combination with LED light source using different wavelengths minimum 2 observation modes:
| 20 | * for better visualization and diagnostics of surface parts of tissues
| 21 | * for better visualization of the vascular network in the deeper parts of the observed tissue
| 22 | Possibility of chromoendoscopy at the moment after image processing by digital staining of the observed tissue by changing the wave lengths in order to single out certain regions of interest to the doctor and facilitate diagnostics
| 23 | Anti blur function (automatically selects the highest quality image from multiple images)
| 24 | LED (Light Emitting Diode) light source
| 25 | Combination of LEDs that emit different wavelengths of light and form intense white light - minimum 4 (four) diodes
| 26 | Compatible with the appropriate video processor that automatically controls the light intensity, as well as special modes chromatography
| 27 | Wireless connection for endoscopes incorporated in the light connector:
| 28 | * optical communication sensor
| 29 | * coil for wireless power supply of the endoscope
| 30 | * optical laser for image transfer from the endoscope
| 31 | Service life of LED diodes minimum 10,000 working hours
| 32 | Air Supply Pump - Integrated irrigation pump with 4 operating modes - Off / Low / Mid / Hi
| 33 | Power supply: 100-240V; 50-60Hz; 1.2-0.7A
| 34 | Class 1 LED product
| 35 | Video endoscope gastroscope
| 36 | The latest generation endoscope with a CMOS chip (camera) on the distal part
| 37 | Video endoscope compatibility with processor and LED light source
| 38 | There is no special electrical connection to the processor
| 39 | Wireless connection with processor incorporated in the light connector:
| 40 | * optical communication sensor
| 41 | * coil for wireless power supply of the endoscope
| 42 | * optical laser for image transfer to the processor
| 43 | Field of view up to 140 °
| 44 | Depth of field of view : 2-100 mm
| 45 | Distal diameter : 9.5 mm
| 46 | Insertion tube diameter up to 9.5 mm
| 47 | Working channel diameter about 3 mm
| 48 | Working length up to 1100 mm
| 49 | Total length up to 1400 mm
| 50 | Bending capacity up/down 210 ° / 90 °
| 51 | Bending capacity left/right 100 ° / 100 °
| 52 | Additional water channel
| 53 | Compatible with a video processor that allows chromatographic observation of tissues
| 54 | Video endoscope colonoscope
| 55 | The latest generation endoscope with a CMOS chip (camera) on the distal part
| 56 | Video endoscope compatibility with processor and LED light source
| 57 | There is no special electrical connection to the processor
| 58 | Wireless connection with processor incorporated in the light connector:
| 59 | * optical communication sensor
| 60 | * coil for wireless power supply of the endoscope
| 61 | * optical laser for image transfer to the processor
| 62 | Field of view about 170 °
| 63 | Depth of field of view about 2-100 mm
| 64 | Diameter of the distal part about 12 mm
| 65 | Insertion tube diameter about 12 mm
| 66 | Working channel diameter about 4 mm
| 67 | Working length up to 1700 mm
| 68 | Bending capacity up/down 180 ° / 180 °
| 69 | Bending capacity left/right 160 ° / 160 °
| 70 | Additional water channel
| 71 | Compatible with a video processor that allows chromatographic observation of tissues
| 72 | LCD medical monitor 26"
| 73 | 1920 (H) x 1080 (H) pixels, 450cd / m2,
| 74 | Inputs: Analogni (PAL/NTSC, Y/C) In/Output, PC (Analogue RGB, DVI)
| 75 | Mobile stand for endoscopic equipment
| 76 | Steel - Aluminum construction.
| 77 | Three shelves and a movable monitor holder.
| 78 | Sliding keyboard shelf
| 79 | Four wheels, two with brakes.
| 80 | A hand-held monitor holder that can be moved
| 81 | Endoscope carrier
| 82 | Mains voltage and earth connection.
| 83 | Suction unit
| 84 | Max. vacuum 90 kPa
| 85 | Bottle 2 l. / Volt. 230 V / 50 Hz
| 86 | the noise level must not exceed 40 decibels
| 87 | Double protection filter
| 88 | Suction between 40 and 45 liters / min.
23 |  | Mobile Fiberoptic LED Bronchoscope |  | 4 |  | 0
| 1 | Field of view 120 °
| 2 | Depth of field of view 1-50 mm
| 3 | Distal diameter up to 5 mm
| 4 | Insertion tube diameter up to 5 mm
| 5 | The diameter of the working channel : 2 mm
| 6 | Working length : 600 mm
| 7 | Bending capacity: up 180 ° and down 130 °
| 8 | Lithium battery for continuous operation of the doctor for at least 60 minutes
| 9 | Accessories for the Bronchoscope
| 10 | Leakage tester for endoscope
| 11 | Endoscope wash brushes
| 12 | Endoscope hand wash set
| 13 | Biopsy channel rubber set 10/1
24 |  | Surgical light with satelites, camera and two monitors |  | 3
| 1 | LED light
| 2 | LED lifespan: minimum 50.000 hours
| 3 | Surgical light housing made of aluminum or better
| 4 | Temperature increase at head height: less than 1°C
| 5 | Safety glass on the lamp, resistant to cleaning agents and scratch-proof
| 6 | Homogenous light beam that doesn’t create shadows.
| 7 | Illumination in standard mode: Main light: min 160.000 Lux Satellite light: min 160.000 Lux, Satellite light: min 110.000 Lux,
| 8 | Technical requirements of the main light and satelites
| 9 | Laminar flow compatibility
| 10 | Electrical light dimming: 25%-100%
| 11 | Patten size for from 19 to 24 cm
| 12 | Pattern of the light must be mechanically and continuosly adjustable from the sterile field
| 13 | Ambient light for MIS at the intensity of 100 Lux at 1 m
| 14 | Light field diameter (d10) at distance of 1 meter = 200 mm or better
| 15 | 4 different color temperature from 3.500 to 5.000 K or better
| 16 | Color rendering index Ra (1-8) more than 95
| 17 | Red rendering index R9 between 91 and 97 for different color temperature
| 18 | Illumination depth without refocusing (L1+L2) at 20% not less than 1.000 mm
| 19 | Total irradiance at maximum intensity must not exceed 500W /sq.m.
| 20 | Power supply - primary voltage 100-240V, 50Hz
| 21 | Power consumption with 24 V max. 60 W
| 22 | The lights must be supplied with a backup system with an autonomy of 1 hour minimum
| 23 | Technical specification of the camera and monitors
| 24 | Full HD control unit with USB module for recording video and images on an external hard disk (USB memory)
| 25 | Both the video and the image will have a ratio of 16: 9 and the optical zoom should not be less than X12
| 26 | The HD video output should be the native signal of the camera in order to not multiply signal processing
| 27 | The video management system will convert this signal if necessary
| 28 | The Camera should be removable
| 29 | The camera controls will be available on the light head as well as on the wall control Panel with, at a minimum, the following functions: ON/OFF Rotation of the camera without stop (infinite) Zoom +/-
| 30 | HD monitors 27"
| 31 | Expandable anchor, to be cut on site
| 32 | Customization of arm length