## Lot - US systems
Bidder: |  |  |  | Offered price (currency): | Offered price: | 0
LOT 9 | Ultrasound for intreventional procedures
ID | Equipment name (Technical Specification Requested) | QTY | Technical Specification Offered (model) | Unit price | DAP Unit Price | Total Price per line item
1 | PORTABLE COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM | 1 |  |  |  | 0
2 | COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM WITH 4D TEE PROBE | 4 |  |  |  | 0
3 | COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM | 2 |  |  |  | 0
4 | PREMIUM VASCULAR ULTRASOUND SYSTEM | 1 |  |  |  | 0
5 | PREMIUM 4D COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM 1 | 2 |  |  |  | 0
6 | PREMIUM 4D COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM 2 | 1 |  |  |  | 0
7 | PREMIUM 4D COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM 3 | 2 |  |  |  | 0
8 | PORTABLE VASCULAR ULTRASOUND SYSTEM | 1 |  |  |  | 0
9 | PREMIUM VASCULAR ULTRASOUND SYSTEM | 2 |  |  |  | 0
10 | Holder and tray for TEE probes | 9 |  |  |  | 0
|  |  |  | Total Price per LOT: | Total Price per LOT: | 0

## 1.Portable Cardiac US
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | PORTABLE COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM |  | 1 |  | 0
| 1 | Portable ultrasound system (laptop style) with digital beamformer and min. 500000 effective channels
| 2 | Cart on wheels adjustable in height with minimum 3 probe connectors
| 3 | Color 15" high resolution LCD screen
| 4 | The device supports transducers with a frequency range of 1 to 15 MHz or more
| 5 | Console weight maximum 8kg (with battery)
| 6 | Integrated three-lead ECG module
| 7 | Integrated battery with scan time minimum 45 min.
| 8 | Integrated SSD or HDD
| 9 | Maximum depth min. 30 cm (probe dependent)
| 10 | Physical TGC
| 11 | Full alphanumeric QWERTY keyboard integrated on the control panel (not acceptable on the touch screen), together with an ergonomic handle for carrying the device.
| 12 | Advanced image processing technique to remove speckle in real-time
| 13 | Compound imaging
| 14 | Automatic tissue and automatic spectral optimization
| 15 | Automated software, based on "speckle tracking" technology for assessing the global function of the left ventricle and regional deformation of the wall.
| 16 | System supports the following transducer types:
| 17 | Sector phased array
| 18 | Convex array
| 19 | Linear array
| 20 | TEE
| 21 | Matrix 4D Volume TEE
| 22 | Intraoperative
| 23 | Intracardiac echo (ICE) catheters
| 24 | Operating Modes
| 25 | 2D tissue (B-Mode)
| 26 | 2D color flow
| 27 | 2D angio flow (Power Doppler)
| 28 | Continuous wave Doppler (CW)
| 29 | M-mode
| 30 | Color M-mode
| 31 | Anatomical M-mode
| 32 | Pulsed wave Doppler (PW)
| 33 | Tissue Doppler Imaging (TDI)
| 34 | Hardware and software options which can be added later:
| 34.1 | Matrix 4D volume TEE probe Frequency Range: 2.0 - 8.0 MHz or wider Field of View (FOV): 90°
| 34.2 | LVO Contrast
| 34.3 | Fully compatible with offline review/reporting station for advanced post-processing analysis
| 34.4 | Possibility for integration with the appropriate angio operating theatre (C-arm), which results in the integration of the ultrasound device and the angio operating theatre.
| 35 | Connections and archiving
| 35.1 | Ethernet network connection
| 35.2 | DICOM 3.0
| 36 | Transducers:
| 36.1 | Matrix and/or single crystal phased array probe Frequency Range: 1.0 - 5.0 MHz or wider Field of View (FOV): 90°
| 36.2 | Linear probe Frequency Range: 3.0 – 12.0 MHz or wider Number of elements: min. 160 Field of view (FOV): 40 mm (+/- 5 mm)
| 36.3 | Convex probe Bandwidth: 1.0– 5.0 MHz or wider Number of elements: min. 160 Field of view (FOV): min. 70°
| 37 | B&W printer

## 2.CardiacUS
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
2 |  | COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM WITH 4D TEE PROBE |  | 4 |  | 0
| 1 | Ultrasound system, not portable, with wheels and front handles
| 2 | Advanced beamforming technology, software or hardware based, with minimum 4000000 effective channels and dynamic range of minimum 280 dB
| 3 | Advanced image processing technique for removing speckle in real-time
| 4 | Automatic tissue optimization by single keystroke
| 5 | Four equal active probe ports,in pinless connectors
| 6 | 21" High-Definition (HD), LED/LCD display on articulated arm
| 7 | Color touch screen min. 12"
| 8 | Touch keyboard
| 9 | Physical or on touch-panel TGC with min. 8 sliders
| 10 | Control panel, adjustable by height, with rotation capability (min. 175° from the center)
| 11 | Weight max. 85 kg
| 12 | Tissue imaging maximum depth range of minimum 30 cm - probe specific
| 13 | Maximum frame rate minimum 1900 fps, depending on probe, settings and applications
| 14 | System frequency 1 – 22 MHz or wider
| 15 | Transducer Types:
| 16 | Sector phased array
| 17 | 4D TEE volume matrix array
| 18 | TEE probes
| 19 | Linear array
| 20 | Convex array
| 21 | Intraoperative probe
| 22 | Pencil probes
| 23 | Operating Modes
| 24 | 2D(B mode)
| 25 | M-mode
| 26 | Color M-mode
| 27 | Anatomical M-mode
| 28 | Color doppler
| 29 | Pulsed wave (PW) doppler with HPRF
| 30 | Continuous wave doppler (CW)
| 31 | 2D angio flow or Power doppler
| 33 | Tissue doppler imaging
| 34 | 4D full-volume scanning on 4D TEE probe
| 35 | Quantitative data LV for global and segmental wall motion ( bulls-eye view )
| 36 | LVO Contrast
| 37 | Minimum two independent simultaneous scan planes on 4D probes with possibility to use Color Doppler
| 39 | The mitral valve quantification tool based on 4D acquisition
| 40 | Archiving and connections:
| 42 | DICOM 3.0
| 43 | Ethernet network connection
| 44 | Display Port or HDMI, USB
| 45 | Hardware and software options which can be added later:
| 45.1 | Fully compatible with offline review/reporting station for advanced post-processing analysis
| 45.2 | Linear 2D transducer for superficial soft tissue: frequency from 2 to 22 MHz or wider, single crystal technology or matrix technology, transducer width: min. 50 mm, min. 1900 elements
| 45.3 | The device is upgradeable with a program for the fusion of CT, MRI DICOM data with live ultrasound image(with the possibility of automatic overlapping according to anatomy markers) and navigation for carrying out therapeutic procedures (biopsies, ablations, injections) with display of instrument position on CT, MRi image combined with fusion
| 46 | Transducers:
| 46.1 | 2D matrix and/or single crystal phased array transducer Bandwidth: 1.0 – 5.0 MHz or wider Field of View: min. 90°
| 46.2 | Matrix 4D volume TEE transducer Bandwidth: 2.0 – 7.0 MHz or wider Field of View: min. 90°
| 46.3 | Linear array transducer Bandwidth: 3.0 – 12.0 MHz or wider Field of View: 40 mm (+/- 5 mm)
| 47 | Battery which provides a transportation mode that keeps the system ready to scan after connecting to the power outlet
| 48 | ECG cable
| 49 | B/W video printer

## 3.CardiacUS
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
3 |  | COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM |  | 2 |  | 0
| 1 | Ultrasound system, not portable, with wheels and front handles
| 2 | Advanced beamforming technology, software or hardware based, with minimum 4000000 effective channels and dynamic range of minimum 280 dB
| 3 | Advanced image processing technique for removing speckle in real-time
| 4 | Automatic tissue optimization by single keystroke
| 5 | Four equal active probe ports,in pinless connectors
| 6 | 21" High-Definition (HD), LED/LCDdisplay on articulated arm
| 7 | Color touch screen min. 12"
| 8 | Touch keyboard
| 9 | Physical or on touch-panel TGC with min. 8 sliders
| 10 | Control panel, adjustable by height, with rotation capability (min. 175° from the center)
| 11 | Weight max. 85 kg
| 12 | Tissue imaging maximum depth range of minimum 30 cm - probe specific
| 13 | Maximum frame rate minimum 1900 fps, depending on probe, settings and applications
| 14 | System frequency 1 – 22 MHz or wider
| 15 | Transducer Types:
| 16 | Sector phased array
| 17 | 4D TEE volume matrix array
| 18 | TEE probes
| 19 | Linear array
| 20 | Convex array
| 21 | Intraoperative probe
| 22 | Pencil probes
| 23 | Operating Modes
| 24 | 2D (B mode)
| 25 | M-mode
| 26 | Color M-mode
| 27 | Anatomical M-mode
| 28 | 2D color doppler
| 29 | Pulsed wave (PW) doppler with HPRF
| 30 | Continuous wave doppler (CW)
| 31 | 2D angio flow or Power doppler
| 32 | Tissue doppler imaging
| 33 | 4D full-volume scanning with 4D TEE probe
| 34 | Quantitative data LV for global and segmental wall motion ( bulls-eye view )
| 35 | LVO Contrast
| 36 | Hardware and software options which can be added later:
| 36.1 | Fully compatible with offline review/reporting station for advanced post-processing analysis
| 36.2 | Matrix 4D volume TEE transducer Bandwidth: 3.0 – 7.0 MHz or wider Field of View: min. 90°
| 36.3 | Linear 2D transducer for superficial soft tissue: frequency from 2 to 22 MHz or wider, single crystal technology or matrix technology, transducer width: min. 50 mm, min. 1900 elements
| 36.4 | The device is upgradeable with a program for the fusion of CT, MRI DICOM data with live ultrasound image(with the possibility of automatic overlapping according to anatomy markers) and navigation for carrying out therapeutic procedures (biopsies, ablations, injections) with display of instrument position on CT, MRi image combined with fusion
| 37 | Archiving and connections:
| 37.1 | DICOM 3.0
| 37.2 | Ethernet network connection
| 37.3 | Display Port or HDMI, USB
| 38 | Transducers:
| 38.1 | 2D matrix and/or single crystal phased array transducer Bandwidth: 1.0 – 5.0 MHz or wider Field of View: min. 90°
| 38.2 | Linear array transducer Bandwidth: 3.0 – 12.0 MHz or wider Field of View: 40 mm (+/- 5 mm)
| 39 | Battery which provides a transportation mode that keeps the system ready to scan after connecting to the power outlet
| 40 | ECG cable
| 41 | B/W video printer

## 4.Vascular
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
4 |  | PREMIUM VASCULAR ULTRASOUND SYSTEM A |  | 1 |  | 0
| 1 | Ultrasound system, not portable, with wheels and front handles
| 2 | Advanced imageformer (beamformer), software or hardware based, with min. 4500000 processing channels
| 3 | Four equal active probe ports,in pinless connectors
| 4 | LED/LCD/WLED display min. 24'', wide screen, high-resolution, with translation (independent of console), swivel and fold down
| 5 | Color touch display screen min. 12"
| 6 | Control panel,interactive backlight, adjustable by height and depth, with rotation capability (min. 175° from the center)
| 7 | Full-sized, backlit alphanumeric keyboard
| 8 | Integrated SSD or HDD
| 9 | Integrated speakers
| 10 | Maximum depth of field min. 30 cm (probe dependent)
| 11 | Maximum dynamic range min. 320 dB
| 12 | Maximum framerate min. 2700 fps (Hz)
| 13 | Harmonic imaging
| 14 | Supported transducer types:
| 15 | Convex array
| 16 | Linear array
| 17 | Matrix linear array
| 18 | Sector phased array
| 19 | Micro convex array
| 20 | Intraoperative
| 21 | Operating Modes
| 22 | B-Mode
| 23 | M-Mode
| 24 | Anatomical M-Mode
| 25 | Color Doppler Mode
| 26 | Power Doppler Imaging
| 27 | PW Doppler
| 28 | Extended field of view imaging
| 29 | Speckle reduction imaging which provides multiple levels of speckle reduction
| 30 | Automatic B-Mode and spectral optimization
| 31 | Automatic positioning in PW: adjusts ROI position, sample volume position and steering
| 32 | Easy, fast visualization of tiny vessels
| 33 | Contrast imaging
| 34 | Spatial compounding
| 35 | Automatic IMT measurement
| 36 | DICOM 3.0
| 37 | Complete user manual available on board by pressing F1 or some other key
| 38 | Hardware and software options which can be added later:
| 38.1 | Upgradeable to 2D and 3D linear transducer with frequency range from 3 to 14 MHz or wider , with min. 55.000 crystal elements, equipped with matrix technology. The transducer supports simultaneous display technology in min. two cross section planes in real time, with the use of Color Doppler allowed.
| 38.1 | Upgradeable to abdominal electronic transducer for 2D and 4D exams, having crystals arranged in several rows, with the frequency range from 1 to 6MHz or wider with min. 9.000 elements. The transducer supports simultaneous display technology in min. two cross section planes in real time, with the use of Color Doppler allowed.
| 39 | Battery which provides a transportation mode that keeps the system ready to scan after connecting to the power outlet
| 40 | Transducers:
| 40.1 | Linear probe Bandwidth: 3.0 – 12.0 MHz or wider Number of elements: min. 160 Field of view (FOV): 40 mm (+/- 5 mm)
| 40.2 | Convex probe Bandwidth: 1.0 – 5.0 MHz or wider Number of elements: min. 160 Field of view (FOV): min. 110°
| 40.3 | Sector probe Bandwidth: 1.0 – 5.0 MHz or wider Field of view (FOV): min. 90º
| 41 | USB ports
| 42 | Ethernet
| 43 | Display Port(DP)
| 44 | Digital B&W thermal printer

## 5. Cardio
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
5 |  | PREMIUM 4D COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM |  | 2 | 0 | 0
| 1 | 4D ultrasound system, not portable, with wheels, front and rear handles for easy transport
| 2 | Advanced beamforming technology, software based, without limitations of focal zones, with dynamic range of minimum 400 dB
| 3 | Four active probe ports
| 4 | OLED display min. 22", wide screen, High-Definition (HD), flicker-free, with translation (independent of console): horizontal bidirectional, vertical height adjustment and swivel to any viewing direction
| 5 | Ultra-high-resolution, wide-screen format, color, multi-touch LCD screen min. 12"
| 6 | Floating operating keyboard adjustable in three dimensions: height, rotation, extension
| 7 | Maximum depth range not less then 50 cm – probe specific
| 8 | Maximum frame rate in 2D min. 5000 fps, depending on probe, settings and applications
| 9 | Maximum volume rate min. 1000 in 4D TTE, depending on probe, settings and applications
| 10 | Maximum volume rate min. 3000 in 4D TEE, depending on probe, settings and applications
| 11 | LVO contrast imaging
| 12 | Low MI contrast imaging
| 13 | Integrated HDD
| 14 | Automatic tissue optimization with single keystroke
| 15 | Single push, automatic, real-time optimization of PW or CW spectrum scale and baseline
| 16 | Technology for harmonic tissue imaging providing improved lateral and contrast resolution
| 17 | Real-time triplex mode: 2D+CFM+CW
| 18 | An advanced image processing technique to remove speckle in real-time examining
| 19 | Multiple angle imaging – multiple co-planar images from different angles combined into a single image in real-time
| 20 | Transducer Types:
| 21 | Convex array
| 22 | Linear array
| 23 | Sector phased array
| 24 | Single crystal matrix array
| 25 | 4D volume matrix phased array
| 26 | TEE transducers
| 27 | 4D TEE volume matrix array
| 28 | Pencil transducers
| 29 | Operating Modes
| 30 | 2D tissue (B-Mode)
| 31 | M-Mode
| 32 | Anatomical M-mode
| 33 | Curved anatomical M-mode
| 34 | 2D Color Flow Mode (CFM)
| 35 | 2D Virtual Apex imaging or similar
| 36 | 2D Angio flow (Power Doppler)
| 37 | PW Doppler with High PRF
| 38 | TVI, TVD, Tissue Tracking
| 39 | Continuous Wave Doppler (CW)
| 40 | 4D tissue
| 41 | 4D Color Flow
| 42 | 4D full volume scanning single-beat and multi-beat
| 43 | Simultaneous display of minimum 12 slices extracted from the 4D volume data (tissue and/or color) with combination of short-axis and long-axis standard views
| 44 | Two and three independent simultaneous scan planes
| 45 | Parametric imaging tool which gives quantitative data for global and segmental wall motion and allows comprehensive assessment at a glance by combining three longitudinal views into one comprehensive bulls-eye view
| 46 | Simultaneous display of three independent random slices through the 4D volume (tissue and color)
| 47 | Stress Echo
| 48 | The left atrium quantification tool based on 4D acquisition
| 49 | The left ventricule quantification tool based on 4D acquisition
| 50 | The right ventricule quantification tool based on 4D acquisition
| 51 | The mitral valve quantification tool based on 4D acquisition
| 52 | The aortic valve quantification tool based on 4D acquisition
| 53 | 4D Strain with support for the following parameters: Area, longitudinal, circumferential, radial, twist and torsion
| 54 | Tool that enables placements of markers/annotations into a 4D ultrasound volume data set
| 55 | Acquisition and visualization providing enhanced display of anatomical structures using advanced shadowing techniques in combination with depth illuminating colors
| 56 | Storage and transfer formats: JPEG, MPEG, AVI, DICOM, Raw DICOM and VolDicom
| 57 | Ethernet network connection
| 58 | DICOM 3.0
| 59 | Transducers:
| 60 | 2D Active Matrix Single Crystal Phased Array Transducer, Bandwhidth: 1.5 – 4.5 MHz (+/- 0.5 MHz), Field of View: min. 120°
| 61 | Active Matrix 4D Volume Phased Array Transducer, Bandwhidth: 1.5 – 5.5 MHz (+/- 0.5 MHz), Field of View: min. 90°
| 62 | Active Matrix 4D Volume TEE Transducer, Bandwhidth: 3.0 – 8.0 MHz (+/- 0.5 MHz), Field of View: min. 90°
| 63 | Pencil Transducer Bandwidth: 2.0 MHz
| 64 | Thermal B&W printer

## 6.US Cardio
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
6 |  | PREMIUM 4D COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM |  | 1 | 0 | 0
| 1 | 4D ultrasound system, not portable, with wheels, front and rear handles for easy transport
| 2 | Advanced beamforming technology, software based, without limitations of focal zones, with dynamic range of minimum 400 dB
| 3 | Four active probe ports
| 4 | OLED display min. 22", wide screen, High-Definition (HD), flicker-free, with translation (independent of console): horizontal bidirectional, vertical height adjustment and swivel to any viewing direction
| 5 | Ultra-high-resolution, wide-screen format, color, multi-touch LCD screen min. 12"
| 6 | Floating operating keyboard adjustable in three dimensions: height, rotation, extension
| 7 | Maximum depth range not less then 50 cm – probe specific
| 8 | Maximum frame rate in 2D min. 5000 fps, depending on probe, settings and applications
| 9 | Maximum volume rate min. 1000 in 4D TTE, depending on probe, settings and applications
| 10 | Maximum volume rate min. 3000 in 4D TEE, depending on probe, settings and applications
| 11 | LVO contrast imaging
| 12 | Low MI contrast imaging
| 13 | Integrated HDD
| 14 | Automatic tissue optimization with single keystroke
| 15 | Single push, automatic, real-time optimization of PW or CW spectrum scale and baseline
| 16 | Technology for harmonic tissue imaging providing improved lateral and contrast resolution
| 17 | Real-time triplex mode: 2D+CFM+CW
| 18 | An advanced image processing technique to remove speckle in real-time examining
| 19 | Multiple angle imaging – multiple co-planar images from different angles combined into a single image in real-time
| 20 | Transducer Types:
| 21 | Convex array
| 22 | Linear array
| 23 | Sector phased array
| 24 | Single crystal matrix array
| 25 | 4D volume matrix phased array
| 26 | TEE transducers
| 27 | 4D TEE volume matrix array
| 28 | Pencil transducers
| 29 | Operating Modes
| 30 | 2D tissue (B-Mode)
| 31 | M-Mode
| 32 | Anatomical M-mode
| 33 | Curved anatomical M-mode
| 34 | 2D Color Flow Mode (CFM)
| 35 | 2D Virtual Apex imaging or similar
| 36 | 2D Angio flow (Power Doppler)
| 37 | PW Doppler with High PRF
| 38 | TVI, TVD, Tissue Tracking
| 39 | Continuous Wave Doppler (CW)
| 40 | 4D tissue
| 41 | 4D Color Flow
| 42 | 4D full volume scanning single-beat and multi-beat
| 43 | Simultaneous display of minimum 12 slices extracted from the 4D volume data (tissue and/or color) with combination of short-axis and long-axis standard views
| 44 | Two and three independent simultaneous scan planes
| 45 | Parametric imaging tool which gives quantitative data for global and segmental wall motion and allows comprehensive assessment at a glance by combining three longitudinal views into one comprehensive bulls-eye view
| 46 | Simultaneous display of three independent random slices through the 4D volume (tissue and color)
| 47 | Stress Echo
| 48 | The left atrium quantification tool based on 4D acquisition
| 49 | The left ventricule quantification tool based on 4D acquisition
| 50 | The right ventricule quantification tool based on 4D acquisition
| 51 | The mitral valve quantification tool based on 4D acquisition
| 52 | The aortic valve quantification tool based on 4D acquisition
| 53 | 4D Strain with support for the following parameters: Area, longitudinal, circumferential, radial, twist and torsion
| 54 | Tool that enables placements of markers/annotations into a 4D ultrasound volume data set
| 55 | Acquisition and visualization providing enhanced display of anatomical structures using advanced shadowing techniques in combination with depth illuminating colors
| 56 | Storage and transfer formats: JPEG, MPEG, AVI, DICOM, Raw DICOM and VolDicom
| 57 | Ethernet network connection
| 58 | DICOM 3.0
| 59 | Transducers:
| 60 | 2D Active Matrix Single Crystal Phased Array Transducer, Bandwhidth: 1.5 – 4.5 MHz (+/- 0.5 MHz), Field of View: min. 120°
| 61 | Active Matrix 4D Volume Phased Array Transducer, Bandwhidth: 1.5 – 5.5 MHz (+/- 0.5 MHz), Field of View: min. 90°
| 62 | Active Matrix 4D Volume TEE Transducer, Bandwhidth: 3.0 – 8.0 MHz (+/- 0.5 MHz), Field of View: min. 90°
| 63 | Pencil Transducer Bandwidth: 2.0 MHz
| 64 | Thermal B&W printer
| 65 | Software package (installed on a standard PC, configured according to the minimum technical specification of the software package manufacturer) for basic and advanced viewing, reporting and quantitative analysis capabilities for 2D and 4D ultrasound images as if the images were still on the ultrasound system, with the following software:
| 66 | Parametric imaging tool which gives information about global and segmental strain and combines three longitudinal views into one comprehensive bulls-eye view
| 67 | Automated 2D EF measurement tool based upon a 2D-speckle tracking algorithm
| 68 | Software for the analysis of longitudinal, radial and circumferential myocardial deformation, rotation and twist from data obtained on ultrasound system
| 69 | Strain and strain rate imaging which gives information
| 70 | about synchronicity of myocardial motion with additional features:
| 71 | - Efficient segment specific TSI time measurements
| 72 | Automatic measurement of intima media thickness of the carotid artery as IMT or corresponding
| 73 | Automated measurement of LV volume and EF from volumetric data
| 74 | Software that uses the speckle tracking analysis of volumetric data (4D) to calculate the following parameters (global and/or segmental): longitudinal, circumferential, radial and area strain, as well as twist and torsion
| 75 | Automated alignment, segmentation and measurement of aortic annulus from volumetric data sets
76
| 77 | Semiautomated mitral valve quantification
| 78 | Semiautomated left atrial quantification package which offers the ability to visualize and quantify the left atrium volumes, ejection fraction and strain, based on images acquired with the 4D TTE probes
| 79 | Semiautomated right ventricular quantification
| 80 | Visualization providing enhanced display of anatomical structures using advanced shadowing techniques in combination with depth color maps
| 81 | Up to four ultrasound scanners compatible with software package can be simultaneously connect to the patient archive

## 7.US Cardio
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
7 |  | PREMIUM 4D COLOR DOPPLER CARDIAC ULTRASOUND SYSTEM (52) |  | 2 | 0 | 0
| 1 | 4D ultrasound system, not portable, with wheels, front and rear handles for easy transport
| 2 | Advanced beamforming technology, software based, without limitations of focal zones, with dynamic range of minimum 400 dB
| 3 | Four active probe ports
| 4 | OLED display min. 22", wide screen, High-Definition (HD), flicker-free, with translation (independent of console): horizontal bidirectional, vertical height adjustment and swivel to any viewing direction
| 5 | Ultra-high-resolution, wide-screen format, color, multi-touch LCD screen min. 12"
| 6 | Floating operating keyboard adjustable in three dimensions: height, rotation, extension
| 7 | Maximum depth range not less then 50 cm – probe specific
| 8 | Maximum frame rate in 2D min. 5000 fps, depending on probe, settings and applications
| 9 | Maximum volume rate min. 1000 in 4D TTE, depending on probe, settings and applications
| 10 | Maximum volume rate min. 3000 in 4D TEE, depending on probe, settings and applications
| 11 | LVO contrast imaging
| 12 | Low MI contrast imaging
| 13 | Integrated HDD
| 14 | Automatic tissue optimization with single keystroke
| 15 | Single push, automatic, real-time optimization of PW or CW spectrum scale and baseline
| 16 | Technology for harmonic tissue imaging providing improved lateral and contrast resolution
| 17 | Real-time triplex mode: 2D+CFM+CW
| 18 | An advanced image processing technique to remove speckle in real-time examining
| 19 | Multiple angle imaging – multiple co-planar images from different angles combined into a single image in real-time
| 20 | Transducer Types:
| 21 | Convex array
| 22 | Linear array
| 23 | Sector phased array
| 24 | Single crystal matrix array
| 25 | 4D volume matrix phased array
| 26 | TEE transducers
| 27 | 4D TEE volume matrix array
| 28 | Pencil transducers
| 29 | Operating Modes
| 30 | 2D tissue (B-Mode)
| 31 | M-Mode
| 32 | Anatomical M-mode
| 33 | Curved anatomical M-mode
| 34 | 2D Color Flow Mode (CFM)
| 35 | 2D Virtual Apex imaging or similar
| 36 | 2D Angio flow (Power Doppler)
| 37 | PW Doppler with High PRF
| 38 | TVI, TVD, Tissue Tracking
| 39 | Continuous Wave Doppler (CW)
| 40 | 4D tissue
| 41 | 4D Color Flow
| 42 | 4D full volume scanning single-beat and multi-beat
| 43 | Simultaneous display of minimum 12 slices extracted from the 4D volume data (tissue and/or color) with combination of short-axis and long-axis standard views
| 44 | Two and three independent simultaneous scan planes
| 45 | Parametric imaging tool which gives quantitative data for global and segmental wall motion and allows comprehensive assessment at a glance by combining three longitudinal views into one comprehensive bulls-eye view
| 46 | Simultaneous display of three independent random slices through the 4D volume (tissue and color)
| 47 | The left atrium quantification tool based on 4D acquisition
| 48 | The left ventricule quantification tool based on 4D acquisition
| 49 | 4D Strain with support for the following parameters: Area, longitudinal, circumferential, radial, twist and torsion
| 50 | Tool that enables placements of markers/annotations into a 4D ultrasound volume data set
| 51 | Acquisition and visualization providing enhanced display of anatomical structures using advanced shadowing techniques in combination with depth illuminating colors
| 52 | Storage and transfer formats: JPEG, MPEG, AVI, DICOM, Raw DICOM and VolDicom
| 53 | Ethernet network connection
| 54 | DICOM 3.0
| 55 | Transducers:
| 56 | 2D Active Matrix Single Crystal Phased Array Transducer, Bandwhidth: 1.5 – 4.5 MHz (+/- 0.5 MHz), Field of View: min. 120°
| 57 | Active Matrix 4D Volume Phased Array Transducer, Bandwhidth: 1.5 – 5.5 MHz (+/- 0.5 MHz), Field of View: min. 90°
| 58 | Active Matrix 4D Volume TEE Transducer, Bandwhidth: 3.0 – 8.0 MHz (+/- 0.5 MHz), Field of View: min. 90°
| 59 | Pencil Transducer Bandwidth: 2.0 MHz
| 60 | Thermal B&W printer

## 8.US Portable Vasc
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
8 |  | Portable ultrasound for Vascular Surgery |  | 1 | 0 | 0
| 1 | Laptop style portable ultrasound system
| 2 | Height adjustable docking cart with 3 probe ports and extended battery which allows minimum 3.5 hours of operation without mains power
| 3 | 15" high-resolution (minimum 1024 x 768) color LCD
| 4 | Console weight maximum 5,5 kg (with battery)
| 5 | Integrated solid state drive (SSD)
| 6 | Maximum scan depth minimum 30 cm (probe dependent)
| 7 | Maximum focus number minimum 8 (probe dependent)
| 8 | Alphanumeric keyboard
| 9 | Speckle reduction imaging which provides multiple levels of speckle reduction as SRI-HD or corresponding
| 10 | Spatial compounding with minimum 9 angles
| 11 | Auto tissue optimization
| 12 | Auto color flow optimization
| 13 | Auto spectrum optimization
| 14 | Convex view (Wider field of view) on linear probes
| 15 | High density write zoom, (HD Zoom)
| 16 | Report writer package
| 17 | Supported transducer types:
| 18 | Linear Array
| 19 | Linear Array with 4 programmable buttons
| 20 | Convex Array
| 21 | Microconvex Array
| 22 | Phased Array
| 23 | TEE
| 24 | Operating modes:
| 25 | B-Mode
| 26 | M-Mode
| 27 | Color Flow Mode (CFM)
| 28 | Power Doppler Imaging (PDI)
| 29 | Pulse Wave Doppler (PWD)
| 30 | Harmonic imaging as Coded Harmonic Imaging (CHI)
| 31 | On-board User Manual
| 32 | Patient information database
| 33 | Customizable user interface
| 34 | Hardware and software options which can be added later:
| 35 | Wide-band high frequency linear transducer Bandwidth: 10 - 22 MHz (+/- 0,5 MHz) Footprint: 8 x 20 mm (+/- 5 mm)
| 36 | Software which provides better hemodynamics visualization by combining effects of B-mode and color flow Doppler
| 37 | Console interfaces: 3 x USB, LAN, HDMI
| 38 | Docking cart interfaces: 4 x USB, DVI, speakers
| 39 | Transducers:
| 40 | Linear probe Bandwidth: 3.5 - 9 MHz or wider Footprint: 9 x 50 mm (+/- 5 mm)
| 41 | Convex probe Bandwidth: 2.0 - 5.0 MHz or wider Footprint: 18 x 65 mm (+/- 5 mm)
| 42 | Phased array probe Bandwidth: 2.0 - 4.0 MHz or wider Field of View: min. 120°
| 43 | Printer

## 9.US Vascular 41 i 45
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
9 |  | PREMIUM VASCULAR ULTRASOUND SYSTEM |  | 2 | 0 | 0
| 1 | Ultrasound system, not portable, with wheels and front and rear handles
| 2 | Advanced imageformer (beamformer), software based, with min. 11000000 processing channels
| 3 | Four active probe ports, not counting pencil probe connector, plus one inactive probe storage port
| 4 | OLED display min. 22'', wide screen, high-resolution (min. 1920 x 1080) with translation (independent of console), swivel and fold down
| 5 | High-Resolution, color, touch, display screen min. 12" with brightness adjustment and user-configurable layout
| 6 | Operator keyboard adjustable in three dimensions: height, rotation and extension
| 7 | Full-sized, backlit alphanumeric keyboard
| 8 | Integrated SSD min. 1TB
| 9 | Integrated DVD-R multi drive
| 10 | Integrated speakers with sub-woofer for premium sound
| 11 | Integrated gel warmer
| 12 | Maximum depth of field min. 50 cm (probe dependent)
| 13 | Maximum dynamic range min. 400 dB
| 14 | Maximum framerate min. 9500 fps (Hz)
| 15 | Harmonic imaging
| 16 | Supported transducer types:
| 17 | Convex array
| 18 | Micro convex array
| 19 | Linear array
| 20 | Sector phased array
| 21 | Matrix array
| 22 | Probes with multiple technological solutions implemented: single crystal, acoustic amplifier and cool stack
| 23 | Operating Modes
| 24 | B-Mode
| 25 | M-Mode
| 26 | Anatomical M-Mode
| 27 | Color Flow Mode (CFM)
| 28 | Power Doppler Imaging (PDI)
| 29 | PW Doppler
| 30 | Extended field of view imaging with min. 160 cm scan length
| 31 | Convex field of view available on linear and sector probes
| 32 | Speckle reduction imaging which provides multiple levels of speckle reduction
| 33 | Automatic B-Mode and spectral optimization
| 34 | Automatic positioning in CF and PW: adjusts ROI position, sample volume position and steering
| 35 | Method for showing flow in grayscale using B-Mode technology (not the doppler technology), which is based on reflection of blood cells
| 36 | Visualization of tiny vessels, displaying as a 3D effect
| 37 | Contrast imaging with time intensity curve analysis and parametric imaging
| 38 | Side-by-side comparison of previous ultrasound and other modality exams during live scanning
| 39 | Variable angle spatial compounding
| 40 | Automatic IMT measurement
| 41 | DICOM 3.0
| 42 | Hardware and software options which can be added later:
| 43 | Volume Navigation or corresponding option that includes fusion of CT / PET / MR / US acquisition with real-time ultrasound imaging and automatic synchronization (registration) of ultrasound imaging and CT / MR acquisition during fusion process, with sensor needle guidance
| 44 | Software for measuring liver attenuation (attenuation coefficient) by auto measure algorithm with reference B-mode
| 45 | CINE memory min. 1GB
| 46 | Transducers:
| 47 | Convex probe Bandwidth: 1.0 – 6.0 MHz (+/- 0.5MHz) Number of elements: min. 192 Field of view (FOV): min. 80°
| 48 | Linear probe Bandwidth: 2.0 – 10.0 MHz (+/- 0.5MHz) Number of elements: min. 192 Field of view (FOV): 45 mm (+/- 5 mm)
| 49 | Sector probe Bandwidth: 1.0 – 5.0 MHz (+/- 0.5MHz) Number of elements: min. 285 Field of view (FOV): min. 120º
| 50 | USB 3.0 ports
| 51 | Ethernet
| 52 | HDMI
| 53 | Digital B&W thermal printer

## 6.Holder
Bidder:
LOT 9 |  | Ultrasound System |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
10 |  | Holder and tray for TEE probes |  | 9 | 0 | 0
| 1 | Holder and tray for disinfection and holding of TEE probes