## Lot 13 MR
Bidder: |  |  |  | Offered price (currency): | 0
LOT 12 | Magnetic Resonance Imaging 1.5 T System with additional equipment
ID | Equipment name | QTY | Technical Specification Offered(model) | Total Price per line item
1.1. | Magnetic Resonance Imaging 1.5 T System | 1 |  |  | 0
|  |  |  | Total Price per Lot: | 0

## SPECS
Bidder:
LOT 12 |  |  |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
|  | MRI 1.5 T System |  | 1
| 1 | Superconducting magnet with active shielding
| 2 | Magnetic Field Strength, 1.5 Tesla
| 3 | Magnetic field homogeneity typical value in accordance with the V.R.M.S. method:
| 4 | - DSV, for spherical volume with 30 cm diameter:: less than 0.15 ppm
| 5 | - DSV, for spherical volume with 40 cm diameter: less than 0.65 ppm
| 6 | - DSV, for spherical volume with 50cm diameter: less than 5.5 ppm
| 7 | Zero Boil-off technology, 0 liters per year during normal operation - no loss of helium during the normal operation of the apparatus
| 8 | Shimming on patient
| 9 | Limit distance of the 5 Gauss line from the isocentre
| 10 | - Radially, along the X and Y axes - less than 2.5 m
| 11 | - Axially, along the Z axis - less than 4.0 m
| 12 | GANTRY
| 13 | Magnet length - less than 200 cm
| 14 | Diameter at isocentre, at least 70 cm
| 15 | Field of view - not less than 50 cm in each direction ± 5%
| 16 | Patient video surveillance system
| 17 | In-bore lighting
| 18 | In-bore ventilation
| 19 | On bore operation console on both sides of the patient table
| 20 | Patient communication system
| 21 | GRADIENT SYSTEM
| 22 | Maximum amplitude value in each orthogonal plane - not less than 44 mT / m
| 23 | Maximum slew rate at the maximum amplitude in each orthogonal plane - not less than 200 T/m/s
| 24 | Both maximum amplitude and Slew Rate values described in 3.1. and 3.2. must be met simultaneously
| 25 | Minimum value for TR time in Turbo (Fast) Spin Echo technique and its acquisition matrix of 256x256 - not longer than 6 msec.
| 26 | Minimum value for TE time in Turbo (Fast) Spin Echo technique and its acquisition matrix of 256x256 - less than 2.0 msec.
| 27 | Minimum value for TR time in the EPI (Echo Planar Imaging) technology and its acquisition matrix of 256x256 - less than 10 msec.
| 28 | The minimum value for TE in the EPI ( Echo Planar Imaging ) technology and its acquisition matrix of 256x256 - less than 2.7 msec.
| 29 | RF SYSTEM
| 30 | Water cooled RF amplifier
| 31 | Digital Radiofrequency technology.
| 32 | - Signal digitization is performed in the system bore or in the RF coil;
| 33 | - Digital signal transfer with optical fibers technology
| 34 | Number of independent RF channels in the gantry in a single FOV, min 64
| 35 | Number of physical coil elements that can be simultaneously connected during one exam, minimum 128
| 36 | Receiver signal resolution - not less than 32 bits
| 37 | Receiver dynamic range - not less than 160 dB
| 38 | RF COILS
| 39 | Number of coil elements (integrated or independent) in each requested RF coil, if it is not requested differently should correspond to the number of physical presented elements of that offered coil and not sum of the coil elements of the offered coil and another coil(s).
| 40 | Built-in coil for the whole body
| 41 | Multi-channel single coil or a combination of the two coils for thoracic, abdominal, pelvic and cardiac imaging with coverage (SI) of at least 50 cm and at least 16 elements per coil
| 42 | Head and neck imaging coil with a minimum of 19 elements
| 43 | Spine imaging coil with at least 40 elements and minimum 100 cm coverage, installed into the patient table
| 44 | Integrated coil technology (TDI, Tim 4G, Atlas, dStream, etc.)
| 45 | Dedicated coil for peripheral vascular examinations with at least 36 elements
| 46 | PATIENT TABLE
| 47 | Minimum table height, not higher than 70 cm
| 48 | Total scanning range - not less than 180 cm
| 49 | Most exams performed in head and feet first orientations
| 50 | Maximum patient weight for scanning, not less than 220 kg
| 51 | Movement of the table in horizontal direction (stepping) to avoid any re-positioning of the patient or RF coils during peripheral imaging.
| 52 | Fully dockable patient table
| 53 | PHYSIOLOGICAL PARAMETERS DISPLAY UNIT
| 54 | ECG with triggering
| 55 | Pulse
| 56 | Respiration
| 57 | OPERATOR CONSOLE AND POSTPROCESSING SERVER WITH 3 CLIENT WORKSTATIONS
| 58 | Four workstations - one for recording, three for data processing
| 59 | Direct connection with other digital diagnostic devices in DICOM format with specified DICOM capabilities:
| 60 | DICOM STORAGE
| 61 | DICOM QUERY/RETRIEVE
| 62 | DICOM PRINT
| 63 | Color LCD medical monitors, for the operator console and the workstations, one monitor the screen diagonal not less than 19" for operator console and 6 monitors (2 for each client) the screen diagonal not less than 24"
| 64 | Long-time archiving device (DVD-RW or CD-RW) on the operator console
| 65 | Reconstruction speed, minimum 60.000 2D FFT per second for 256x256 matrix and 100% FOV
| 66 | Server needs to support a mínimum of 30.000 DICOM images simultaneously and have RAM memory of mínimum 64 GB capacity
| 67 | The following application softwares can simultaneously be used by 3 client workstations
| 68 | · Multimodality advanced 3D visualization and processing software
| 69 | · MPR views with different slice thickness,
| 70 | · Rendering modes for all Thick Slab: MIP, MinIP, Volume Rendering
| 71 | · 3D point synchronization in any orientation on the fly
| 72 | · Real time reformat
| 73 | · Display of curved, lumen and cross section views
| 74 | · Multiple VR objects merge up to 8 Volume Rendered models into a single 3D view
| 75 | · Auto-center fly through in air or vessel contrast studies (airways, colon, angiography)
| 76 | · Lumen view unfolded 3D view of a previously outlined structure
| 77 | · Advanced segmentation tools such as Scalpel, Threshold, Dilate, Erode, Subtraction, Intersection, Addition.
| 78 | Quantitative determination of diffusion and perfusion parameters on the operator console and on the server
| 79 | Integrating multiple images (e.g., multiple spine, upper and lower abdominal, peripheral blood vessels) on the operator console or on the server
| 80 | Advanced cardiac package with special sequences and protocols for advanced cardiac imaging including 3D and 4D functionalities. It should support advanced techniques for ventricular function imaging, dynamic imaging, tissue characterization, coronary imaging, and more, on operator console
| 81 | The advanced software package to allow adjustments of scanning parameters to the patient's condition (e.g. heart rate, arhytmia, breathing capabilities, etc.) on the operator console
| 82 | Technique for rapid Cardiac cine imaging in 2D or 3D with advanced accelaration and reconstruction technique. The technique should allow for scan time reduction and full heart coverage in a sigle breath hold on the operator console.
| 83 | T1 and T2 myocardial tissue quantification maps on operator console and on server
| 84 | Evaluation package for integration (composing) of several diagnostic images in 1 “Full Format” (skeleton and blood vessels) on operator console or on server
| 85 | 3D data processing - MIP techniques, 3D rendering, 3D visualization capabilities in all planes on the operator console and on the server
| 86 | Possibility of image fusion on the operator console or on the server
| 87 | Possibility of data processing obtained by Single voxel, 2D and 3D Multivoxel spectroscopy on the operator console and on the server
| 88 | Possibility of processing Diffusion tensor imaging and tractography, as well as the possibility of fusion with anatomical images on the operator console and on the server
| 89 | Time intensity curves analysis and fusion of parameter maps with anatomy on the operator console and on the server
| 90 | Arterial Spin Labeling Quantization Possibility of Quantification of a Perfusion Without Contrast (Arterial Spin Labeling) either on the operator console or on the server
| 91 | Quantitative determination of diffusion and perfusion parameters on the operator console and on the server
| 92 | The possibility of processing of cardiac images on minimum two workstations at the same time.
| 93 | The minimum post-processing capabilities should include:
| 94 | · Review and edit of patient data
| 95 | · Identify series for review
| 96 | · Create and share comments
| 97 | · Series composer
| 98 | · Multiple layouts, synchronization and cross-referencing
| 99 | · Measurement tools
| 100 | · Annotation tools
| 101 | · MPR, 3D/4D volume rendering
| 102 | · Ventricular and atrial assessment
| 103 | · Fully automated contour detection for volumes and ejection fraction calculation
| 104 | · Flow analysis (volumes velocities, gradients) with offset and aliasing correction and Qp/Qs calculation
| 105 | · Perfusion analysis fore rest and stress studies with Polar map display and MPRI
| 106 | · Delayed Enhancement studies with auto scar, edema and MVO analysis
| 107 | · T1 analysis with motion correction and registration and T1 and ECV maps generation
| 108 | · T2 analysis with motion correction and registration and T2 maps generation
| 109 | · T2* analysis with T2* maps generation and iron content reporting
| 110 | · 4D Flow visualization and analysis (Ventricular Flow, Pulse Wave Velocity, Wall Shear Stress, Energy Loss)
| 111 | SEQUENCES AND TECHNIQUES
| 112 | (Turbo/Fast) Spin-echo
| 113 | (Turbo/Fast) Gradient-echo
| 114 | TSE / FSE single-shot
| 115 | Steady-state pulse sequence
| 116 | Echo Planar Imaging (EPI) - single and multi-shot
| 117 | Inversion recovery (IR)
| 118 | Spectral fat suppression
| 119 | Separation of fat and water using the "Dixon" technique or similar ones
| 120 | Dual Echo
| 121 | Standard MR angiographic techniques: Time Of Flight, Phase Contrast, Contrast-Enhanced MRA
| 122 | Angiography without contrast with the option of imaging peripheral, renal and carotid arteries
| 123 | Dynamic (time-resolved) 3D MRA
| 124 | Bolus Tracking
| 125 | Elimination of artifacts due to patient movement for
| 126 | examinations of the abdomen, head and bone/joint system using the "radial k-space filling" technique in all three planes
| 127 | Diffusion Weighted Imaging (DWI)
| 128 | Diffusion tensor imaging (DTI) with tractography (Fiber Tracking) and minimum 150 directions
| 129 | Perfusion with contrast
| 130 | Perfusion without contrast, 3D ASL
| 131 | MR spectroscopy Singlevoxel , 2D and 3D Multivoxel techniques
| 132 | Susceptibility Weighted Imaging
| 133 | Reduction of the acoustic noise using the 'zero TE' or ultra short TE supports the usage of zero TE or ultra short TE techniques for neurological and MSK imaging - such as Silenz, Silent Suite, or equivalent - the technique must be specified
| 134 | Interactive real-time imaging
| 135 | Automatic examination of the head - (automatic plane setting, acquisition, processing and data transfer)
| 136 | Double Inversion Recovery Black Blood cardiovascular Imaging
| 137 | Triple Inversion Recovery Black Blood cardiovascular Imaging
| 138 | 2D Cine Cardiac Imaging
| 139 | 2D Flow Phase Contrast
| 140 | 4D Flow imaging
| 141 | Delayed enhancement cardiac imaging in 2D, 3D, Phase Sensitive and Single Shot
| 142 | Rest and Stress perfusion Imaging
| 143 | T1, T2 and T2* cardiac mapping techniques
| 144 | Non contrast enhanced coronaries angiography
| 145 | Prospective and retrospective gating
| 146 | Analysis of time curves (Time intensity curves analysis) - Option of fusing parameter maps with anatomical images
| 147 | Techniques for 3D post-processing :
| 148 | Multi Planar Reconstruction (MPR)
| 149 | Maximum Intensity Projection (MIP)
| 150 | Volume rendering (VRT)
| 151 | ADDITIONAL EQUIPMENT, ACCESSORIES AND WORKS
| 152 | Anti-magnetic patient trolley
| 153 | Oxygen monitor for the continuous monitoring of oxygen levels.
| 154 | MRI contrast media injector for use in 1.5T magnetic field with minimum two injection heads, one for contrast, the other for a saline solution, with maximum flow rate of minimum 10 ml/sec.
| 155 | RF cabin for system protection from electromagnetic interference (attenuation min. 80 dB): complete cabin with antistatic floor, door, window, lightning, all the necessary filters and all the finishing works, ready for work with patients
| 156 | Air Conditioning and Ventilation equipment for the examination room
| 157 | Air Conditioning equipment for the technical room
| 158 | Chiller for MR system with a closed cooling water system.
| 159 | Uninterrupted power supply (UPS), 5 pcs, one for the operator console four for server and client workstations
| 160 | Special power distribution cabinet for the MRI device
| 161 | Complete preparation of the premises where the MR system will be installed, including all civil works needed for installation of the system, its operation and full functionality, based on Turn-key solution.
| 162 | Application trraining on site for medical staff of duration of minimum 10 days.
| 163 | Application training in producer's cardiac reference site with same or similar system for 3 phisicians and 3 technicians in total of min. 5 days

## Scoring
BID EVALUATION
| CRITERIA FOR EVALUATION | Maximum No. of ponders
1. | LOWEST OFFERED PRICE | 70
This criterion is expressed according to the formula: Number of points of the bidder = (Lowest offered price) / (offered price) * 70
2. | TECHNICAL AND TECHNOLOGICAL ADVANTAGES | 30
This criterion is evaluated so that the number of points for each of the selected technical characteristics is determined below.
| TOTAL : | 100
SCORING SYSTEM FOR QUALITY EVALUATION
No. | Parametar/Characteristic | Maximum No. of ponders | Yes or No, or enter the value where required | Proof (in form of manufacturer statement or brochure or technical/product data)
1 | Magnetic field homogeneity in accordance with the V.R.M.S. method:DSV, for spherical volume with 30cm diameter: less than 0.15 ppm (typical value) | 2
Number of points of the bidder = (minimum offered value) x 2 points / (offered value)
2 | Magnetic field homogeneity in accordance with the V.R.M.S. method:DSV, for spherical volume with 40cm diameter: less than 0.65 ppm (typical value) | 2
Number of points of the bidder = (minimum offered value) x 2 points / (offered value)
3 | Magnetic field homogeneity in accordance with the V.R.M.S. method:DSV, for spherical volume with 50cm diameter: less than 5.5 ppm (typical value) | 2
Number of points of the bidder = (minimum offered value) x 2 points / (offered value)
4 | Minimum value for TR time in Turbo (Fast) Spin Echo technique and its acquisition matrix of 256x256 - not longer than 6.0 msec. | 2
Number of points of the bidder = (minimum offered value) x 2 points / (offered value)
5 | Minimum value for TR time in the EPI (Echo Planar Imaging) technology and its acquisition matrix of 256x256 - less than 10 msec. | 2
Number of points of the bidder = (minimum offered value) x 2 points / (offered value)
6 | Number of independent RF channels in the gantry in a single FOV, min 64 | 10
Number of points of the bidder = (offered value) x 10 points / (maximum offered value)
7 | Offered multi-channel single coil for thoracic, abdominal, pelvic and cardiac imaging which is requested in spec 5.2. with the following features: - Single coil, not combination of two, - Coverage (SI), minimum 60 cm in the z- direction, (or longitudinal direction) - Minimum 30 elements - Weight resting on the patient's body (weight of the coil itself and weight of all additions such as cable, plug, etc.), minimum 1.8 kg. | 10
No - 0 points, Yes - 10 points
| Total number of ponders | 30