## Price Schedule
Bidder: |  |  |  | Offered price (currency): | 0
LOT 6 | LABORATORY
ID | Equipment name (Technical Specification Requested) | QTY | Technical Specification Offered (model) | Unit price | Total price per line item
1 | ANTIPLATELET THERAPY RESPONSE ASSESMENT | 1 |  |  | 0
2 | BIOCHEMISTRY ANALYSER | 2 |  |  | 0
3 | BLOOD GAS ANALYSER | 2 |  |  | 0
4 | COAGULATION ANALYSER | 2 |  |  | 0
5 | FULLY AUTOMATED BLOOD BANK INSTRUMENT | 1 |  |  | 0
6 | FULLY AUTOMATED HEMOSTASIS ANALYZER | 1 |  |  | 0
7 | HEMATOLOGY ANALYZER small | 1 |  |  | 0
8 | HEMATOLOGY ANALYZER | 1 |  |  | 0
9 | IMMUNOASSAY ANALYZER HIGH THROUGHPUT | 1 |  |  | 0
10 | IMMUNOASSAY ANALYZER | 1 |  |  | 0
11 | Automatic blood glucose analyzer | 1 |  |  | 0
12 | Laboratory centrifuge | 4 |  |  | 0
13 | Refrigirated centrifuge | 2 |  |  | 0
14 | Plasma Thawer | 1 |  |  | 0
15 | Binocular microscope | 2 |  |  | 0
16 | General Purpose Laboratory Refrigerator 1 | 4 |  |  | 0
17 | General Purpose Laboratory Refrigerator 2 | 7 |  |  | 0
18 | General Purpose Laboratory Freezer (-18 to -25C) | 6 |  |  | 0
19 | Blood Storage Refrigerator (4° C) | 2 |  |  | 0
20 | Laboratory Freezer (-86°С) | 4 |  |  | 0
21 | Mini vortex | 1 |  |  | 0
22 | Analytical balance 1 | 1 |  |  | 0
23 | Laminar air flow cabinet | 1 |  |  | 0
24 | Laboratory fume hood | 1 |  |  | 0
25 | Barcode scanner and reader for test tubes | 2 |  |  | 0
26 | Blood Roller Mixer | 1 |  |  | 0
|  |  |  | Total Price per Lot: | 0

## LAB
Bidder:
LOT 6 |  | LABORATORY |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | ANTIPLATELET THERAPY RESPONSE ASSESMENT |  | 1 |  | 0
| 1 | Automated test platform
| 2 | Monitor platelet response to: Aspirin and P2Y12 inhibitors
| 3 | Measures platelet reactivity via light transmission aggregometry (LTA)
| 4 | Uses whole blood, closed tube samples with no pipetting or sample preparation
| 5 | Specific assessment of P2Y12 inhibition with Prostaglandin E1
| 6 | Sample tube is inserted directly on test device
| 7 | Interpret with single-use test device
| 8 | Test device contains all necessary reagents
| 9 | Storage temperature 2 – 25 °C
| 10 | Testing at point-of-care
2 |  | BIOCHEMISTRY ANALYSER |  | 2 |  | 0
| 1 | Throughput minimum 1200 tests per hour including electrolytes
| 2 | Measurement methods: photometry, potentiometry, turbidimetry
| 3 | Photometry: Minimum 11 different wavelengths
| 4 | Minimal reaction volume less than 130 μL
| 5 | Minimal sample volume 1.5 μL or less
| 6 | Volume for all three electrolytes (Na, K, Cl) less than 21 μL
| 7 | Minimum 60 different simultaneously processed analytes on board
| 8 | Minimum 100 positions for reagents
| 9 | Ability to load reagents without stopping the system
| 10 | Minimal sample capacity 120 positions with ability for continuous loading
| 11 | Cooled positions for reagents, calibrators and controls
| 12 | Ability to measure serum indices - lipemia, icterus, hemolysis
| 13 | Clot detection
| 14 | Premanant glass reaction cuvettes with automatic washing
| 15 | Automatic rerun option for samples outside linearity range
| 16 | Ability to run emergency samples
| 17 | Ability to program automatic power-on of analyser
| 18 | Software based on Windows operating system
| 19 | Bidirectional communication with LIS
| 20 | Ability to integrate system with immunochemistry analyser
3 |  | BLOOD GAS ANALYSER |  | 2 |  | 0
| 1 | Cartridge based system for blood gas testing
| 2 | Minimum test panel: pH, pCO2, pO2, K+, Na+, Ca2+, Cl-, Glu, Lac, Hct, tHb, O2Hb, COHb, MetHb, HHb, sO2, HCO3-, TCO2, BE(B), BEecf, Anion Gap.
| 3 | Disposable, all-in-one single multi-use cartridge includes all testing components: sensors, CO-Ox optical cell, lysing solution, controls, tubing, waste bag, sampler.
| 4 | Time to results: up to 45 sec from sample aspiration
| 5 | Sample type: heparinized whole blood (arterial, venous, capillary)
| 6 | On-board use-life for cartridge: up to 31 days
| 7 | Identification of errors including micro-clots and interference and flagging analytes if an error is detected
| 8 | Active quality process control program designed to provide continuous monitoring of the analytical process before, during, and after sample measurement with real-time, automatic error detection, automatic correction of the system and automatic documentation of all corrective actions.
| 9 | Selection of analytes depending on the needs of the customer
| 10 | Room-temperature storage of cartridge.
| 11 | Zero maintenance – no hands-on troubleshooting or corrective actions required
| 12 | Integrated printer
4 |  | COAGULATION ANALYSER |  | 2 |  | 0
| 1 | Fully automated, bench top analyzer with eight optical reading channels for coagulometric, chromogenic and immunological measurements.
| 2 | Optical-based clot detection system
| 3 | “Random access” analyzer
| 4 | Closed-Tube Sampling via cap-piercing system
| 5 | STAT samples loaded on any rack, in any position, at any time
| 6 | Assay-specific HIL (Hemolysis, Bilirubin and Lipemia) sample check
| 7 | Sample aspiration clog check and flagging abnormal sample aspiration, prompting user to check for clots
| 8 | Tube-fill volume check
| 9 | Continuous loading and unloading of samples and reagents, cuvette loading and waste disposal with no system interruption.
| 10 | Barcode reader for automatic identification of samples and reagents.
| 11 | Automatic measurement of on-board volume of the reagent
| 12 | Fully automated detection of total immunoglobulin against platelet factor 4-heparin complexes
| 13 | Rerun and reflex testing
| 14 | Fully automated reporting of factor assays with Parallelism
| 15 | LIS connection
| 16 | Touch screen display monitor
5 |  | FULLY AUTOMATED BLOOD BANK INSTRUMENT |  | 1 |  | 0
| 1 | Fully automated system, true continuous & random access system: allow continuous, uninterrupted access to samples, reagents, microplates.
| 2 | Should allow continuous refilling of system liquid (without interruption) and waste removal
| 3 | Should use microplate technology based on solid phase red cells adherence technology
| 4 | Should use IgG specific technology
| 5 | Should be able to perform antibody screening, antibody identification, crossmatch, blood grouping, Rh phenotyping, DAT, Weak D.
| 6 | Should be able to do platelet antibody crossmatch and screening, antigen screening, TPHA
| 7 | Should have facility for automatic weak D testing of Rh negative samples-reflex test
| 8 | Should be able to load more than 200 samples or more at any given point in time
| 9 | Should have different security access levels for staff with individual passwords
| 10 | Should be able to run pooled cells-, 2 cell- and 3 cell Antibody screening
| 11 | Should have antibody screening and identification panels with shelf life of more than 12 weeks
| 12 | Should have integrated waste containers
| 13 | Should have clot detection, liquid level detection, automatic agitation of red cells and reagent inventory monitoring system
| 14 | Should have STAT functionality for emergency samples
| 15 | Should be able to check on board reagent inventory before starting the run and alert in case of absence of reagents
| 16 | Should have an incorporated wash step for processing of hemolyzed, lipemic or icteric samples
6 |  | FULLY AUTOMATED HEMOSTASIS ANALYZER |  | 1 |  | 0
| 1 | Proven thromboelastometry technology based on a fixed cylindrical cup and permanently oscillating vertical axis
| 2 | Four channels
| 3 | Fully automated, no pipetting and preparation
| 4 | Sample tube fits directly into cartridge, avoiding blood handling
| 5 | Simple cartridge system minimizes operator involvement
| 6 | Large integrated touchscreen allows easy viewing of results
| 7 | Cartridges stored at room temperature
| 8 | Results of coagulation status can be available within 10 minutes
| 9 | System designed for in vitro diagnostics in the Point of Care field and laboratories
| 10 | Cartridges detect:
| 11 | Factor deficiency
| 12 | Discriminate between thrombocytopenia and fibrinogen deficiency or fibrin polymerization disorders
| 13 | Hyperfibrinolysis
| 14 | Direct FX inhibitors and direct thrombin (FIIa) inhibitor effects
| 15 | Cartridge portfolio contains five tests offered in two test combinations
| 16 | Available tests:
| 17 | INTEM C – Fast assessment of clot formation, fibrin polymerization and fibrinolysis via the intrinsic pathway
| 18 | EXTEM C – Fast assessment of clot formation, fibrin polymerization and fibrinolysis via the extrinsic pathway
| 19 | FIBTEM C – Fast analysis without platelets; qualitative assessment of fibrinogen status
| 20 | APTEM C – In-vitro fibrinolysis inhibition; assessment of the possible effect of antifibrinolytic drugs
| 21 | HEPTEM C – Specific detection of heparin when compared with INTEM C via heparin neutralization
| 22 | Sigma complete
| 23 | FEBTEM C, EXTEM C, INTEM C, APTEM C
| 24 | Complete + hep
| 25 | FIBTEM C, EXTEM C, INTEM C, HEPTEM C
7 |  | HEMATOLOGY ANALYZER small |  | 1 |  | 0
| 1 | 1. 5-part diff hematology analyzer
| 2 | Whole blood aspiration volume up to 12 Μl
| 3 | 3 required reagents, cinanide- free, azide- free and formladehide- free
| 4 | Single diliution for leucocyte differentiation
| 5 | Paremeteres: WBC, RBC, HGB, HCT, MCV, MCH, MCHC, RDW-SD, RDW-CV, PLT, MPV, PCT, PDW, LY%, LY#, MO%, MO#, NE%, NE#, EO%, EO#, BA%, BA#
| 6 | Open vial sampling
| 7 | Throughput of 60 samples per hour
| 8 | Automatic aspiration probe cleaning
| 9 | Automatci Start Up
| 10 | Worklist option
| 11 | Windows oprative system
| 12 | Automatic Quality Assurance: control charts, limits, XB i XM controls
| 13 | Interlaboratory quality assurance program option
| 14 | Data base– capacity to store patient results including differential scatter plots and histograms
| 15 | Graphic representation for all cell populations- scatter plots and histograms
| 16 | LIS connectivity
8 |  | HEMATOLOGY ANALYZER |  | 1 |  | 0
| 1 | The newest hematology analyzer from manufaturer's portfolio
| 2 | 5-part diff stand-alone hematology analyzer
| 3 | Throughtput of 100 samples per hour with autoloader capasity of 100 samples minimum
| 4 | Whole blood aspiration volume- up to 170 μL
| 5 | Dead volume up to 0.5mL
| 6 | Single aspiration pathway- single proble for all sample types
| 7 | Auto rerun and reflex testing available
| 8 | Impedance method for RBC, WBC and PLT counting
| 9 | Automatic NRBC counting
| 10 | Automatic reticulocyte counting without manual sample preparation
| 11 | Reagent status available at any time
| 12 | Quad diluent configuration
| 13 | 4 required reagents; environmentally friendly methanol- and ethlylene-glycol-free
| 14 | Modular system; optional integration with at least one hematology analyzer and/or slide maker and stainer system
| 15 | IVD parameter for early sepsis detection
| 16 | LIS connectivity
9 |  | IMMUNOASSAY ANALYZER HIGH THROUGHPUT |  | 1 |  | 0
| 1 | Completely automated immunoassay analyzer with throughput 200 tests/hour
| 2 | Continuous loading samples, consumables and reagents without stoping the system
| 3 | On board refrigerator storage (4 do 10 ˚C) of 50 position for reagents
| 4 | Open-pack stability (assay dependent) 14-64 days
| 5 | Calibration curve stability (assay dependent) 14-63 days
| 6 | Possibility to load primary tube directly to the system
| 7 | Primary tube releases after maximum 25 second (aspiration time)
| 8 | Calibration curves and parameters displayed on screen and printouts
| 9 | Monitoring quality control over time (Levey-Jannings chart).
| 10 | Bi-directional communication and connection to the LIS
10 |  | IMMUNOASSAY ANALYZER |  | 1 |  | 0
| 1 | Completely automated bench-top immunoassay analyzer with throughput minimum 100 tests/hour
| 2 | Continuous loading samples and reagents without stoping the system
| 3 | On board refrigerator storage (4 do 10 ˚C) of minimum 20 position for reagents
| 4 | Open-pack stability (assay dependent) 14-64 days
| 5 | Calibration curve stability (assay dependent) 14-63 days
| 6 | Possibility to load primary tube directly to the system
| 7 | Samples: serum, plasma, urine.
| 8 | Calibration curves and parameters displayed on screen and printouts
| 9 | Monitoring quality control over time (Levey-Jannings chart).
| 10 | Bi-directional communication and connection to the LIS
11 |  | Automatic blood glucose analyzer |  | 1 |  | 0
| 1 | Automated device for glucose measurement
| 2 | For blood, plasma or serum sample
| 3 | Min 20 places for samples
| 4 | Touch screen user interface
| 5 | Possibility to connect to hospital or lab information systems (LIS).
12 |  | Laboratory centrifuge |  | 4 |  | 0
| 1 | Max. capacity 4 x 100 ml / 32 x 15 ml
| 2 | Max. RPM (speed) 6,000 min-1
| 3 | max. RCF 4,226
| 4 | Running time 1 – 99 min
| 5 | Ergonomic operating and information panel
| 6 | Display switchable from RCF to RPM
| 7 | Lid locking and holding
| 8 | Lid dropping protection
| 9 | Automatic rotor recognition
| 10 | Imbalance switch-off
| 11 | Emergency lid lock release
| 12 | Centrifuging chamber of stainless steel
| 13 | Swing-out rotor - 90 degrees
| 14 | Rotor capacity 68 x (12x75 mm) 5 mL tubes
13 |  | Refrigirated centrifuge |  | 2 |  | 0
| 1 | Maximum speed 15000 RPM
| 2 | Maximum RCF 21382
| 3 | Microprocessor control, LCD display
| 4 | Ergonomic operating and information panel
| 5 | Switches from RPM to RCF
| 6 | Possibility of manually stopping centrifugation
| 7 | Aceeleration and braking ramps
| 8 | Selectable program memory
| 9 | Temperature range: -20°C to +40°C
| 10 | Running time 1-99 min
| 11 | Fixed angle rotor – 45 degrees
| 12 | Rotor capacity 32 x (16x100) tubes
| 13 | Centrifuging chamber of stainless steel
| 14 | Maximum capacity 4x200 mL
14 |  | Plasma Thawer |  | 1 |  | 0
| 1 | Thaws FFP bags and warms blood, erythro-cyte concentrates (EC), cryopreserved preparations, cryopreserved stem cells (HPC) and infusion solutions
| 2 | Thaws 4bags at a time
| 3 | Dry warming (no direct contact to water)
| 4 | Leak detection sensors
| 5 | Large display with plain text
| 6 | No risk of contamination due to dry thawing and warming (no direct contact to water)
| 7 | Safety Electronic and mechanical overtemperature alarm
| 8 | Adjustable temperature range from +37°C to +45°C
| 9 | Visual and acoustic signal after heating duration is complete or in case of error
15 |  | Binocular microscope |  | 2 |  | 0
| 1 | Microscope with LED illumination
| 2 | Fine focus drive
| 3 | Adjustable torque for focus drive
| 4 | 7 objectives nosepiece
| 5 | Ceramic coated microscope stage
| 6 | XY stage control telescopic, torque adjustable,
| 7 | Condenser 0.90/1.25 Oil PH,DF
| 8 | Tube lens focal lenght more then 190mm,
| 9 | Objectives fluotar: 5x/0.15; 10x/0.32; 20x/0.55; 40x/0.80;
| 10 | Eyepieces with 10x magnification and 25mm field of view, both with diopter ring
| 11 | Dust cover
| 12 | Digital color camera - same manufacturer as microscope
| 13 | Live image: Full HD resolution 1920x1080 pixels with 30 frames per second
| 14 | Captured image: resolution 2592x1944p, 5Mpixels
| 15 | Video recording: HD resolution 1920x1080 pixels
| 16 | Remote control for camera
| 17 | Photo and video recording directly on SD card
| 18 | PC brand name, procesor I3, 8GB, HD 1TB, Win 10 64BIT
| 19 | Full HD LED Monitor 24", IPS
16 |  | General Purpose Laboratory Refrigerator |  | 4 |  | 0
| 1 | General purpose laboratory refrigerator
| 2 | With solid doors
| 3 | Volume 150-160 liters
| 4 | Temerature range: from +1C to + 10C
| 5 | Cabinet Material Corrosion-resistant
| 6 | Automatic defrost
| 7 | HCFC- and CFC-free insulation
| 8 | Includes 3 shelves, 1 basket
| 9 | Digital display, integrated temperature controller
| 10 | High- and low-temperature alarms, Door-ajar alarm
| 11 | Standard door locks, reversible door
| 12 | External dimensions: 600x600x850mm ± 10%
17 |  | General Purpose Laboratory Refrigerator |  | 7 |  | 0
| 1 | General purpose laboratory refrigerator
| 2 | With solid doors
| 3 | Volume 250-300 liters
| 4 | Temerature range: from +1C to + 10C
| 5 | Cabinet Material Corrosion-resistant
| 6 | Automatic defrost
| 7 | HCFC- and CFC-free insulation
| 8 | Includes 5 shelves, 1 basket
| 9 | Digital display, integrated temperature controller
| 10 | High- and low-temperature alarms, Door-ajar alarm
| 11 | Standard door locks, reversible door
| 12 | External dimensions: 600x650x1600mm ± 10%
18 |  | General Purpose Laboratory Freezer (-18 to -25C) |  | 6 |  | 0
| 1 | With solid doors
| 2 | Volume 150-160 liters
| 3 | Temperature range: from -18 to -25C
| 4 | Cabinet Material Corrosion-resistant, spark-proof
| 5 | Automatic defrost
| 6 | HCFC- and CFC-free insulation
| 7 | Standard acces port
| 8 | 3 baskets
| 9 | Integrated temperature controller
| 10 | Digital display
| 11 | High- and low-temperature alarms, Door-ajar alarm
| 12 | Standard door locks, reversible door
| 13 | External dimensions: 600x600x850 ± 10%
19 |  | Blood Storage Refrigerator (4° C) |  | 2 |  | 0
| 1 | High Performance Blood Bank Refrigerator
| 2 | Temperature range 3-7C. Factory Temperature Setpoint 4°C
| 3 | Capacity 385 blood bags (450ml)
| 4 | Refrigeration: Vapor compression system Forced-Air Cooled
| 5 | Defrost - automatic
| 6 | Insulation 5.08 cm - High Density HFC-blown Polyurethane Foam
| 7 | Self-closing glass door
| 8 | Doorlock
| 9 | Four casters for easy mobility; the front two are lockable
| 10 | Standard, built-in chart recorder
| 11 | Graphic thermometer shows high or low temperature, flashing bar top bar for overtemperature
| 12 | Control Sensor im,ersed i ethylen glycole solution
| 13 | Microprocessor controller
| 14 | Remote Alarm Terminals (set highest and lowes temperature)
| 15 | Alarm for Battery backup, Door ajar
| 16 | Dimensions 2000 H x 950D x 700 W mm ± 10%
| 17 | Noise level 52 dB
| 18 | GMP Clean Room Class A / ISO 6 (ISO EN 14644-1) compatible with appropriate pre-install preparation
20 |  | Laboratory Freezer (-86°С) |  | 4 |  | 0
| 1 | Up-right freezer, nominal capacity: 600-650l
| 2 | Temperature range: -40 °C / -86 °C
| 3 | Supply voltage: 230V/50Hz.
| 4 | Structure and door: Internal chamber in stainless steel
| 5 | Exterior in steel coated with anti-corrosive film.
| 6 | Models are provided with castors for easy moving.
| 7 | Isolation: 127-140 mm HCFC free foam and gas,
| 8 | Cooling air
| 9 | Pressure equalization valves to reduce ice production.
| 10 | 1 full exterior doors of the same structure as the rest of the freezer coated with magnet tires, interior of steel coated with 4-5 shelves and 4-5 separate compartments with internal, separate doors. An ergonomic handle that allows to open the door with one hand
| 11 | Regulatory and control system:
| 12 | Controller equipped with 3 independent microprocessors: one regulating the cooling system, the other that regulates the alarm system and the third one that regulates the data transmission.
| 13 | Access to data is protected by a three-level access code (user, service, and administration).
| 14 | Color display, size 7 "with display of set temperature, chamber temperature and alarm status.
| 15 | Ecomode function that allows the operation of the freezer to save energy.
| 16 | Real-time function that allows daily display of temperature variation.
| 17 | Automatically store temperature data
| 18 | Battery back up for a minimum of 30 hours in case of power failure.
| 19 | Equipped with a built-in CO2 safety system, which automatically turns on in the event of a power failure and maintains a temperature range of -40-70 ºC for 48 hours.
21 |  | Mini vortex |  | 1 |  | 0
| 1 | Speed range: 0-3000rpm
| 2 | Operating mode – Continual / Pulse
| 3 | Maximal load 0,5kg
| 4 | Built-in safety features for overload
| 5 | Possibility of adding different platforms for tubes and microplates
22 |  | Analytical balance 1 |  | 1 |  | 0
| 1 | Tarring throughout the entire weighing range.
| 2 | LCD display.
| 3 | Sturdy white plastic housing with stainless steel weighing platform.
| 4 | Weighing range 0 -5 000 gr.
| 5 | Accuracy: 1 gr.
23 |  | Laminar air flow cabinet |  | 1 |  | 0
| 1 | Working place -two seats
| 2 | Micro processor maintains steady blower speed..
| 3 | Airflow alarms.
| 4 | Tempered glass side walls.
| 5 | H14 HEPA filters providing superior product protection.
| 6 | Pre-filter EU3.
| 7 | Average airflow velocity (adjustable): 0.45 m/s.
| 8 | The air cleanliness within the work zone is better than ISO Class 5
| 9 | The 10° sloped front.
| 10 | Antimicrobial coating on all painted surfaces minimizes contamination.
| 11 | 2 electrical outlets.
| 12 | Apparatus for UV lamp with timer (UV lamp and protective cover included).
| 13 | Base included.
24 |  | Laboratory fume hood |  | 1 |  | 0
| 1 | Fume cupboard for the containment and removal of toxic vapours and aerosol, providing operator safety
| 2 | The bench-mounted fume cupboard consists of a self-supporting underbench unit or a supporting steel frame and an upper part
| 3 | Made of steel, covered with Epoxy coating
| 4 | Safety glass front window
| 5 | With carbon filters and prefilter
| 6 | Microprocessor controlled monitoring system
| 7 | Air flow vol/air treated m3/h 600
| 8 | Cabinet is supplied with acid resistant PVC work surface
| 9 | In compliance with EN 14175.3 standards
| 10 | Interior lighting
| 11 | Two waterproof sockets
| 12 | Widith of digestor min 90cm
25 |  | Barcode scanner and reader for test tubes |  | 2 |  | 0
| 1 | Barcode scanner for healtcare purposes, medication administration to blood verification and administration, specimen collection, admission, inventory managemen
| 2 | Blutooth connection
| 3 | Must have table holder
| 4 | Direct thermal printer for healthcare purposes
| 5 | Direct thermal (d) and Thermal transfer (t) printing ofbar codes, text, and graphics
| 6 | 203 dpi print resolution (8 dots/mm)
| 7 | Maximum print width is 104 mm
| 8 | IEC 60601-1 Certified
26 |  | Blood Roller Mixer |  | 1 |  | 0
| 1 | For blood and other sample mixing
| 2 | Adjustable speed
| 3 | Able to use with different tube diameter