## Lot 7 Tele
Bidder: |  |  |  | Offered price: | 0
LOT 7 | TELEMETRIC PATIENT MONITORING SYSTEM
ID | Equipment name | QTY | Technical Specification Offered(model) |  | Total Price per line item
1 | Telemtreic Patient Monitoring System with 60 months license | 1 |  |  | 0
|  |  |  | Total Price per Lot: | 0

## Telemetric Patient Monitoring
Bidder:
LOT 7 |  |  |  |  | Date:
[5 empty] | Total DAP price: | 0
Line item No. | ID | Technical specification requested | Technical Specification offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in technical documentation
1 |  | TELEMETRIC PATIENT MONITORING SYSTEM |  | 1 |  | 0
| 1 | The system for real time online patient telemonitoring should include the wearable units for monitoring vital parameteres and access to the software/interface for receiving, following and analysing patients data
| 2 | Number of patients which have to be monitored:120
| 3 | Devices for monitoring vital parameters-120 pcs
| 4 | Parameters that should be monitored in real time:
| 4.1 | ECG (minimum 12-lead ECG)
| 4.2 | Body temperature
| 4.3 | Blood pressure
| 4.4 | SpO2
| 4.5 | Heart rate
| 4.6 | Respiration rate
| 4.7 | Body position and activity
| 4.8 | Should have Built-in stethoscope
| 5 | Monitoring units have to be mobile
| 6 | System should include:
| 6.1 | Wearable patient unit-monitoring device
| 6.2 | Transmitter
| 6.3 | Baterries
| 7 | The system should records the collected signals and values in the internal memory of monitor device for min 36 hours with one battery set
| 8 | Data transfer: 3G, 4G, GPRS, WiFi
| 9 | deleted
| 10 | All the patients phisiological data is constantly monitored and stored automatically with cloud based platform
| 11 | Online access to the system through PC, tablet, cellular devices with user account
| 12 | The number of users/staff with access to the system should not be limited
| 13 | Visualization interface which allows to be observed up to 60 patients in one single screen
| 14 | System should include alert system for adverse conditions
| 15 | Automatic report system, displaying the trend of the vital signs as well as detection of certain cardiac conditions of the patients
| 16 | The visualization interface for history data of the vital parameter and tools for measurements of the ECG details (amplitude, intervals)
| 17 | Intelligent self adapting alarm system based on NEWS 2 standard rules
| 18 | Multivariate machine learning model for prediction of deterioration based on the input of different parameters, calculated continuously
| 19 | Trend analysis algorithm for follow up of the vital parameters trend and directions
| 20 | Automatic ranging of the patients according their risk in the operator's interface in order to highlight the priority of the each patient
| 21 | Patient historical data
| 22 | Administrative form for patient input as well as lab results, image diagnostics, life vital parameters and comorbidity data
| 23 | AI- driven pathology patterns detection
| 24 | System calibration, licenses and software update
| 25 | Monthly system calibration and software license (including cloud space and software updates), irrespective of the number of users, for period of 60 months after delivery and device activation have to be included in the price