## PriceSchedule
Bidder: |  |  |  | Offered price:
LOT 1_Endoscopic Surgery and Diagnostics
Item No. | Equipment name | Qty | Model, Manufacturer and Country of Origin | Unit price | Total Price per line item
1 | Laparoscopy 4K tower with system for inteligent CO2 flow control | 1
2 | Arthroscopy tower with FESS accessories | 1
3 | Electric tourniquet pump | 1
4 | Video Laringoscopy sistem with flexibile nasolaringosope | 1
5 | Videoendoscopic tower with gastroscope and colonoscope | 1
6 | Instrument set for cystoscopy with LED light source | 1
7 | Urology holmioum laser | 1
8 | Pneumatic lithotriptor | 1
9 | Uretherorenoscope set | 1
10 | Electrorenoscope with histeroscopy instrument set | 1
11 | Flexible ureterorenoscope | 1
12 | Basic rectroscopy instrument set | 2
13 | Endourology visualisation system with histeroscopy pump | 1
14 | Rigid endoscopes 2,7 mm, 180mm lenght for ENT department | 2
Total price per Lot |  |  |  |  | 0

## TechSpecification
Bidder: | [5 empty] | Total DAP price: | 0
LOT 1_Endoscopic Surgery and Diagnostics
Item No. | ID | Technical Specification Requested | Technical Specification Offered | Qty | Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Laparoscopy 4K tower with system for inteligent CO2 flow control |  | 1 |  | 0
| 1.1 | Integrated Visualization System comprises of camera control unit integrated in monitor with minimal display resolution 4096x2160 and minimum size of 32"
| 1.2 | Video outputs including at minimum UHD3840x2160p, display port 1.2, HDMI 3G-SDI
| 1.3 | 3 CMOS camera head compatible with 4K monitor resolution, with programmable function, autoclavable or equivalent
| 1.4 | Electronic LED light guide or equivalent |  | 2
| 1.5 | Image capture system with minimal resolution 1920x1080p and 1TB internal memory. Data output by minimum 4 x USB 3.0. HDMI video input/output. Touch screen data input
| 1.6 | Equipment cart with minimum 4 shelves, adjustable height, flex monitor holder, and mount possibility for secondary monitor
| 1.7 | Insufflator with minimum CO2 flow of 40l/min and minimum 3 adjustable flow level. Minimum set pressure 15 mmHg. Different modes including at least mode with continuous smoke and aerosol evacuation
| 1.8 | Suction-irigation pump with flow rating 2000ml/min and set pressure up to 150mmHg
| 1.9 | Laparoscope dia 10mm 0° compatible with resolution of camera control unit
| 1.10 | Laparoscope dia 10mm 30° compatible with resolution of camera control unit
| 1.11 | Consumables include: tri-lumen tube set with integraded ULPA filter, tri-lumen port dia 5-12mm and min lenght 100mm, single lumen insufflation tube set, suction-irigation tube sets, irrigation cannula
2 |  | Arthroscopy tower with FESS accessories |  | 1 |  | 0
| 2.1 | Integrated Visualization System comprises of camera control unit integrated in monitor with minimal display resolution 4096x2160 and minimum size of 32"
| 2.2 | Video outputs including at minimum UHD3840x2160p, display port 1.2, HDMI 3G-SDI
| 2.3 | 3 CMOS camera head compatible with 4K monitor resolution, with programmable function, autoclavable or equivalent
| 2.4 | Electronic LED light guide or equivalent |  | 2
| 2.5 | Image capture system with minimal resolution 1920x1080p and 1TB internal memory. Data output by minimum 4 x USB 3.0. HDMI video input/output. Touch screen data input
| 2.6 | ENT motor control unit with color display - touch screen, with accesores: Two-Pedal Footswithch Shaver handpiece with cleaning adaptor, Max. 10,000 rpm for shaver blades, max. 12,000 rpm for sinus shaver, Straight suction channel, Integrated irrigation channel, Adjustable handle Minimum seven functions: Shaver system for surgery of the paranasal sinuses and anterior skull base; INTRA Drill Handpieces (40,000 rpm and 80,000 rpm); Sinus Shaver; Micro Saws; Intranasal Drill; Dermatome; High-Speed Handpieces (60,000 rpm and 100,000 rpm) Two motor outputs: Two motor outputs enable simultaneous connection of two motors: For example, a shaver and micro motor Soft start function ntegrated irrigation and coolant pump Maximum number of revolutions and motor torque: Microprocessor-controlled revolutions per minute; the preselected parameters are maintained all the time during drilling Maximum number of revolutions can be preset
| 2.7 | Roller pump for suction or irrigation, with ENT/NEURO softver for cleaning telescope lens, with acessories: irigation tubes FC, reusable, sterilizable, Pump Tube, consumables/accessories, package of 10, Membrane, consumables/accessories, package of 20, Tubing Set, Irrigation, CV, for single use, sterile, 5 packages of 10
| 2.8 | Equipment cart with minimum 4 shelves, adjustable height, flex monitor holder, and mount possibility for secondary monitor
| 2.9 | Arthroscope dia 4mm 30° deg with sheat compatible with resolution of camera control unit
| 2.10 | Arthroscope dia 4mm 0° deg with sheat compatible with resolution of camera control unit
| 2.11 | Suction-irigation pump with flow rating 2000ml/min and set pressure up to 150mmHg
| 2.12 | Arthroscopic resection system with two ports and displays and shaver handpiece with minimum forward speed 12000 rpm and oscillate speed 2500cpm
| 2.13 | Set of standard arthroscopic instruments diametar 3.4 or equivalent including linear, wide bite, left and right punch, scissors, aligator grasper and sterilization case
| 2.14 | FESS onsumables include: Tray, for Cleaning, Sterilization and Storage of Accessoires, Shaver Blade, serrated cutting edge, 4 mm, 12 cm, sterile, 2 pcs, Shaver Blade, double serrated cutting edge, 4 mm, 12 cm, sterile, 2 pcs, Shaver Blade, straight cutting edge, 4 mm, 12 cm, sterile, 2 pcs, Shaver Blade, curved 35°, 4 mm, 12 cm, sterile, 2 pcs, Suction Shaver Blade, curved 40°, 4 mm, 12 cm, sterile, 2 pcs, Shaver Blade, curved 40°, sterilizable, 4 mm, 12 cm, sterile, 2 pcs, Sinus Burr, curved 40°, diamond head, 4 mm, 12 cm, sterile, 2 pcs
| 2.15 | Consumables include: suction-irigation tube sets and shaver blades and burs
3 |  | Electric tourniquet pump |  | 1 |  | 0
| 3.1 | Electric tourniquet pump, dual port, personalized, with touch-screen user interface and limb occlusion pressure sensor. Back-up battery power supply
| 3.2 | Reusable contour cuffs for arm, thigh, lower leg and IVRA
| 3.3 | Reusable limb sleeve for skin protection in minimum three sizes
4 |  | Video Laringoscopy sistem with flexibile nasolaringosope |  | 1 |  | 0
| 4.1 | HD Videolaryngoscope,3.6mm insertion tube,angulation range 130°/130° with enhancement of visibility of capillaries and other structures on the mucosal surface by using the narrow band observation light
| 4.2 | Video system with integrated LED light, enhancement of visibility of capillaries and other structures on the mucosal surface by using the narrow band
| 4.3 | FULL HD LCD Monitor min. size of 27"
| 4.4 | Trolley with integrated LCD monitor arm
| 4.5 | Leakage tester
5 |  | Videoendoscopic tower with gastroscope and colonoscope |  | 1 |  | 0
| 5.1 | Video processo minimum image quality HD, LED light source, min. 3 light sources, Output: min. 2x DVI, 1x RGB, 1x S Video, 1x Video
| 5.2 | Compatible with CCD and CMOS chip camera
| 5.3 | Control all CPU functions via the front panel, ability to connect and control the processor via the keyboard
| 5.4 | External memory connector for storing on USB
| 5.5 | Possibility to connect a digital printer and digital DVD recorder
| 5.6 | Possibility to memorize: min. 45 patients, min. 20 clinical procedures, min. 20 attending physicians, DICOM interface incorporated in the processor
| 5.7 | Possibility to increase contrast in red color to improve visibility of abnormalities, inflamation and delineation
| 5.8 | Posilibility to combine special light wavelengts to improve contrast imaging for characterisation and improved visuelization of superficial vascular and mucosal paterns
| 5.9 | Posibility to enhance slight colour differences (vascular and mucosal) without tissue staining by digitally selecting three wavelenghts of light
| 5.10 | Anti blur function (automatically selects the highest quality image from multiple images)
| 5.11 | Service life of LED diodes minimum 10.000 working hours
| 5.12 | Air Supply Pump - Integrated irrigation pump with at least 4 operating modes
| 5.13 | Video gastroscope with CMOS chip positioned directly in the tip
| 5.15 | Field of view min. 140 °, Depth of field of view: min. 2-100 mm
| 5.16 | Distal diameter: max. 9.2 mm, Insertion tube diameter : max. 9.3 mm
| 5.17 | Working channel diameter: min. 2.8 mm, Working length up to 1100 mm
| 5.18 | Total length up to 1400 mm
| 5.19 | Bending capacity up/down min. 210°/ min. 90 °, Bending capacity left/right min. 100°/ min. 100 °
| 5.20 | Additional water channel
| 5.21 | Compatible with chromatographic observation of tissues described under 1.11, 1.12 and 1.13
| 5.22 | Video colonoscope with CMOS chip positioned directly in the tip
| 5.24 | There is no special electrical connection to the processor
| 5.25 | Field of view min. 170 °, depth of field of view min. 2-100 mm
| 5.26 | Diameter of the distal part max. 12 mm, insertion tube diameter max.12 mm
| 5.27 | Working channel diameter min. 3.8 mm, working length min. 1650 mm
| 5.28 | Bending capacity up/down min. 180 °/ min. 180 °, bending capacity left/right min. 160 ° / min. 160 °
| 5.29 | Additional water channel
| 5.30 | Compatible with chromatographic observation of tissues described under 1.11, 1.12 and 1.13
| 5.31 | LCD medical monitor min. 27", Native Resolution : min. 1920x1080
| 5.32 | Contrast ratio: min. 1000:1, Viewing angle min: 178°
| 5.33 | Interface min.: DVI-D / 3G-SDI / RS-232C
| 5.34 | Mobile stand for endoscopic equipment with three shelves and a movable monitor holder, sliding keyboard shelf, Endoscope carrier
| 5.35 | Mains voltage and earth connection.
| 5.36 | Suction unit Max. Vacuum min. 90 kPa, Suction between 40 and 45 liters/min
| 5.37 | Bottle min. 2l
6 |  | Instrument set for cystoscopy with LED light source |  | 1 |  | 0
| 6.1 | Sheath for cystoscope 16Fr., Obturator for cystoscope 16Fr.
| 6.2 | Sheath for cystoscope 17.5Fr., Obturator for cystoscope 17.5Fr.
| 6.3 | Sheath for cystoscope 19.5Fr., Obturator for cystoscope 19.5Fr.
| 6.4 | Sheath for cystoscope 21Fr., Obturator for cystoscope 21Fr.
| 6.5 | Adapter for cystoscope for diagnostic |  | 2
| 6.6 | Adapter with intoduction port |  | 2
| 6.7 | Working insert for cystoscope 30/70 ° with Albarran-lever |  | 1
| 6.8 | Flexible graspin forceps 7Fr., fine-toothed, double action, reusable |  | 1
| 6.9 | Flexible biopsy forceps 7Fr., double action, reusable |  | 1
| 6.10 | Flexible graspin forceps 10Fr., fine-toothed, double action, reusable |  | 1
| 6.11 | Flexible biopsy forceps 10Fr., double action, reusable |  | 1
| 6.12 | Fiber optic light cable, diameter min. 2.5 mm, length min.230 cm |  | 2
| 6.13 | HD Telescope 0°, diam.4mm, WL ≈300mm |  | 1
| 6.14 | HD Telescope 30°, diam.4mm, WL ≈300mm |  | 2
| 6.15 | HD Telescope 70°, diam.4mm, WL ≈300mm |  | 2
| 6.16 | Sterilization container |  | 3
| 6.17 | Led Light Source 180W: Working life of lamp approx. 30.000h -Brightness control 0-100% |  | 1
7 |  | Urology holmioum laser |  | 1 |  | 0
| 7.1 | Output min.30W
| 7.2 | Touch screen display
| 7.3 | For laser lithotripsy of bladder, ureter and kidney stones, ablation and coagulation
| 7.4 | Holmium YAG laser
| 7.5 | Wavelength 2100nm
| 7.6 | Energy range 200-5000mJ
| 7.7 | Frequency range 3-25 Hz
| 7.8 | Pulse length 150-800 μs
| 7.9 | Laser fibre 200µm, reusable |  | 2
| 7.10 | Laser fibre 272µm, reusable |  | 2
| 7.11 | Laser googles |  | 2
| 7.12 | Fiber stripper |  | 1
| 7.13 | Ceramic scissors |  | 1
| 7.14 | Footswitch |  | 1
8 |  | Pneumatic lithotriptor |  | 1 |  | 0
| 8.1 | Frequency range 1-12 Hz
| 8.2 | Selection of impulse frequency in 12 steps
| 8.3 | Foot pedal
| 8.4 | Pneumatic handpiece
| 8.5 | Adjustment interface
| 8.6 | Probe cap 0.8-2mm
| 8.7 | Compressed air tubing for handpiece
| 8.8 | Silicone probe guides, 12pcs
| 8.9 | Silicone seals, 20pcs
| 8.10 | Pneumatic probe Ø 2.0 x 425 mm, Pneumatic probe Ø 0.8 x 605 mm, Pneumatic probe Ø 1.0 x 605 mm
| 8.11 | LithoVac body with adjustment interface
| 8.12 | Suction kit with silicone tube and connector
| 8.13 | Suction tube Ø 1.6 x 595 mm, Suction tube Ø 0.8 x 668 mm
| 8.14 | Compressed Air Tubing set 3m
| 8.15 | Compressor for compressed air (medical quality) 8 bar / 4 l, 220V
| 8.16 | Operating telescope 20° with working channel 3.5mm
| 8.17 | Sheath 24Fr.
| 8.18 | Graspin forceps 3.5mm, double action
| 8.19 | Insert graspin forceps 3.5mm double action
| 8.20 | Fiber optic light cable, diameter min. 2.5 mm, length min.200 cm
| 8.21 | Disinfection basin
9 |  | Uretherorenoscope set |  | 1 |  | 0
| 9.1 | Ureterorenoscope 12° 8/9.8FR WL 430mm
| 9.2 | Semi-rigid, with lateral ocular lens, working channel 1x5 (2x3)Fr, atraumatic distal tip, TL 600mm, fiber image carrier
| 9.3 | Graspin forceps 5Fr., WL 550mm, double action, reusable
| 9.4 | Insert graspin forceps 5FR, WL 550MM, double action, reusable
| 9.5 | Disinfection basin
| 9.6 | Fiber optic light cable, diameter min. 2.5 mm, length min.200 cm
10 |  | Electrorenoscope with histeroscopy instrument set |  | 1 |  | 0
| 10.1 | Diagnostic hysteroscope
| 10.2 | Outer sheath diametar 5 mm
| 10.3 | Working channel 5Fr.
| 10.4 | Autoclavable 134°
| 10.5 | Grasping forceps 5Fr., working length 340mm
| 10.6 | Scissors insert with sheath 5 Fr., working length 340mm
| 10.7 | Grasping insert with sheath 5Fr., working length 340mm
| 10.8 | Sterilization basket
| 10.9 | Bipolar hystero-resectoscope
| 10.10 | HD Telescope 30°, diam.4mm, WL ≈300mm
| 10.11 | Outer sheath 24 Fr., rotatable, distal end straight, continuous irrigation
| 10.12 | Internal sheath 22 Fr., rotatable, slanted distal end, continuous irrigation, with obturator
| 10.13 | Working element passive bipolar
| 10.14 | Cutting electrode bipolar 24Fr, reuseble |  | 3
| 10.15 | Coagulation electrode bipolar, reuseble |  | 2
| 10.16 | Hook electrode bipolar, reuseble |  | 2
| 10.17 | Bipolar connecting cable, 3 m
| 10.18 | Protective sterilisation sleeve for resection or vaporisation electrodes
| 10.19 | Sterilization basket
| 10.20 | Fiber optic light cable, diameter min. 2.5 mm, length min.200 cm
11 |  | Flexible ureterorenoscope |  | 1 |  | 0
| 11.1 | Video ureterorenoscope
| 11.2 | Outer sheath diametar 8.7Fr.
| 11.3 | Working channel 3.6Fr.
| 11.4 | Deflection +270°/-270°, (up/down)
| 11.5 | Working length 680mm
| 11.6 | Two functions buttons
| 11.7 | Distal Sensor, Fixfocus
| 11.8 | LED illumination
| 11.9 | Adapter with instrument insertion tap
12 |  | Basic rectroscopy instrument set |  | 2 |  | 0
| 12.1 | Rectoscope Tube, outer diameter 20 mm, working length 30 cm, consisting of: Rectoscope Tube, Obturator
| 12.2 | Rectoscope Tube, outer diameter 16 mm, working length 30 cm, consisting of: Rectoscope Tube, Obturator
| 12.3 | Illumination Head, with fiber optic light transmission and LUER-Lock hub for rubber insufflation bulb, consisting of: Illumination Head, Glass Window Plug, for rectoscope tubes
| 12.4 | Rubber insufflation bulb, with detachable tube, with LUER-lock connector, consisting of: Rubber bulb, Rubber tube
| 12.5 | Sponge Holder, working length 40 cm
| 12.6 | Suction Tube, with thumb control, working length 40 cm
| 12.7 | Biopsy Forceps, medium jaws, single action jaws, rotating, insulated, with connector pin for unipolar coagulation, standard model for mucosa biopsy, size 5 mm, working length 36 cm, consisting of: Plastic Handle, without ratchet, Insulated Metal outer Tube, Working Insert
| 12.8 | Fiber Optic Light Cable, with straight connector, diameter 3.5 mm, length 200 cm
| 12.9 | LED Cold Light Fountains, Illumination technology LED, - Color temperature approx. 6400 K, - Average lamp service life diameter 30,000 h, - Light intensity adjustment manual, - Degree of safety: CF-Defib
| 12.10 | Plastic Container, for sterilization and storage, perforated, with transparent lid, with silicone mat, external dimensions aprox. 500 x 250 x 60 mm
13 |  | Endourology visualisation system with histeroscopy pump |  | 1 |  | 0
| 13.1 | Full HD camera control unit Touch screen display Resolution: min.1920 x 1080 pixel, USB port on camera controller for directly saving images and videos to USB storage media
| 13.2 | Full HD 3CCD Camera Head
| 13.3 | Fully autoclavable 134°C
| 13.4 | C-Mount Zoom Coupler, focal length ≈ f = 13-29mm
| 13.5 | Led Light Source
| 13.6 | Working life of lamp approx. 30.000h
| 13.7 | Brightness control 0-100%
| 13.8 | Fiber optic light cable, diameter min. 5 mm, length min.300cm |  | 2
| 13.9 | Full HD Medical Monitor min.27" |  | 2
| 13.10 | Diagonal:min 27 inch
| 13.11 | Resolution-min:1920 x 1080Pixel
| 13.12 | Suction and irrigation pump Laparoscopy:
| 13.13 | Irrigation ≈ 1.8l/min. (± 10%) Suction ≈ 2 l/min. (± 10%), Hysteroscopy:
| 13.14 | Hysteroscopy: Irrigation ≈ 800ml/min. Irrigation pressure up to 200mmHg Suction ≈ 2 l/min. (± 10%)
| 13.15 | Endourology: Irrigation 10 - 800 ml/min Output nominal pressure range 15 - 90 mmHg
| 13.16 | Tubing set, reusable |  | 2
| 13.17 | Universal video trolley with integrated isolating transformer
| 13.18 | Insufflator CO2
| 13.19 | Touch screen display
| 13.20 | With smoke evucation, integrated CO2 gas heating function (37°C)
| 13.21 | Gas flow rate: primary filling 1 l/min., high flow 2 l/min. to 45 l/min. (± 20 %) preselectable in 1 l/min steps
| 13.22 | HD Telescope 30°, diam.10mm, WL ≈300mm |  | 2
14 |  | Rigid endoscopes 2,7 mm, 180mm lenght for ENT department |  | 2 |  | 0
| 14.1 | Telescope 0°, enlarged view, diameter 2. 7 mm, length 18 cm, autoclavable, fiber optic light transmission incorporated, with plastic containter |  | 1
| 14.2 | Telescope 30°, enlarged view, diameter 2.7 mm, length 18 cm, autoclavable, fiber optic light transmission incorporated, color code: red |  | 1