## PriceSchedule
Bidder: |  |  |  | Offered price: | 0
LOT 2_OP Room and Surgical Instruments
Item No. | Equipment name | Qty | Model, Manufacturer and Country of Origine | Unit price | Total Price per line item
1 | Electrosurgical unit, dual output power 300W | 4 |  |  | 0
2 | Electrosurgical unit for open and laparascopic surgery | 2 |  |  | 0
3 | Ultrasonic generator for bipolar and unltrasonic sealing | 2 |  |  | 0
4 | Bipolar vessel sealing device | 1 |  |  | 0
5 | Device for surgical smoke and aerosol evacuation with accessorizes | 6 |  |  | 0
6 | Radiofrequency surgical generator | 1 |  |  | 0
7 | Cryomatic System | 1 |  |  | 0
8 | System for autotransfusion | 1 |  |  | 0
9 | OR suction | 6 |  |  | 0
10 | BASIC SET HIP | 2 |  |  | 0
11 | SEPARATE INSTRUMENT HIP | 1 |  |  | 0
12 | BASIC SET LOWER EXTREMITIES | 3 |  |  | 0
13 | SEPARATE INSTRUMENT LOWER EXTREM | 1 |  |  | 0
14 | BASIC SET KNEE | 2 |  |  | 0
15 | SEPARATE INSTRUMENT KNEE | 1 |  |  | 0
16 | BASIC SET UPPER EXTREMITIES | 3 |  |  | 0
17 | SEPARATE INSTRUMENT UPPER EXTR | 1 |  |  | 0
18 | OSTEOTOMES SET | 2 |  |  | 0
19 | STERILE CONTAINERS FOR TEXTILE | 20 |  |  | 0
20 | ADENOTOMY SET | 1 |  |  | 0
21 | TONSILECTOMY SET | 4 |  |  | 0
22 | TRACHEOTOMY SET | 2 |  |  | 0
23 | LAPAROTOMY SET | 5 |  |  | 0
24 | AMPUTATION SET | 2 |  |  | 0
25 | BASIC PLASTIC SET | 2 |  |  | 0
26 | DRESSING SET SMALL | 10 |  |  | 0
27 | MAMMECTOMY SET | 1 |  |  | 0
28 | MINOR BASIC SET, GENERAL | 5 |  |  | 0
29 | UROLOGY BLADDER SET | 1 |  |  | 0
30 | UROLOGY KIDNEY SET | 1 |  |  | 0
31 | VARICOSE VEIN STRIPPING SET | 1 |  |  | 0
32 | VASCULAR BASIC SET | 1 |  |  | 0
33 | THORACOTOMY SET (LUNG) | 1 |  |  | 0
34 | CHANGE OF BANDAGE SET | 5 |  |  | 0
35 | CURETTAGE SET | 2 |  |  | 0
36 | GYNECOLOGY SET(ABDOMINAL HYSTERECTOMY) | 1 |  |  | 0
37 | CAESARIAN SECTION SET | 2 |  |  | 0
38 | SUPPLEMENT SET FOR VAGINAL HYSTERECTOMY | 1 |  |  | 0
39 | Instruments set for Laparoscopy | 2 |  |  | 0
40 | Instruments set for Laparoscopy-colorectal Surgery | 1 |  |  | 0
41 | Instruments set for Laparoscopy in Gynecology | 1 |  |  | 0
42 | FESS instrument set | 1 |  |  | 0
43 | Instrument set for Tonsillectomy | 1 |  |  | 0
44 | Diagnostic Instrument set for ENT | 2 |  |  | 0
45 | Basic set for Ear Micro Surgery | 1 |  |  | 0
46 | Instrument sets Laryngomicroscopy (LMS) | 1 |  |  | 0
47 | Oftalmology instruments for cataract injurries, shutters and vitreoretinal | 1 |  |  | 0
48 | OPERATING TABLE FOR GENERAL SURGERY | 1 |  |  | 0
49 | OPERATING TABLE FOR COLORECTAL SURGERY | 1 |  |  | 0
50 | OPERATING TABLE FOR ENT SURGERY | 1 |  |  | 0
51 | OPERATING TABLE FOR OPHTAMOLOGY | 1 |  |  | 0
52 | OPERATING TABLE FOR UROLOGY | 1 |  |  | 0
53 | OPERATING TABLE FOR GYNECOLOGY | 1 |  |  | 0
54 | OPERATING TABLE FOR ORTHOPEDY | 1 |  |  | 0
55 | CEILING LED OPERATING LIGHT WITH SATELITE | 6 |  |  | 0
56 | MOBILE LED EXAMINATION LIGHT | 7 |  |  | 0
Total price per Lot |  |  |  |  | 0

## TechSpecification
Bidder: | [5 empty] | Total DAP price: | 0
LOT 2_OP Room and Surgical Instruments
Item No. | ID | Technical Specification Requested | Technical Specification Offered | Qty | Unit price | Total Price per line item | Insert page no. in techical documentation
1 |  | Electrosurgical unit, dual output power 300W |  | 4 |  | 0
| 1.1 | At least two independent monopolar accessory receptacles with corresponding displays, hand controlled and minimum one combination receptacle
| 1.2 | At least two monopolar cutting modes including Pure 300W in maximum and Blend 200W in maximum
| 1.3 | At least two monopolar coagulation modes including Spray 80W in maximum and Standard 120W in maximum
| 1.4 | Bipolar accessory receptacle with corresponding display and minimum two bipolar modes including Macro and Micro
| 1.5 | Output power settings with power cut adjustments keys on control panel
| 1.6 | Dispersive electrode status/ alarm indicator on control panel
| 1.7 | ARM sensor for dispersive electrode contact quality monitoring
| 1.8 | Monopolar and bipolar footswitch
| 1.9 | Accessories should include monopolar and bipolar cable, dispersive electrode with cable, bipolar forceps
2 |  | Electrosurgical unit for open and laparascopic surgery |  | 2 |  | 0
| 2.1 | Microprocessor controlled electrosurgical unit
| 2.2 | Minimum 20 million measurements per second as a function of real-time current control in relation to the applied power to the tissue
| 2.3 | Output power minimum 400 W when cutting
| 2.4 | Output power minimum 200 W during coagulation
| 2.5 | Wi-Fi connection for service mode or operating room integration system
| 2.6 | Equipotency connection
| 2.7 | User color TFT or LCD touch screen with a diagonal of at least 25 cm and to have a precise and clear display of each activation and program.
| 2.8 | Possibility to set a minimum of 300 programs
| 2.9 | Number of monopolar cutting modes minimum 5 modes
| 2.10 | Number of monopolar coagulation modes minimum 5 modes
| 2.11 | A mode that enables the simultaneous operation of two positive monopolar electrodes on two operating fields of one patient
| 2.12 | Number of bipolar cutting modes minimum 2 modes
| 2.13 | Number of bipolar coagulation modes minimum 2 mod
| 2.14 | Number of thermofusion modes 1 mode and more
| 2.15 | Number of terminals for monopolar and bipolar active instruments minimum 4 terminals plus 1 terminal for neutral electrode
| 2.16 | The connectors on the device are easily interchangeable and have the possibility of combining the settings of the connectors according to the choice of the end user
| 2.17 | Possibility to select the connector to which it can be used for both mopolar or bipolar instrument
| 2.18 | That the HF generator can be upgraded with: Argon generator, Water knife, smoke evacuator or irrigation pump
| 2.19 | Composition of the electrosurgical unit and nesessary accessories:
| 2.20 | HF generator with appropriate monitor
| 2.21 | Mobile trolley
| 2.22 | Foot pedal with 3 button
| 2.23 | Neutral self-adhesive electrode with equipotential additional protective ring, 50/1
| 2.24 | Neutral electrode cable
| 2.25 | Positive electrode handle with keys and cable
| 2.26 | Set of monopolar electrodes
| 2.27 | Bipolar forceps 190 mm
| 2.28 | Cable for Bipolar forceps
| 2.29 | Bipolar instrument for ligation, for open surgery, 210 mm
| 2.30 | Bipolar instrument for ligation and cutting in one act, 350 mm, single use
3 |  | Ultrasonic generator for bipolar and unltrasonic sealing |  | 2 |  | 0
| 3.1 | Generator for ultrasonic and bipolar electrosurgical power. Indicated for soft tissue incisons during laparoscopic and open procedures
| 3.2 | The generator should be delivered with a power supply cable, adapter for ultrasonic instruments and a two-pedal footswitch. Minimum two additional different cables, with piezoelectric ceramics that convert electrical energy into mechanical motion, for minimum use of 95 times
| 3.3 | Bipolar instruments output: 100 VAC RMS maximum, 135 watts maximum, 300 - 490 kHz. No neutral electrode is required or equivalent Ultrasonic instruments output: 150 VAC RMS maximum, 35 watts continuous, 30 - 80 kHz or equivalent
| 3.4 | Menu should be available on touch screen panel. Ultrasonic power levels adjustmant on minimum 5 levels
| 3.5 | During activation, ultrasonic blade is being displaced in minimal range from 33 to 65 microns
| 3.6 | Generator should detect changes between instrument jaws and provides the user with a) audible feedback, b) reduction in power level and c) improved thermal management or equvalent
| 3.7 | There should be at least five ultrasonic shears for open surgery, shaft lengths between 9 and 23 cm. There are at least two ultrasonic hook/spatula instruments for open surgery, adjustable shaft lengths: 4 – 9 cm. |  | 2
| 3.8 | There are should be at least five ultrasonic instruments for laparoscopic procedures, shaft lengths between 32 and 45 cm. |  | 2
| 3.9 | There are should be at least four advanced bipolar instruments for open and laparoscopic procedures, shaft lengths between 20 and 45 cm. |  | 2
4 |  | Bipolar vessel sealing device |  | 1 |  | 0
| 4.1 | Lektrafuse HF generator, Plug & Play, 2 sealing modes, for use in open and laparoscopic surgery. Operating frequency 460kHz, output power 2 x 150W on 50 Ohm |  | 1
| 4.2 | Bipolar instrument 5mm 36cm, Maryland jaw, 360° shaft rotation, non articulating jaw, sealing length 21mm, cutting lenght 20mm, jaw closes at the tip first, for sealing vessels up to and including 7 mm |  | 2
5 |  | Device for surgical smoke and aerosol evacuation with accessorizes |  | 6 |  | 0
| 5.1 | Surgical smoke evacuator with minimum three operating modes with flow setting of minimum 800 LPM and noise level less than 55db
| 5.2 | ULPA filter with four stages of filtration and tree ports for different clinical applications, filter life tracking option
| 5.3 | Safety occlusion alarm
| 5.4 | Possibility of suction control in minimum 10 levels using touchscreen interface
| 5.5 | Adjustable delay option up to 10 seconds for capturing of residual smoke
| 5.6 | Accessories should include automatic activation cable, replacement filter, translucent fluid trap and electrosurgical evacuation pencils for different clinical applications
6 |  | Radiofrequency surgical generator |  | 1 |  | 0
| 6.1 | Radiofrequency surgical generator with output frequency 4.0MHz Monopolar and 1.7MHz Bipolar. Output Power Bipolar 90Watts.
| 6.2 | Foot Controlled Handpiece
| 6.3 | 3-Button Fingerswitch Handpiece
| 6.4 | Disposable 3-Button Fingerswitch Handpiece
| 6.5 | Dual Footswitch & Cable
| 6.6 | Neutral plate
| 6.7 | Bipolar Cable
| 6.8 | Smoke Evacuation System
| 6.9 | Mobile cart
| 6.10 | Power cord
| 6.11 | Instruction Manual
7 |  | Cryomatic System |  | 1 |  | 0
| 7.1 | Cryo Console for use in ophthalmic surgery Footswitch, high pressure & exhaust hose
| 7.2 | Pin Index Yoke for E Size Carbon Dioxide Gas Cylinder
| 7.3 | Adaptor for F Size Carbon Dioxide Gas Cylinder
| 7.4 | Probe Sterilisation Boxes
| 7.5 | 2.5mm Standard Retinal Pencil reusable
| 7.6 | Mobile trolley
| 7.7 | Power cord
| 7.8 | Instruction Manual
8 |  | System for autotransfusion |  | 1 |  | 0
| 8.1 | Full-color touch screen LCD, TFT ( at least 8")
| 8.2 | Internal memory ,USB ports for simultaneous download of multiple cases in MS® Excel format
| 8.3 | Integrated ink printer
| 8.4 | RS232 ports for external PC connection
| 8.5 | Dual RBC detector technology
| 8.6 | Integrated Hct sensor (inlet and outlet Hct)
| 8.7 | The four wheels coupled with the front and rear handles
| 8.8 | Waste line color indicator
| 8.9 | Automatic start at operator determined target volume in the reservoir
| 8.10 | Continuous processing of bowls in sequence
| 8.11 | Automatic stop when reservoir is empty
| 8.12 | Operating protocols: POPT, PSTD, POST-OP
| 8.13 | preoperative sequestration protocols (PPP,PRP1, PRP2)
| 8.14 | HCT 60-65%
| 8.15 | Wash time approx. 5min
| 8.16 | Removal rate >95%
| 8.17 | modular vacuum pump with possibility to be used as standalone
| 8.18 | 50 to -300 mmHg (Intra-operative mode -High vacuum capacity)
| 8.19 | -10 to -100 mmHg (Post-operative mode Gentle aspiration for wound drainage connection)
| 8.20 | Availability of disposables in sizes: 55ml, 125ml, 175ml, 225ml
| 8.21 | Availability of Blood Collection Reservoir with 4 liter capacity and internal defoaming polyurethane layer (30 ppi), Intermediate woven, non-woven filter (40 μm
9 |  | OR suction |  | 6 |  | 0
| 9.1 | Aspirator should be portable
| 9.2 | maximum suction flow 5 lit/min, maximum vacuum -90 kPa
| 9.3 | Maximum noise: 55 dB
| 9.4 | Aspirator can be mounted on rail
| 9.5 | Aspirator has 2 secretion jars. Jars are integrated on holder, so they are also portable together with aspirator
| 9.6 | Capacity of secretion jars 3l,. Jars are autoclavable on temperatures up to 134 °C
| 9.7 | Aspirator is driven by compressed Air or vacuum from central gas supply
| 9.8 | With bronchial aspirator should be dellivered: - gas supply hose
10 |  | BASIC SET HIP |  | 2 |  | 0
| 10.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 10.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 10.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 6
| 10.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 1
| 10.5 | Scalpel handle No.4L for blades fig. 18-36 215 mm, 8 1/2" total length |  | 1
| 10.6 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 10.7 | Titanit "Metzenbaum-Fino" diss.scissors curved 230 mm 9 1/8" total length |  | 1
| 10.8 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 10.9 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 10.10 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 10.11 | Surgical forceps, standard str., tip 2,5 mm, 1x2 teeth, flat handle 160 mm, 6 1/4" total length |  | 2
| 10.12 | Surgical forceps, standard str., tip 3,0 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 10.13 | Surgical forceps str., tip 1,8 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 1
| 10.14 | Atr. Pinzette n. DeBakey, ger. Safety-Pin, tip 2,8mm 200 mm, 8" total length |  | 2
| 10.15 | "DeBakey" atr. forceps, str. Safety-Pin with Press-Stopper, tip 2,8mm 240 mm, 9 1/2" total length |  | 1
| 10.16 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 4
| 10.17 | "Rochester-Pean" haemostatic forceps straight, tip 3,0 mm, serrated jaw 160 mm, 6 1/4" total length |  | 2
| 10.18 | "Kocher-Ochsner" haemostatic forceps str., tip 2,7mm,1x2 teeth, serrated jaw 160 mm, 6 1/4" total length |  | 4
| 10.19 | "Halsted" haemostatic forceps curved, tip 1,7 mm, serrated jaw 185 mm, 7 1/4" total length |  | 2
| 10.20 | "Kocher-Ochsner" haemostatic forceps str., tip 3,7mm,1x2 teeth, serrated jaw 225 mm, 9" total length |  | 2
| 10.21 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 2
| 10.22 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,5 mm, serrated jaw 230 mm, 9 1/8" total length |  | 1
| 10.23 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 2
| 10.24 | Power TC needle holder a. Hegar strong model 205 mm 8 1/8" total length |  | 1
| 10.25 | Roux hook woundretractor double , big 165mm, 6 1/2" total length |  | 2
| 10.26 | "Kocher" retractor 60x20 mm 230 mm, 9 1/8" total length |  | 2
| 10.27 | "Mikulicz" abdominal hook 85x37 mm 250 mm, 10" total length |  | 2
| 10.28 | "Kocher" retractor 6-teeth, sharp, blade 15x30 mm 220 mm, 8 3/4" total length |  | 2
| 10.29 | "Körte" retractor 8-teeth, sharp, blade 26x40 mm 240 mm, 9 1/2" total length |  | 2
| 10.30 | "Israel" retractor 6-teeth, blunt, blade 53x62 mm 255 mm, 10 1/8" total length |  | 2
| 10.31 | "Volkmann" retractor sharp, blade 20x0,7 mm 220 mm, 8 3/4" total length |  | 1
| 10.32 | Bone levers, 17mm width 240 mm, 9 1/2" total length |  | 2
| 10.33 | "Lange-Hohmann" bone lever 22,5 mm wide 250 mm, 10" total length |  | 2
| 10.34 | "Verbrugge-Müller" bone lever 44,0 mm wide 240 mm, 9 1/2" total length |  | 2
| 10.35 | "Hohmann" bone lever, 70 mm blade 250 mm, 10" total length |  | 1
| 10.36 | "Schumacher" bone lever 18 / 35 mm wide 290 mm, 11 1/2" total length |  | 1
| 10.37 | "Simon" uterine scoop fig. 2, sharp, mouth 8,5x12,0 mm 240 mm, 9 1/2" total length |  | 1
| 10.38 | "Simon" uterine scoop fig. 4, sharp, mouth 11,5x15,5 mm 240 mm, 9 1/2" total length |  | 1
| 10.39 | "Endo" bone curette curved, mouth 31x40 mm 240 mm, 9 1/2" total length |  | 1
| 10.40 | "Langenbeck" bone holding forceps straight 205 mm, 8 1/8" total length |  | 1
| 10.41 | "Röttgen-Ruskin" bone rongeur curved 240 mm, 9 1/2" total length |  | 1
| 10.42 | Langenbeck periosteal elevator narrow 10 mm, 195mm 7 3/4" total length |  | 1
| 10.43 | "Langenbeck" raspatory 16,0 mm wide, sharp 190 mm, 7 1/2" total length |  | 1
| 10.44 | "Semb" raspatory 15,0 mm wide 230 mm, 9 1/8" total length |  | 1
| 10.45 | "Wagner" raspatory 17,0 mm wide, PPSU-handle 330 mm, 13 1/8" total length |  | 1
| 10.46 | "Perthes" reamer square, sharp pointed tip 215 mm, 8 1/2" total length |  | 1
| 10.47 | "Lexer" osteotome, with PPSU-handle, 20 mm wide, 220 mm, 8 3/4" total length |  | 1
| 10.48 | "Lexer" osteotome, with PPSU-handle, 30 mm wide, 220 mm, 8 3/4" total length |  | 1
| 10.49 | "Lambotte" osteotome straight, mouth 30 mm 250 mm, 10" total length |  | 1
| 10.50 | "Lexer" osteotome angled, 10 mm wide, with PPSU-handle 275 mm, 11" total length |  | 1
| 10.51 | "Lexer" osteotome angled, 15 mm wide, with PPSU-handle 275 mm, 11" total length |  | 1
| 10.52 | Spoon chisel angled, 16 mm wide, with PPSU-handle 270 mm, 10 3/4" total length |  | 1
| 10.53 | Swan neck chisel double curved, with PPSU-handle 285 mm, 11 1/4" total length |  | 1
| 10.54 | "Caspar" tamper Ø 8,0 mm, with PPSU-handle 200 mm, 8" total length |  | 1
| 10.55 | Tamper Ø 12 mm, PPSU-handle 200 mm, 8" total length |  | 1
| 10.56 | "Wagner" mallet, Ø 35 mm with 1 pair plastic discs 235 mm, 9 1/4" total length |  | 1
| 10.57 | Removal femur instrument with T-handle PPSU 270 mm, 10 3/4" long |  | 1
| 10.58 | Spoon for Luxation femur width 25 mm 315 mm, 12 1/2" total length |  | 1
| 10.59 | Resano capsulas scissors angled, 250 mm 10" total length |  | 1
| 10.60 | Bircher-Ganske meniscus forceps upcurved, 200 mm, 8" total length |  | 1
| 10.61 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 10.62 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 10.63 | Cont. 1/1 bottom non-perforated 600x272x190mm |  | 1
| 10.64 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 10.65 | Wire basket 1/1, fall handle, 480x250x100mm |  | 1
11 |  | SEPARATE INSTRUMENT HIP |  | 1 |  | 0
| 11.1 | "Lane" bone holding forceps mouth width 13 mm, witch ratchet 320 mm, 12 3/4" total length |  | 1
| 11.2 | "Cobb" raspatory 19,0 mm wide 285 mm, 11 1/4" total length |  | 1
| 11.3 | Swan neck chisel double curved, with PPSU-handle 285 mm, 11 1/4" total length |  | 1
| 11.4 | Bircher-Ganske meniscus forceps to the side angled 200 mm, 8" total length |  | 1
| 11.5 | Resano capsulas scissors angled, 250 mm 10" total length |  | 1
12 |  | BASIC SET LOWER EXTREMITIES |  | 3 |  | 0
| 12.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 12.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 10
| 12.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 6
| 12.4 | Scalpel handle No.4L for blades fig. 18-36 215 mm, 8 1/2" total length |  | 1
| 12.5 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 2
| 12.6 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 1
| 12.7 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 12.8 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 12.9 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 12.10 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 12.11 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 12.12 | Surgical forceps, standard str., tip 2,5 mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 12.13 | Surgical forceps, standard str., tip 3,0 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 12.14 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 12.15 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 12.16 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 6
| 12.17 | "Rochester-Pean" haemostatic forceps straight, tip 3,0 mm, serrated jaw 160 mm, 6 1/4" total length |  | 6
| 12.18 | "Kocher-Ochsner" haemostatic forceps str., tip 2,7mm,1x2 teeth, serrated jaw 160 mm, 6 1/4" total length |  | 6
| 12.19 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 2
| 12.20 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,3 mm, serrated jaw 220 mm, 8 3/4" total length |  | 1
| 12.21 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,5 mm, serrated jaw 230 mm, 9 1/8" total length |  | 1
| 12.22 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 210 mm 8 1/4" total length |  | 1
| 12.23 | Power TC needle holder a. Hegar strong model 205 mm 8 1/8" total length |  | 1
| 12.24 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 1
| 12.25 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 12.26 | "Kocher" retractor 60x25 mm 230 mm, 9 1/8" total length |  | 2
| 12.27 | "Langenbeck" retractor blunt, blade 40x13 mm 210 mm, 8 1/4" total length |  | 2
| 12.28 | "Volkmann" retractor 4 teeth, 8,5x19 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 12.29 | "Kocher" retractor 6-teeth, sharp, blade 15x30 mm 220 mm, 8 3/4" total length |  | 2
| 12.30 | "Körte" retractor 8-teeth, sharp, blade 26x40 mm 240 mm, 9 1/2" total length |  | 2
| 12.31 | "Volkmann" retractor sharp, blade 20x0,7 mm 220 mm, 8 3/4" total length |  | 1
| 12.32 | "Israel" retractor 6-teeth, blunt, blade 53x62 mm 255 mm, 10 1/8" total length |  | 1
| 12.33 | "Wagner" bone lever 8,0 mm wide 220 mm, 8 3/4" total length |  | 2
| 12.34 | Bone levers, 17mm width 240 mm, 9 1/2" total length |  | 2
| 12.35 | "Verbrugge-Müller" bone lever 44,0 mm wide 240 mm, 9 1/2" total length |  | 2
| 12.36 | Hohmann bone lever 13/35 mm width 285 mm, 11 1/4" total length |  | 2
| 12.37 | "Simon" uterine scoop fig. 2, sharp, mouth 8,5x12,0 mm 240 mm, 9 1/2" total length |  | 1
| 12.38 | "Volkmann" bone curette fig.1, mouth 6,8x9,5 mm 170 mm, 6 3/4" total length |  | 1
| 12.39 | "Volkmann" bone curette fig.4, mouth 11,5x15,5 mm 170 mm, 6 3/4" total length |  | 1
| 12.40 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 12.41 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 12.42 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 12.43 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 12.44 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 12.45 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
13 |  | SEPARATE INSTRUMENT LOWER EXTREM |  | 1 |  | 0
| 13.1 | "Overholt-Geissendörfer" dissecting forc curved, tip 1,8 mm, serrated jaw 195 mm, 7 3/4" total length |  | 1
| 13.2 | Mixter discetting forceps, curved 220 mm, 8 3/4" total length |  | 1
| 13.3 | "WEITLANER" WOUNDRETRACTOR, SHARP 3X4 TEETH, 200 MM, 8" TOTAL LENGTH |  | 1
| 13.4 | Weitlaner retractor sharp 5X6 teeth 255 mm, 10 1/8" total length |  | 1
| 13.5 | "Crile" haemostatic forceps straight, tip 1,6 mm, serrated jaw 140 mm, 5 1/2" total length |  | 1
| 13.6 | Power TC wire forceps, tip conic 170 mm, 6 3/4" total length |  | 2
| 13.7 | "Deschamp" ligature needle curved right 230 mm, 9 1/8" total length |  | 1
| 13.8 | "Deschamp" ligature needle curved left 280 mm, 11 1/8" total length |  | 1
| 13.9 | "Deschamp" ligature needle curved right 280 mm, 11 1/8" total length |  | 1
14 |  | BASIC SET KNEE |  | 2 |  | 0
| 14.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 14.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 4
| 14.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 4
| 14.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 2
| 14.5 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 2
| 14.6 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 14.7 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 14.8 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 14.9 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 14.10 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 14.11 | Surgical forceps, standard str., tip 2,5 mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 14.12 | Surgical forceps, standard str., tip 3,0 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 14.13 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 14.14 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 6
| 14.15 | "Kocher-Ochsner" haemostatic forceps str., tip 2,7mm,1x2 teeth, serrated jaw 160 mm, 6 1/4" total length |  | 6
| 14.16 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,3 mm, serrated jaw 220 mm, 8 3/4" total length |  | 1
| 14.17 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 14.18 | "Hegar-Mayo" power TC needle holder straight, tip 2,5 mm, needle size 1-4 mm 180 mm, 7 1/8" total length |  | 1
| 14.19 | Power TC needle holder a. Hegar strong model 205 mm 8 1/8" total length |  | 1
| 14.20 | "Roux" retractor, double ended 19x22 / 23x30 mm 145 mm, 5 3/4" total length |  | 2
| 14.21 | "Langenbeck" retractor blunt, blade 40x13 mm 210 mm, 8 1/4" total length |  | 2
| 14.22 | "Langenbeck" retractor blunt, blade 55x20 mm 220 mm, 8 3/4" total length |  | 2
| 14.23 | "Volkmann" retractor sharp, blade 20x0,7 mm 220 mm, 8 3/4" total length |  | 1
| 14.24 | "Volkmann" retractor 4 teeth, 8,5x19 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 14.25 | "Weitlaner" self-retaining retractor semi-sharp, 3/4 teeth 165 mm, 6 1/2" total length |  | 1
| 14.26 | Dissector"Langenbeck" 8mm width, 200mm, 8" total length |  | 1
| 14.27 | "Hohmann" bone lever 17,0 mm wide 240 mm, 9 1/2" total length |  | 2
| 14.28 | Knochenhebel n.Lange-Hohmann, modif., gr 33 mm wide, 280 mm, 11 1/8" total length |  | 2
| 14.29 | "Volkmann" bone curette fig.1, mouth 6,8x9,5 mm 170 mm, 6 3/4" total length |  | 1
| 14.30 | "Weller" meniscus holding forceps curved to side 200 mm, 8" total length |  | 1
| 14.31 | "Langenbeck" bone holding forceps straight 205 mm, 8 1/8" total length |  | 1
| 14.32 | "Ruskin" bone rongeur curved, 5x15 mm jaw 180 mm, 7 1/8" total length |  | 1
| 14.33 | "Röttgen-Ruskin" bone rongeur curved 240 mm, 9 1/2" total length |  | 1
| 14.34 | "Ruskin-Liston" bone cutter 25° angeld, length of cut 25,0 mm 190 mm, 7 1/2" total length |  | 1
| 14.35 | Raspatory 13,0 mm wide, PPSU handle 190 mm, 7 1/2" total length |  | 1
| 14.36 | "Lexer" osteotome, with PPSU-handle, 10 mm wide, 220 mm, 8 3/4" total length |  | 1
| 14.37 | "Lexer" osteotome, with PPSU-handle, 15 mm wide, 220 mm, 8 3/4" total length |  | 1
| 14.38 | "Lexer" Osteotome straight, mouth 25 mm, with PPSU-handle 220 mm, 8 3/4" total length |  | 1
| 14.39 | "Wagner" mallet, Ø 35 mm with 1 pair plastic discs 235 mm, 9 1/4" total length |  | 1
| 14.40 | "Caspar" tamper Ø 5,0 mm, with PPSU-handle 200 mm, 8" total length |  | 1
| 14.41 | "Caspar" tamper Ø 8,0 mm, with PPSU-handle 200 mm, 8" total length |  | 1
| 14.42 | Power TC extraction plier for wires mouth 7,0 mm 180 mm, 7 1/8" total length |  | 1
| 14.43 | Double ended probe tip Ø 2,0 mm 180 mm, 7 1/8" total length |  | 1
| 14.44 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 14.45 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 14.46 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 14.47 | Cont. 1/1 bottom non-perforated 600x272x190mm |  | 1
| 14.48 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 14.49 | Wire basket 1/1, fall handle, 480x250x100mm |  | 1
15 |  | SEPARATE INSTRUMENT KNEE |  | 1 |  | 0
| 15.1 | Smillie meniscotome, left 170 mm, 6 3/4" total length |  | 1
| 15.2 | Smillie meniscotome, right 170 mm, 6 3/4" total length |  | 1
| 15.3 | "Officer" surgical forceps str., tip 4,7 mm, 4x5 teeth, flat handle 150 mm, 6" total length |  | 1
| 15.4 | Surgical forceps str., tip 4,5 mm, 3x4 teeth, flat handle 200 mm, 8" total length |  | 1
| 15.5 | Surgical forceps str., tip 5,0 mm, 4x5 teeth, flat handle 145 mm, 5 3/4" total length |  | 1
| 15.6 | Surgical forceps, russian model str, ring-Ø 7,0 mm, serrated, flat handl 250 mm, 10" total length |  | 2
16 |  | BASIC SET UPPER EXTREMITIES |  | 3 |  | 0
| 16.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 16.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 16.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 6
| 16.4 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 2
| 16.5 | "Titanit" Metzenbaum-Fino scissors curved 145 mm 5 3/4" total length |  | 1
| 16.6 | "Titanit" Jameson-Werbe scissors curved, fine blunt baldes, 135 mm 5 1/4" total length |  | 1
| 16.7 | "Titanit" Tuula Pankreas diss. scissors curved, 150 mm 6" total length |  | 1
| 16.8 | Power TC Iris- and ligature scissors, sh./bl., curved, 110 mm 4 1/4" total length |  | 1
| 16.9 | Power TC ligature scissors, curved, wavecut, 145 mm 5 3/4" total length |  | 1
| 16.10 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 16.11 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 16.12 | "DeBakey" atr. forceps, straight safety-pin, tip 1,5mm 155 mm, 6 1/8" total length |  | 2
| 16.13 | "Micro-Adson" surgical forceps str., tip 0,9 mm, 1x2 teeth, flat handle 120 mm, 4 3/4" total length |  | 2
| 16.14 | "Micro-Adson" anatomic forceps str., tip 1,2 mm, serrated, flat handle 120 mm, 4 3/4" total length |  | 2
| 16.15 | Anatomic splinter forceps str., tip 0,5 mm, serrated, flat handle 145 mm, 5 3/4" total length |  | 2
| 16.16 | "Micro-Halsted" haemostatic forceps curved, tip 0,8 mm, serrated jaw 125 mm, 5" total length |  | 6
| 16.17 | "Micro-Halsted" haemostatic forceps str.,tip 1,2mm,1x2 teeth, serrated jaw 125 mm, 5" total length |  | 6
| 16.18 | "Leriche" haemostatic forceps straight, tip 1,6 mm, 1x2 teeth 150 mm, 6" total length |  | 2
| 16.19 | "Leriche" haemostatic forceps curved, tip 1,3 mm, serrated jaw 150 mm, 6" total length |  | 2
| 16.20 | "Micro-Adson" haemostatic forceps curved, tip 1,0 mm, serrated jaw 140 mm, 5 1/2" total length |  | 1
| 16.21 | "Baby-Mixter" dissecting forceps curved, tip 1,4 mm, serrated jaw 140 mm, 5 1/2" total length |  | 1
| 16.22 | "Baby-Adson" dissecting forceps curved, tip 1,7 mm, serrated jaw 140 mm, 5 1/2" total length |  | 1
| 16.23 | "Halsey" power TC needle holder straight, tip 1,5 mm, needle size 5-7 mm 130 mm, 5 1/8" total length |  | 1
| 16.24 | "Hegar-Mayo" power TC needle holder straight, tip 2,2 mm, needle size 2-4 mm 150 mm, 6" total length |  | 1
| 16.25 | Wound retractor, delicate model small curve, 1 tooth, sh. 165 mm, 6 1/2" total length |  | 2
| 16.26 | Wound-/trachea retractor large curve, 1 tooth, sharp 165 mm, 6 1/2" total length |  | 2
| 16.27 | RETRACTOR, SMALL CURVE,2 PRONGS SHARP, 165 MM, 6 1/2" TOTAL LENGTH |  | 2
| 16.28 | Retractor, small curve,3 prongs sharp, 160 mm, 6 1/4" total length |  | 2
| 16.29 | Retractor, small curve,4 prongs sharp, 165 mm, 6 1/2" total length |  | 2
| 16.30 | Wound retractor, delicate model small curve, 2 teeth, 4,5x4,2mm, blunt 165 mm, 6 1/2" total length |  | 2
| 16.31 | Retractor, small curve,3 prongs blunt, 165 mm, 6 1/2" total length |  | 2
| 16.32 | Wound retractor, delicate model small curve, 4 teeth, 4,5x9,5mm, blunt 165 mm, 6 1/2" total length |  | 2
| 16.33 | Iris hook "Rollet", 2-teeth, sharp, 130 mm, 5 1/8" total length |  | 2
| 16.34 | Iris retractor "Rollet", 4-teeth, sharp 130 mm, 5 1/8" total length |  | 2
| 16.35 | Langenbeck-Green retractor, fine, 16X6mm 160 mm, 6 1/4" total length |  | 2
| 16.36 | Langenbeck-Green retractor, fine,6X25 mm 160 mm, 6 1/4" total length |  | 2
| 16.37 | "Desmarres" lid retractor 8x8 mm 140 mm, 5 1/2" total length |  | 1
| 16.38 | "Desmarres" lid retractor 12x16 mm 140 mm, 5 1/2" total length |  | 1
| 16.39 | "Alm" woundretractor blunt, 4/4 teeth 100 mm, 4" total length |  | 1
| 16.40 | "Weitlaner" self-retaining retractor semi-sharp, 2/3 teeth 110 mm, 4 1/4" total length |  | 1
| 16.41 | Probes straight 130 mm, 5 1/8" total length |  | 1
| 16.42 | "Bowman" eye probe Ø 0,7 + 0,8 mm, fig. 00/0 130 mm, 5 1/8" total length |  | 1
| 16.43 | "Bowman" eye probe Ø 0,9 + 1,1 mm, fig. 1/2 130 mm, 5 1/8" total length |  | 1
| 16.44 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 16.45 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 16.46 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 16.47 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 16.48 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 16.49 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
17 |  | SEPARATE INSTRUMENT UPPER EXTR |  | 1 |  | 0
| 17.1 | "Rose" wound retractor double-end, blunt 130 mm, 5 1/8" total length |  | 2
| 17.2 | Woundretractor double 20x15mm and 20x19mm valves 205 mm total length |  | 2
| 17.3 | "Weitlaner" self-retaining retractor semi-sharp, 3/4 teeth 165 mm, 6 1/2" total length |  | 1
| 17.4 | Wound retractor Mollison sharp 4x4prongs 155mm, 6 1/8" total length |  | 1
| 17.5 | "Cottle" mallet, 1 flat 1 convex side head-Ø 30 mm, 235 g 185 mm, 7 1/4" total length |  | 1
| 17.6 | Flat nose plier mouth 1,5 mm 190 mm, 7 1/2" total length |  | 1
| 17.7 | Power TC wire cutting pliers ang, wire: soft up to 3,0 mm, hard 2,5mm 230 mm, 9 1/8" total length |  | 1
18 |  | OSTEOTOMES SET |  | 2 |  | 0
| 18.1 | "Lexer" osteotome, with PPSU-handle, 7 mm breit, 220 mm, 8 3/4" total length |  | 1
| 18.2 | "Lexer" osteotome, with PPSU-handle, 10 mm wide, 220 mm, 8 3/4" total length |  | 1
| 18.3 | "Lexer" osteotome, with PPSU-handle, 15 mm wide, 220 mm, 8 3/4" total length |  | 1
| 18.4 | "Lexer" osteotome, with PPSU-handle, 20 mm wide, 220 mm, 8 3/4" total length |  | 1
| 18.5 | "Lexer" osteotome, with PPSU-handle, 30 mm wide, 220 mm, 8 3/4" total length |  | 1
| 18.6 | "Lexer" chisel 10 mm wide, PPSU-handle 220 mm, 8 3/4" total length |  | 1
| 18.7 | "Lexer" chisel 15 mm wide, PPSU-handle 220 mm, 8 3/4" total length |  | 1
| 18.8 | "Lexer" chisel 25 mm wide, PPSU-handle 220 mm, 8 3/4" total length |  | 1
| 18.9 | Cont. 3/4 bottom non-perforated 475x272x100 mm |  | 1
| 18.10 | Cont. 3/4 flex lid system red, lid pps biostop system 475x285 mm |  | 1
| 18.11 | Wire basket 3/4, middle size, with fall handle 405x250x50 mm |  | 1
19 |  | STERILE CONTAINERS FOR TEXTILE |  | 20 |  | 0
| 19.1 | basic Lid for container, perforated |  | 1
| 19.2 | Cont. 1/1 bottom perforated 600x272x235mm |  | 1
| 19.3 | Teflon long time filter for max 2000cycles |  | 4
20 |  | ADENOTOMY SET |  | 1 |  | 0
| 20.1 | "Davis-Boyle" frame for tongue spatula for children, w. interchangeable inserts |  | 1
| 20.2 | "Davis-Boyle" frame for tongue spatula for adults, with interchangeable inserts |  | 1
| 20.3 | "Russel-Davis" tongue spatula extra narrow, blade 18x57 mm |  | 1
| 20.4 | "Russel-Davis" tongue spatula extra narrow, blade 22x67 mm |  | 1
| 20.5 | "Russel-Davis" tongue spatula extra narrow, blade 25x75 mm |  | 1
| 20.6 | "Russel-Davis" tongue spatula blade 38x85 mm |  | 1
| 20.7 | "Russel-Davis" tongue spatula blade 40x92 mm |  | 1
| 20.8 | "Yankauer" suction tube curved, Ø 10 mm 275 mm, 11" total length |  | 1
| 20.9 | "Metzenbaum" power TC diss. scissors bl./bl., cur., tip 2,8 mm 200 mm, 8" total length |  | 1
| 20.10 | Surgical forceps, standard str., tip 3,0 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 1
| 20.11 | "Hegar-Mayo" power TC needle holder straight, tip 2,5 mm, needle size 1-4 mm 180 mm, 7 1/8" total length |  | 1
| 20.12 | "Beckmann" adenoid curette fig. 0, mouth 13,8 mm 220 mm, 8 3/4" total length |  | 1
| 20.13 | "Beckmann" adenoid curette fig. 1, mouth 15,0 mm 220 mm, 8 3/4" total length |  | 1
| 20.14 | "Beckmann" adenoid curette fig. 2, mouth 15,8 mm 220 mm, 8 3/4" total length |  | 1
| 20.15 | "Beckmann" adenoid curette fig. 3, mouth 18,8 mm 220 mm, 8 3/4" total length |  | 1
| 20.16 | "Beckmann" adenoid curette fig. 4, mouth 22,5 mm 220 mm, 8 3/4" total length |  | 1
| 20.17 | "Beckmann" adenoid curette fig. 5, tip 24,1 mm 220 mm, 8 3/4" total length |  | 1
| 20.18 | Laboratory bowl, heavy-weight model B8, capacity 1,0 L, Ø 150 mm 75 mm height |  | 1
| 20.19 | Cont. 1/2 bottom non-perforated 310x272x116 mm |  | 1
| 20.20 | Cont. 1/2 flex lid system red, lid pps biostop system 310x280 mm |  | 1
| 20.21 | Wire basket tray 1/2, half size with fall handle 255x250x50 mm |  | 1
| 20.22 | Silicone mat blue perforated 250 x 235 mm |  | 1
21 |  | TONSILECTOMY SET |  | 4 |  | 0
| 21.1 | "Brünings" tongue spatula width of endings 19,5/15,2 mm, light cur 190 mm, 7 1/2" total length |  | 1
| 21.2 | "Davis-Boyle" frame for tongue spatula for children, w. interchangeable inserts |  | 1
| 21.3 | "Davis-Boyle" frame for tongue spatula for adults, with interchangeable inserts |  | 1
| 21.4 | "Russel-Davis" tongue spatula extra narrow, blade 18x57 mm |  | 1
| 21.5 | "Russel-Davis" tongue spatula extra narrow, blade 22x67 mm |  | 1
| 21.6 | "Russel-Davis" tongue spatula extra narrow, blade 25x75 mm |  | 1
| 21.7 | "Russel-Davis" tongue spatula blade 38x85 mm |  | 1
| 21.8 | "Russel-Davis" tongue spatula blade 40x92 mm |  | 1
| 21.9 | "Abraham" tonsil holding forceps 205 mm 8 1/8" total length |  | 1
| 21.10 | "Hurd" tonsil dissector and retractor 225 mm, 9" total length |  | 1
| 21.11 | "Yankauer" suction tube curved, Ø 10 mm 275 mm, 11" total length |  | 1
| 21.12 | "Metzenbaum" power TC diss. scissors bl./bl., cur., tip 2,8 mm 200 mm, 8" total length |  | 1
| 21.13 | "Eves" tonsill snare 250 mm, 10" total length |  | 1
| 21.14 | Wire loops 100 pcs / package, 0,4 x 40 mm |  | 1
| 21.15 | Tonsil forceps "Schnidt" (Boettcher) strongly curved 185mm, 7 1/4" long |  | 2
| 21.16 | "Nager" tonsil retractor blade 13x15 mm 230 mm, 9 1/8" total length |  | 1
| 21.17 | Surgical forceps, standard str., tip 3,0 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 1
| 21.18 | "Hegar-Mayo" power TC needle holder straight, tip 2,5 mm, needle size 1-4 mm 180 mm, 7 1/8" total length |  | 1
| 21.19 | Laboratory bowl, heavy-weight model B8, capacity 1,0 L, Ø 150 mm 75 mm height |  | 1
| 21.20 | Cont. 1/2 bottom non-perforated 310x272x116 mm |  | 1
| 21.21 | Cont. 1/2 flex lid system red, lid pps biostop system 310x280 mm |  | 1
| 21.22 | Wire basket tray 1/2, half size with fall handle 255x250x50 mm |  | 1
| 21.23 | Silicone mat blue perforated 250 x 235 mm |  | 1
22 |  | TRACHEOTOMY SET |  | 2 |  | 0
| 22.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 22.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 4
| 22.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 2
| 22.4 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 1
| 22.5 | "Titanit" Metzenbaum-Fino scissors curved 145 mm 5 3/4" total length |  | 1
| 22.6 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 22.7 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 22.8 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 22.9 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 22.10 | "Halsted-Mosquito" haemostatic forceps cur.,tip 1,5mm,1x2 teeth, serrated jaw 125 mm, 5" total length |  | 4
| 22.11 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 4
| 22.12 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 2
| 22.13 | "Pean" haemostatic forceps straight, tip 2,5 mm, serrated jaw 140 mm, 5 1/2" total length |  | 2
| 22.14 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,0 mm, serrated jaw 210 mm, 8 1/4" total length |  | 1
| 22.15 | "Baby-Mixter" dissecting forceps curved, tip 1,4 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 22.16 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 22.17 | "Langenbeck" retractor blunt, blade 33x14 mm 210 mm, 8 1/4" total length |  | 2
| 22.18 | "Roux" retractor, double ended 19x22 / 23x30 mm 145 mm, 5 3/4" total length |  | 2
| 22.19 | Wound-/trachea retractor large curve, 1 tooth, sharp 165 mm, 6 1/2" total length |  | 2
| 22.20 | Wound-/trachea retractor large curve, 4 teeth, 7,5x14,0 mm, sharp 165 mm, 6 1/2" total length |  | 2
| 22.21 | Laborde tracheal dilator delicate pattern, curved mouth 125 mm, 5'' total length |  | 1
| 22.22 | "Weitlaner" self-retaining retractor sharp, 3/4 teeth 135 mm, 5 1/8" total length |  | 1
| 22.23 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 22.24 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 22.25 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 22.26 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
23 |  | LAPAROTOMY SET |  | 5 |  | 0
| 23.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 23.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 23.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 4
| 23.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 1
| 23.5 | Scalpel handle No.4L for blades fig. 18-36 215 mm, 8 1/2" total length |  | 2
| 23.6 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 1
| 23.7 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 23.8 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 23.9 | Titanit "Metzenbaum-Fino" diss.scissors curved 230 mm 9 1/8" total length |  | 1
| 23.10 | "Metzenbaum-Fino" power TC diss.scissors bl./bl., curved, very delicate shanks 260 mm 10 1/4" total length |  | 1
| 23.11 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 23.12 | Power TC ligature scissors, curved, wavecut, 230 mm 9 1/8" total length |  | 1
| 23.13 | "Lexer" power TC dissecting scissor bl./bl., cur., tip 2,8 mm 165 mm, 6 1/2" total length |  | 1
| 23.14 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 23.15 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 23.16 | Surgical forceps str., tip 1,8 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 23.17 | Surgical forceps str., tip 2,2 mm, 1x2 teeth, flat handle 250 mm, 10" total length |  | 2
| 23.18 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 23.19 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 23.20 | "DeBakey" atr. forceps, straight safety-pin with press-stopper, tip 2,0 m 240 mm, 9 1/2" total length |  | 2
| 23.21 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 6
| 23.22 | "Pean" haemostatic forceps straight, tip 2,5 mm, serrated jaw 140 mm, 5 1/2" total length |  | 6
| 23.23 | "Halsted-Mosquito" haemostatic forceps cur.,tip 1,5mm,1x2 teeth, serrated jaw 125 mm, 5" total length |  | 6
| 23.24 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 10
| 23.25 | "Halsted" haemostatic forceps straight, tip 1,7 mm, serrated jaw 185 mm, 7 1/4" total length |  | 2
| 23.26 | "Halsted" haemostatic forceps str., tip 2,2mm,1x2 teeth, serrated jaw 185 mm, 7 1/4" total length |  | 2
| 23.27 | "Kocher-Ochsner" haemostatic forceps str., tip 3,4mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 2
| 23.28 | "Kocher-Ochsner" haemostatic forceps str.,tip 3,8mm,1x2 teeth, serrated jaw 240 mm, 9 1/2" total length |  | 2
| 23.29 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 4
| 23.30 | "Mikulicz" peritoneum forceps curv.,tip 2,7 mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 6
| 23.31 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,5 mm, serrated jaw 230 mm, 9 1/8" total length |  | 2
| 23.32 | "Overholt" haemostic and dissecting forceps delicate, fig.1, 295mm, 11 3/4" total length |  | 2
| 23.33 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,5 mm, serrated jaw 270 mm, 10 3/4" total length |  | 2
| 23.34 | "Mixter" dissect./ligature forceps curved, tip 2,2 mm, serrated jaw 230 mm, 9 1/8" total length |  | 2
| 23.35 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 2
| 23.36 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 2
| 23.37 | "Wagensteen" power TC needle holder 275 mm, 11" total length |  | 1
| 23.38 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 180 mm, 7 1/8" total length |  | 2
| 23.39 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 255 mm, 10 1/8" total length |  | 2
| 23.40 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 23.41 | "Langenbeck" retractor 30x14 mm 210 mm, 8 1/4" total length |  | 2
| 23.42 | "Langenbeck" retractor blunt, blade 40x11 mm 210 mm, 8 1/4" total length |  | 2
| 23.43 | "Kocher-Langenbeck" retractor blunt, blade 55x11 mm 215 mm, 8 1/2" total length |  | 2
| 23.44 | "Volkmann" retractor 4 teeth, 8,5x19 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 23.45 | "Fritsch" abdominal retractor fig. 4, blunt, blade 45x75 mm 250 mm, 10" total length |  | 2
| 23.46 | "Mikulicz" abdominal retractor blunt, blade 151x50 mm 250 mm, 10" total length |  | 2
| 23.47 | "Mikulicz" abdominal retractor blunt, blade 180x50 mm 270 mm, 10 3/4" total length |  | 1
| 23.48 | "Haberer" abdominal/intestinal spatula width of endings 40 mm a. 50 mm, malleab 280 mm, 11 1/8" total length |  | 1
| 23.49 | atraumatic ovum forceps straight, DeBakey serration 250 mm, 10" total length |  | 2
| 23.50 | "Allis" atr. intestinal grasping forceps mouth 8,6 mm 255 mm, 10 1/8" total length |  | 2
| 23.51 | "Volkmann" bone curette fig.0, mouth 5,2x7,3 mm 170 mm, 6 3/4" total length |  | 1
| 23.52 | "Nelaton" probe curved, bluntt 160 mm, 6 1/4" total length |  | 1
| 23.53 | Double ended probe tip Ø 1,5 mm 200 mm, 8" total length |  | 1
| 23.54 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 23.55 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 23.56 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 23.57 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 23.58 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 23.59 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
| 23.60 | Silicone mat, 440x230mm |  | 1
24 |  | AMPUTATION SET |  | 2 |  | 0
| 24.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 24.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 24.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 4
| 24.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 2
| 24.5 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 24.6 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 24.7 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 24.8 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 24.9 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 24.10 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 24.11 | Surgical forceps, standard str., tip 3,0 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 24.12 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 24.13 | "Rochester-Pean" haemostatic forceps straight, tip 3,0 mm, serrated jaw 160 mm, 6 1/4" total length |  | 6
| 24.14 | "Kocher-Ochsner" haemostatic forceps str., tip 2,7mm,1x2 teeth, serrated jaw 160 mm, 6 1/4" total length |  | 6
| 24.15 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 2
| 24.16 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,3 mm, serrated jaw 220 mm, 8 3/4" total length |  | 1
| 24.17 | Power TC needle holder a. Hegar strong model 205 mm 8 1/8" total length |  | 2
| 24.18 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 2
| 24.19 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 24.20 | "Langenbeck" retractor blunt, blade 40x13 mm 210 mm, 8 1/4" total length |  | 2
| 24.21 | Bone retractor, strong, 1 tooth 245mm / 9 3/4" total length |  | 1
| 24.22 | "Körte" retractor 8-teeth, sharp, blade 26x40 mm 240 mm, 9 1/2" total length |  | 2
| 24.23 | Bone levers, 17mm width 240 mm, 9 1/2" total length |  | 2
| 24.24 | Hohmann bone lever 13/35 mm width 285 mm, 11 1/4" total length |  | 2
| 24.25 | "Liston" amputation knife blade 190 mm 330 mm, 13 1/8" total length |  | 1
| 24.26 | "Catlin" amputation knife double-edged, 160 mm cutting length, 320 mm, 12 3/4" total length |  | 1
| 24.27 | "Percy" amputation retractor remov. hand., inner Ø 43mm, Ø tot. 170 m 170 mm, 6 3/4" total length |  | 1
| 24.28 | "Bier" amputation saw with sawblade, blade 270 mm 420 mm, 16 3/4" total length |  | 1
| 24.29 | Bone file and rasp 20,0 mm wide 220 mm, 8 3/4" total length |  | 1
| 24.30 | "Lambotte" raspatory 20,0 mm wide 210 mm, 8 1/4" total length |  | 1
| 24.31 | "Fergusson" bone holding forceps 205mm 8 1/8" total length |  | 1
| 24.32 | "Röttgen-Ruskin" bone rongeur curved 240 mm, 9 1/2" total length |  | 1
| 24.33 | "Stille-Liston" bone cutter straight, length of cut 40,0 mm 280 mm, 11 1/8" total length |  | 1
| 24.34 | "Volkmann" bone curette fig.1, mouth 6,8x9,5 mm 170 mm, 6 3/4" total length |  | 1
| 24.35 | Double ended probe tip Ø 2,0 mm 200 mm, 8" total length |  | 1
| 24.36 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 24.37 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 24.38 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 24.39 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 24.40 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 24.41 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
25 |  | BASIC PLASTIC SET |  | 2 |  | 0
| 25.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 25.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 25.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 4
| 25.4 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 2
| 25.5 | "Titanit" Iris scissors, sh./sh., straight, 110 mm, 4 1/4" total length |  | 1
| 25.6 | "Titanit" Tuula Pankreas diss. scissors curved, 150 mm 6" total length |  | 1
| 25.7 | "Jameson" dissecting scissors curved, 175 mm 7 " total length |  | 1
| 25.8 | "Titanit" Metzenbaum-Fino scissors curved 145 mm 5 3/4" total length |  | 1
| 25.9 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 25.10 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 25.11 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 25.12 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 25.13 | "Micro-Adson" anatomic forceps str., tip 1,2 mm, serrated, flat handle 120 mm, 4 3/4" total length |  | 2
| 25.14 | "Micro-Adson" surgical forceps str., tip 0,9 mm, 1x2 teeth, flat handle 120 mm, 4 3/4" total length |  | 2
| 25.15 | "DeBakey" atr. forceps, straight safety-pin, tip 1,5mm 155 mm, 6 1/8" total length |  | 2
| 25.16 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 25.17 | "Baby-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 100 mm, 4" total length |  | 6
| 25.18 | "Micro-Halsted" haemostatic forceps curved, tip 0,8 mm, serrated jaw 125 mm, 5" total length |  | 12
| 25.19 | "Micro-Halsted" haemostatic forceps str.,tip 1,2mm,1x2 teeth, serrated jaw 125 mm, 5" total length |  | 10
| 25.20 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 2
| 25.21 | "Pean" haemostatic forceps straight, tip 2,5 mm, serrated jaw 140 mm, 5 1/2" total length |  | 2
| 25.22 | sponge forceps curved, with ratchet 175 mm, 7" total length |  | 1
| 25.23 | "Czerny" uterine tenaculum forceps 4x4 teeth, straight, 11 mm wide 200 mm, 8" total length |  | 1
| 25.24 | "Baby-Mixter" dissecting forceps curved, tip 1,4 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 25.25 | "Ryder" power TC needle holder straight, mouth 1,2 mm, serrated jaw 135 mm, 5 1/4" total length |  | 1
| 25.26 | "Halsey" power TC needle holder straight, tip 1,5 mm, needle size 5-7 mm 130 mm, 5 1/8" total length |  | 1
| 25.27 | Power TC needle holder a. Baby-Crile- Wood, straight, jaw 18 mm 150 mm 6" total length |  | 1
| 25.28 | "Hegar-Mayo" power TC needle holder straight, tip 2,5 mm, needle size 1-4 mm 180 mm, 7 1/8" total length |  | 1
| 25.29 | "Langenbeck" retractor 30x14 mm 210 mm, 8 1/4" total length |  | 2
| 25.30 | "Langenbeck" retractor blunt, blade 40x11 mm 210 mm, 8 1/4" total length |  | 2
| 25.31 | Langenbeck retractor 60x20 mm 210 mm, 8 1/4" total length |  | 1
| 25.32 | "Volkmann" retractor 4 teeth, 8,5x19 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 25.33 | "Volkmann" retractor 6 teeth, 8,5x29 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 25.34 | "Senn-Miller" retractor, double ended 8x7 mm (3-teeth sh.) & 18x5,5 mm hooks 160 mm, 6 1/4" total length |  | 2
| 25.35 | Wound-/trachea retractor large curve, 1 tooth, sharp 165 mm, 6 1/2" total length |  | 1
| 25.36 | Wound-/trachea retractor large curve, 1 tooth, blunt 165 mm, 6 1/2" total length |  | 1
| 25.37 | "Desmarres" lid retractor 9x12 mm 140 mm, 5 1/2" total length |  | 2
| 25.38 | "Desmarres" lid retractor 12x16 mm 140 mm, 5 1/2" total length |  | 2
| 25.39 | "Gilles" skin retractor small, Ø 2,5mm, str., sharp 180 mm, 7 1/8" total length |  | 2
| 25.40 | "Gilles" skin retractor large, Ø 3,5mm, str., sharp 180 mm, 7 1/8" total length |  | 2
| 25.41 | "Volkmann" bone curette fig.00, mouth 4,4x6,5 mm 170 mm, 6 3/4" total length |  | 1
| 25.42 | "Volkmann" bone curette fig.1, mouth 6,8x9,5 mm 170 mm, 6 3/4" total length |  | 1
| 25.43 | "Volkmann" bone curette fig.3, mouth 10,0x14,5 mm 170 mm, 6 3/4" total length |  | 1
| 25.44 | Double ended probe tip Ø 1,0 mm 145 mm, 5 3/4" total length |  | 1
| 25.45 | Double ended probe tip Ø 1,5 mm 160 mm, 6 1/4" total length |  | 1
| 25.46 | Double ended probe tip Ø 2,0 mm 180 mm, 7 1/8" total length |  | 1
| 25.47 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 25.48 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 25.49 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 25.50 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 25.51 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 25.52 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
26 |  | DRESSING SET SMALL |  | 10 |  | 0
| 26.1 | "Gross" sponge-/dressing forceps with lock, curved 200 mm, 8" total length |  | 1
| 26.2 | "Backhaus" towel clamp 90 mm, 3 1/2" total length |  | 4
| 26.3 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 1
| 26.4 | Anatomic tissue forceps str., tip 2,2 mm, serrated, flat handle 145 mm, 5 3/4" total length |  | 1
| 26.5 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 1
| 26.6 | "Metzenbaum-Fino" power TC diss.scissors bl./bl., curved, 145 mm 5 3/4" total length |  | 1
| 26.7 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 26.8 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 4
| 26.9 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 2
| 26.10 | Retractor, small curve,4 prongs sharp, 165 mm, 6 1/2" total length |  | 2
| 26.11 | "Desmarres" lid retractor 11x14 mm 140 mm, 5 1/2" total length |  | 2
| 26.12 | "Weitlaner" self-retaining retractor sharp, 2/3 teeth 110 mm, 4 1/4" total length |  | 1
| 26.13 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 26.14 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
27 |  | MAMMECTOMY SET |  | 1 |  | 0
| 27.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 27.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 4
| 27.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 4
| 27.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 2
| 27.5 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 1
| 27.6 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 27.7 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 27.8 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 27.9 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 27.10 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 27.11 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 27.12 | Surgical forceps str., tip 1,8 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 27.13 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 27.14 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 27.15 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 6
| 27.16 | Dissecting forceps Crile, straight 1x2 teeth, 160 mm, 6 1/4" total length |  | 6
| 27.17 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 4
| 27.18 | Sponge forceps with ratchet curved 235 mm, 9 1/4" total length |  | 2
| 27.19 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,0 mm, serrated jaw 210 mm, 8 1/4" total length |  | 2
| 27.20 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,3 mm, serrated jaw 220 mm, 8 3/4" total length |  | 2
| 27.21 | "Baby-Mixter" dissecting forceps curved, tip 1,4 mm, serrated jaw 180 mm, 7 1/8" total length |  | 2
| 27.22 | "Lahey" organ-and tissue grasping forceps, 5X5 teeth, 155mm, 6 1/8" total length |  | 1
| 27.23 | "Czerny" uterine tenaculum forceps 4x4 teeth, straight, 11 mm wide 200 mm, 8" total length |  | 1
| 27.24 | "Leland-Jones" atr. peripheral-clamp 30° ang, tip 2,5mm, 72mm serrated jaw 190 mm, 7 1/2" total length |  | 1
| 27.25 | "DeBakey-Derra" atr. baby anast. clamp large 170 mm, 6 3/4" total length |  | 1
| 27.26 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 27.27 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 1
| 27.28 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 27.29 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 27.30 | "Volkmann" retractor 6 teeth, 8,5x29 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 27.31 | "Kocher-Langenbeck" retractor blunt, blade 55x11 mm 215 mm, 8 1/2" total length |  | 1
| 27.32 | "Kocher-Langenbeck" retractor blunt, blade 80x16 mm 215 mm, 8 1/2" total length |  | 1
| 27.33 | "Brunner" retractor straight, blunt, blade 120x25 mm 275 mm, 11" total length |  | 1
| 27.34 | "Adson" self-retaining retractor straight, sharp, 4/4 teeth 190 mm, 7 1/2" total length |  | 1
| 27.35 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 27.36 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 27.37 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 27.38 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 27.39 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 27.40 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
28 |  | MINOR BASIC SET, GENERAL |  | 5 |  | 0
| 28.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 28.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 4
| 28.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 2
| 28.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 1
| 28.5 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 1
| 28.6 | "Titanit" Metzenbaum-Fino scissors curved 145 mm 5 3/4" total length |  | 1
| 28.7 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 28.8 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 28.9 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 28.10 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 28.11 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 28.12 | Surgical forceps str., tip 1,7 mm, 1x2 teeth, flat handle 180 mm, 7 1/8" total length |  | 2
| 28.13 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 28.14 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 4
| 28.15 | "Pean" haemostatic forceps straight, tip 2,5 mm, serrated jaw 140 mm, 5 1/2" total length |  | 4
| 28.16 | "Halsted-Mosquito" haemostatic forceps cur.,tip 1,5mm,1x2 teeth, serrated jaw 125 mm, 5" total length |  | 4
| 28.17 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 4
| 28.18 | "Kocher-Ochsner" haemostatic forceps str., tip 3,4mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 2
| 28.19 | "Baby-Mixter" dissecting forceps curved, tip 1,4 mm, serrated jaw 140 mm, 5 1/2" total length |  | 1
| 28.20 | "Baby-Mixter" dissecting forceps curved, tip 1,4 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 28.21 | "Halsted" haemostatic forceps curved, tip 1,7 mm, serrated jaw 185 mm, 7 1/4" total length |  | 2
| 28.22 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 28.23 | "Hegar-Mayo" power TC needle holder straight, tip 2,5 mm, needle size 1-4 mm 180 mm, 7 1/8" total length |  | 2
| 28.24 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 28.25 | "Langenbeck" retractor 30x14 mm 210 mm, 8 1/4" total length |  | 2
| 28.26 | "Langenbeck" retractor blunt, blade 40x11 mm 210 mm, 8 1/4" total length |  | 2
| 28.27 | "Volkmann" retractor 4 teeth, 8,5x19 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 28.28 | Wound-/trachea retractor large curve, 2 teeth, 7,5x7,5mm, sharp 165 mm, 6 1/2" total length |  | 2
| 28.29 | Cushing wound retractor fig. 2, blunt, 1 blade 9x10 mm 200 mm, 8" total length |  | 2
| 28.30 | "Allis" atr. intestinal grasping forceps straight, tip 6,5 mm, DeBakey teeth 155 mm, 6 1/8" total length |  | 1
| 28.31 | "Volkmann" bone curette fig.0, mouth 5,2x7,3 mm 170 mm, 6 3/4" total length |  | 1
| 28.32 | "Nelaton" probe curved, bluntt 160 mm, 6 1/4" total length |  | 1
| 28.33 | Double ended probe tip Ø 1,5 mm 145 mm, 5 3/4" total length |  | 1
| 28.34 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 28.35 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 28.36 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 28.37 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 28.38 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 28.39 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
29 |  | UROLOGY BLADDER SET |  | 1 |  | 0
| 29.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 29.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 29.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 6
| 29.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 1
| 29.5 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 1
| 29.6 | Scalpel handle No.4L for blades fig. 18-36 215 mm, 8 1/2" total length |  | 1
| 29.7 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 210 mm, 8 1/4" total length |  | 1
| 29.8 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 29.9 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 29.10 | Titanit "Metzenbaum-Fino" diss.scissors curved 230 mm 9 1/8" total length |  | 1
| 29.11 | "Titanit" Metzenbaum-Fino scissors curved, 260 mm 10 1/4" total length |  | 1
| 29.12 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 29.13 | Power TC ligature scissors, curved, wavecut, 230 mm 9 1/8" total length |  | 1
| 29.14 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 29.15 | "Titanit" "Slim Line" Potts scissors 60° ang., 10 mm blade length, ring handl 173 mm 6 3/4" total length |  | 1
| 29.16 | "Titanit" vascular scissors DeBakey, 60° angled, 225 mm 9" total length |  | 1
| 29.17 | "Titanit" Cooper surgical scissors sh./bl., straight, 145 mm 5 3/4" total length |  | 1
| 29.18 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 29.19 | Surgical forceps str., tip 1,8 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 29.20 | Surgical forceps str., tip 2,2 mm, 1x2 teeth, flat handle 250 mm, 10" total length |  | 2
| 29.21 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 29.22 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 29.23 | "DeBakey" atr. forceps, straight safety-pin with press-stopper, tip 2,0 m 240 mm, 9 1/2" total length |  | 2
| 29.24 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 10
| 29.25 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 10
| 29.26 | "Halsted" haemostatic forceps str., tip 2,2mm,1x2 teeth, serrated jaw 185 mm, 7 1/4" total length |  | 2
| 29.27 | "Kocher-Ochsner" haemostatic forceps str., tip 3,4mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 2
| 29.28 | "Kocher-Ochsner" haemostatic forceps str.,tip 3,8mm,1x2 teeth, serrated jaw 240 mm, 9 1/2" total length |  | 2
| 29.29 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 4
| 29.30 | "Mikulicz" peritoneum forceps curv.,tip 2,7 mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 6
| 29.31 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,0 mm, serrated jaw 210 mm, 8 1/4" total length |  | 2
| 29.32 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,5 mm, serrated jaw 230 mm, 9 1/8" total length |  | 2
| 29.33 | "Overholt-Geissendoerfer" dissecting- forceps, fig.3, 275 mm, 11" |  | 2
| 29.34 | "Gemini" dissect./ligature forceps curved, tip 1,8 mm, serrated jaw 230 mm, 9 1/8" total length |  | 2
| 29.35 | "Gemini" dissect./ligature forceps curved, tip 1,9 mm, serrated jaw 280 mm, 11 1/8" total length |  | 2
| 29.36 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 29.37 | Power TC needle holder a. Hegar strong model 205 mm 8 1/8" total length |  | 2
| 29.38 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 29.39 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 230 mm, 9 1/8" total length |  | 1
| 29.40 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 255 mm, 10 1/8" total length |  | 1
| 29.41 | Power TC needle holder a. Stratte s-curved, tip 2,0 mm, serrated jaw 230 mm 9 1/8" total length |  | 1
| 29.42 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 29.43 | "Kocher" retractor 60x25 mm 230 mm, 9 1/8" total length |  | 2
| 29.44 | "Volkmann" retractor 4 teeth, 8,5x19 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 29.45 | "Fritsch" abdominal retractor fig. 4, blunt, blade 45x75 mm 250 mm, 10" total length |  | 2
| 29.46 | "Mikulicz" abdominal retractor blunt, blade 120x50 mm 250 mm, 10" total length |  | 1
| 29.47 | "Mikulicz" abdominal retractor blunt, blade 151x50 mm 250 mm, 10" total length |  | 1
| 29.48 | "Kelly" wound retractor 150 x 39 mm 260 mm, 10 1/4" total length |  | 1
| 29.49 | "Haberer" abdominal/intestinal spatula width of endings 40 mm a. 50 mm, malleab 280 mm, 11 1/8" total length |  | 1
| 29.50 | "Legueu" bladder spatula curved, blade 35 mm wide 260 mm, 10 1/4" total length |  | 2
| 29.51 | "Tübinger model" vaginal retractor blunt, 1 blade 95x20 mm 265 mm, 10 1/2" total length |  | 1
| 29.52 | "Simon" vaginal retractor blunt, 1 blade 115x26 mm 275 mm, 11" total length |  | 1
| 29.53 | "Kristeller" set vaginal specula 110x30 mm, upper & lower blade 220 mm, 8 3/4" total length |  | 1
| 29.54 | Pelvis retractor "Thorlakson curved, blade with holes,170x50 mm,ergonomic gri 330 mm, 13 1/8" total length |  | 1
| 29.55 | Pelvis retractor "Thorlakson curved, blade with holes,170x65mm,ergonomic grip 330 mm, 13 1/8" total length |  | 1
| 29.56 | Explor. Coagul.. hook, Caspar, 7mm isolated 245 mm, 9 3/4" total length, |  | 1
| 29.57 | Cushing wound retractor fig. 2, blunt, 1 blade 9x10 mm 200 mm, 8" total length |  | 1
| 29.58 | "Emmet" fistel hook, 1 tooth 220 mm, 8 3/4" total length |  | 2
| 29.59 | atraumatic ovum forceps straight, DeBakey serration 250 mm, 10" total length |  | 2
| 29.60 | "Allis" atr. intestinal grasping forceps mouth 6,5 mm 220 mm, 8 3/4" total length |  | 2
| 29.61 | "Allis" atr. intestinal grasping forceps mouth 8,6 mm 255 mm, 10 1/8" total length |  | 2
| 29.62 | "Stockmann" penis clamp 70 mm, 2 3/4" total length |  | 1
| 29.63 | Catheter guide curved 490 mm, 19 1/2", total length |  | 1
| 29.64 | Atrauma uretrine clamp 250 mm, 10" long |  | 1
| 29.65 | Probe Ø 2,5 mm, hexagonal handle 300 mm, 12" total length |  | 1
| 29.66 | Double ended probe tin, malleable, tip Ø 3,5 mm 300 mm, 12" total length |  | 1
| 29.67 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 29.68 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 29.69 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 29.70 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 29.71 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 29.72 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
30 |  | UROLOGY KIDNEY SET |  | 1 |  | 0
| 30.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 30.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 30.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 6
| 30.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 1
| 30.5 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 1
| 30.6 | Scalpel handle No.4L for blades fig. 18-36 215 mm, 8 1/2" total length |  | 1
| 30.7 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 210 mm, 8 1/4" total length |  | 1
| 30.8 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 30.9 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 30.10 | Titanit "Metzenbaum-Fino" diss.scissors curved 230 mm 9 1/8" total length |  | 1
| 30.11 | "Titanit" Metzenbaum-Fino scissors curved, 260 mm 10 1/4" total length |  | 1
| 30.12 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 30.13 | Power TC ligature scissors, curved, wavecut, 230 mm 9 1/8" total length |  | 1
| 30.14 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 30.15 | "Titanit" "Slim Line" Potts scissors 60° ang., 10 mm blade length, ring handl 173 mm 6 3/4" total length |  | 1
| 30.16 | "Titanit" vascular scissors DeBakey, 60° angled, 225 mm 9" total length |  | 1
| 30.17 | "Titanit" Cooper surgical scissors sh./bl., straight, 145 mm 5 3/4" total length |  | 1
| 30.18 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 30.19 | Surgical forceps str., tip 1,8 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 30.20 | Surgical forceps str., tip 2,2 mm, 1x2 teeth, flat handle 250 mm, 10" total length |  | 2
| 30.21 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 30.22 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 30.23 | "DeBakey" atr. forceps, straight safety-pin with press-stopper, tip 2,0 m 240 mm, 9 1/2" total length |  | 2
| 30.24 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 10
| 30.25 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 10
| 30.26 | "Halsted" haemostatic forceps str., tip 2,2mm,1x2 teeth, serrated jaw 185 mm, 7 1/4" total length |  | 2
| 30.27 | "Kocher-Ochsner" haemostatic forceps str., tip 3,4mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 2
| 30.28 | "Kocher-Ochsner" haemostatic forceps str.,tip 3,8mm,1x2 teeth, serrated jaw 240 mm, 9 1/2" total length |  | 2
| 30.29 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 4
| 30.30 | "Mikulicz" peritoneum forceps curv.,tip 2,7 mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 6
| 30.31 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,0 mm, serrated jaw 210 mm, 8 1/4" total length |  | 2
| 30.32 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,5 mm, serrated jaw 230 mm, 9 1/8" total length |  | 2
| 30.33 | "Mixter" dissect./ligature forceps curved, tip 2,2 mm, serrated jaw 230 mm, 9 1/8" total length |  | 2
| 30.34 | "Overholt-Geissendoerfer" dissecting- forceps, fig.3, 275 mm, 11" |  | 2
| 30.35 | "Gemini" dissect./ligature forceps curved, tip 1,8 mm, serrated jaw 230 mm, 9 1/8" total length |  | 2
| 30.36 | "O´Shaugnessy"dissect./ligature forceps curved, tip 3,0 mm, serrated jaw 230 mm, 9 1/8" total length |  | 2
| 30.37 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 30.38 | Power TC needle holder a. Hegar strong model 205 mm 8 1/8" total length |  | 2
| 30.39 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 30.40 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 230 mm, 9 1/8" total length |  | 1
| 30.41 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 255 mm, 10 1/8" total length |  | 1
| 30.42 | Power TC needle holder a. Stratte s-curved, tip 2,0 mm, serrated jaw 230 mm 9 1/8" total length |  | 1
| 30.43 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 30.44 | "Volkmann" retractor 4 teeth, 8,5x19 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 30.45 | "Fritsch" abdominal retractor fig. 4, blunt, blade 45x75 mm 250 mm, 10" total length |  | 2
| 30.46 | "Mikulicz" abdominal retractor blunt, blade 120x50 mm 250 mm, 10" total length |  | 1
| 30.47 | "Mikulicz" abdominal retractor blunt, blade 151x50 mm 250 mm, 10" total length |  | 1
| 30.48 | "Mikulicz" abdominal retractor blunt, blade 180x50 mm 270 mm, 10 3/4" total length |  | 1
| 30.49 | "Haberer" abdominal/intestinal spatula width of endings 40 mm a. 50 mm, malleab 280 mm, 11 1/8" total length |  | 1
| 30.50 | Cushing wound retractor fig. 2, blunt, 1 blade 9x10 mm 200 mm, 8" total length |  | 1
| 30.51 | Guyon atraumatic DeBakey kidney forceps, fig.1 240 mm, 9 1/2" total length |  | 1
| 30.52 | Atraumatic kidney clamp "Guyon" 230 mm, 9 1/8" total length |  | 1
| 30.53 | "DeBakey-Satinsky" atr vascular clamp branche backw. cur., doub. angled 220 mm, 8 3/4" total length |  | 1
| 30.54 | "DeBakey-Satinsky" atr vascular clamp fig.2, branche backw. cur., doub. ang. 240 mm, 9 1/2" total length |  | 1
| 30.55 | "DeBakey-Satinsky" atr vascular clamp fig.3, branche backw. cur., doub. ang. 250 mm, 10" total length |  | 1
| 30.56 | "DeBakey" atr. dissect./ligature forceps spoon shaped, tip 3,0 mm 230 mm, 9 1/8" total length |  | 1
| 30.57 | atraumatic ovum forceps straight, DeBakey serration 250 mm, 10" total length |  | 1
| 30.58 | "Allis" atr. intestinal grasping forceps mouth 6,5 mm 220 mm, 8 3/4" total length |  | 2
| 30.59 | "Allis" atr. intestinal grasping forceps mouth 8,6 mm 255 mm, 10 1/8" total length |  | 2
| 30.60 | KIDNEY STONE CLAMP RANDALL 225 MM, 9" TOTAL LENGTH |  | 1
| 30.61 | Randall kidney stone forceps 225mm 9" total length |  | 1
| 30.62 | Kidney stone forceps "Randall" 225mm, 9" total length |  | 1
| 30.63 | "Randall" kidney stone forceps 190 mm, 7 1/2" total length |  | 1
| 30.64 | Probe Ø 2,5 mm, hexagonal handle 300 mm, 12" total length |  | 1
| 30.65 | DeBakey vascular dilator Ø2,0 mm, malleable, hexagonal handle 190 mm, 7 1/2'' lotal length |  | 1
| 30.66 | "Luer-Körte" gall stone scoop fig. 000, malleable shaft, mouth 2,8 mm 320 mm, 12 3/4" total length |  | 1
| 30.67 | "Luer-Körte" gall stone scoop fig. 0, malleable shaft, mouth 4,3 mm 320 mm, 12 3/4" total length |  | 1
| 30.68 | Double ended probe tin, malleable, tip Ø 3,5 mm 300 mm, 12" total length |  | 1
| 30.69 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 30.70 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 30.71 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 30.72 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 30.73 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 30.74 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
31 |  | VARICOSE VEIN STRIPPING SET |  | 1 |  | 0
| 31.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 31.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 31.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 4
| 31.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 1
| 31.5 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 2
| 31.6 | Scalpel handle, straight, round handle for micro blades of SK0361-SK0417 135 mm, 5 1/2" total length |  | 1
| 31.7 | "Titanit" Metzenbaum-Fino scissors curved 145 mm 5 3/4" total length |  | 1
| 31.8 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 31.9 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 31.10 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 31.11 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 31.12 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 31.13 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 31.14 | Surgical forceps str., tip 1,8 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 31.15 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 31.16 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 31.17 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 6
| 31.18 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 10
| 31.19 | "Halsted-Mosquito" haemostatic forceps cur.,tip 1,5mm,1x2 teeth, serrated jaw 125 mm, 5" total length |  | 6
| 31.20 | "Baby-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 100 mm, 4" total length |  | 2
| 31.21 | "Kocher-Ochsner" haemostatic forceps str., tip 3,0 mm,1x2 teeth, serrated jaw 185 mm, 7 1/4" total length |  | 2
| 31.22 | "Baby-Mixter" dissecting forceps curved, tip 1,4 mm, serrated jaw 180 mm, 7 1/8" total length |  | 2
| 31.23 | Power TC needle holder a. Baby-Crile- Wood, straight, jaw 18 mm 150 mm 6" total length |  | 1
| 31.24 | "Hegar-Mayo" power TC needle holder straight, tip 2,5 mm, needle size 1-4 mm 180 mm, 7 1/8" total length |  | 2
| 31.25 | Retractor, small curve,4 prongs sharp, 165 mm, 6 1/2" total length |  | 2
| 31.26 | "Desmarres" lid retractor 11x14 mm 140 mm, 5 1/2" total length |  | 2
| 31.27 | "Langenbeck" retractor 30x14 mm 210 mm, 8 1/4" total length |  | 2
| 31.28 | "Langenbeck" retractor blunt, blade 40x11 mm 210 mm, 8 1/4" total length |  | 2
| 31.29 | "Kocher-Langenbeck" retractor blunt, blade 55x11 mm 215 mm, 8 1/2" total length |  | 2
| 31.30 | "Adson" self-retaining retractor straight, sharp, 4/4 teeth 190 mm, 7 1/2" total length |  | 1
| 31.31 | "Freer" elevator double en., 3,7/3,7 mm wide,sharp/blunt 185 mm, 7 1/4" total length |  | 1
| 31.32 | "Nabatoff" vein stripping iset complete |  | 1
| 31.33 | "Varady" phlebdissector round ends 180 mm, 7 1/8" total length |  | 1
| 31.34 | "Varady" phlebextractor, standard model fig. 1, with ball tip 180 mm, 7 1/8" total length |  | 1
| 31.35 | "Varady-Hofmann" phlebextractor short handle, No.3 180 mm, 7 1/8" total length |  | 1
| 31.36 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 31.37 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 31.38 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 31.39 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 31.40 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 31.41 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
32 |  | VASCULAR BASIC SET |  | 1 |  | 0
| 32.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 32.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 32.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 6
| 32.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 2
| 32.5 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 125 mm, 5" total length |  | 1
| 32.6 | "Titanit" Jameson-Werbe scissors curved, fine blunt baldes, 135 mm 5 1/4" total length |  | 1
| 32.7 | Power TC Iris- and ligature scissors, sh./bl., curved, 110 mm 4 1/4" total length |  | 1
| 32.8 | "Titanit" Potts-Smith scissors 45°ang., 25 mm blade length, ring handle 185 mm 7 1/4" total length |  | 1
| 32.9 | "Titanit" Metzenbaum-Fino scissors curved 145 mm 5 3/4" total length |  | 1
| 32.10 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 32.11 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 32.12 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 32.13 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 32.14 | "Titanit" Cooper surgical scissors sh./bl., straight, 145 mm 5 3/4" total length |  | 1
| 32.15 | "Glasgow" Potts scissors curv., blunt, delicate blades,fig.1 180 mm 7 1/8" total length |  | 1
| 32.16 | "Stussak" tunneling a. diss. scissors strongly curved,delicate, 33 mm BL 180 mm 7 1/8" total length |  | 1
| 32.17 | "Golf"neo intima scissors,strongly curv. bl./bl., fine flat blades, ring handle 200 mm 8" total length |  | 1
| 32.18 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 32.19 | Surgical forceps str., tip 1,8 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 32.20 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 32.21 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 32.22 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 6
| 32.23 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 6
| 32.24 | "Kocher-Ochsner" haemostatic forceps str., tip 3,4mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 2
| 32.25 | "Overholt" dissecting forceps curv., tip 1,8 mm, serrated jaw,delicate 215 mm, 8 1/2" total length |  | 2
| 32.26 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,0 mm, serrated jaw 210 mm, 8 1/4" total length |  | 2
| 32.27 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,3 mm, serrated jaw 220 mm, 8 3/4" total length |  | 2
| 32.28 | "Micro-Adson" haemostatic forceps curved, tip 1,0 mm, serrated jaw 140 mm, 5 1/2" total length |  | 1
| 32.29 | "Baby-Mixter" dissecting forceps curved, tip 1,4 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 32.30 | "Gemini" dissect./ligature forceps curved, tip 1,2 mm, serrated jaw 135 mm, 5 1/4" total length |  | 1
| 32.31 | "Gemini" dissect./ligature forceps curved, tip 1,4 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 32.32 | Power TC needle holder a. Baby-Crile- Wood, delicate serr.,straight,jaw 16,5mm 150 mm 6" total length |  | 1
| 32.33 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 32.34 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 2
| 32.35 | "Ryder" power TC needle holder straight, mouth 1,2 mm, serrated jaw 150 mm, 6" total length |  | 1
| 32.36 | "Elegance" micro needle holder fine Jaws, 1x3 ratchet ring handle, 225 mm 9" total length |  | 1
| 32.37 | Retractor, small curve,4 prongs sharp, 165 mm, 6 1/2" total length |  | 1
| 32.38 | "Desmarres" lid retractor 11x14 mm 140 mm, 5 1/2" total length |  | 2
| 32.39 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 32.40 | "Langenbeck" retractor 30x14 mm 210 mm, 8 1/4" total length |  | 2
| 32.41 | "Langenbeck" retractor blunt, blade 40x11 mm 210 mm, 8 1/4" total length |  | 2
| 32.42 | "Kocher-Langenbeck" retractor blunt, blade 55x11 mm 215 mm, 8 1/2" total length |  | 2
| 32.43 | "Volkmann" retractor 4 teeth, 8,5x19 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 32.44 | Cushing wound retractor fig. 2, blunt, 1 blade 9x10 mm 200 mm, 8" total length |  | 2
| 32.45 | Adson woundretractor 4X5 teeth semi sharp, spread width 267mm 330 mm, 13 1/8" total length |  | 1
| 32.46 | Woundretractor Weitlaner blunt 3X4 teeth, 240 mm 9 1/2" total length |  | 1
| 32.47 | "Adson-Baby" self-retaining retractor with joint, blunt, 3/4 teeth 140 mm, 5 1/2" total length |  | 1
| 32.48 | "Freer" elevator double en., 3,7/3,7 mm wide,sharp/blunt 185 mm, 7 1/4" total length |  | 1
| 32.49 | Intima dissector round handle, 3 mm wide, blunt 185 mm, 7 1/4" total length |  | 1
| 32.50 | "Davis" dissector curved, double ended, 245 mm, 9 3/4", total length |  | 1
| 32.51 | "Crile" micro and suture hook tip 90° curved, blunt, long mouth part 220 mm, 8 3/4" total length |  | 1
| 32.52 | Double ended probe tip Ø 1,5 mm 180 mm, 7 1/8" total length |  | 1
| 32.53 | "Schmied" vascular suction cannulae Ø 3 mm, malleable, silver plated 150 mm, 6" total length |  | 1
| 32.54 | "DeBakey" mini bulldog clamp str, serr. jaw 14mm, closing pre. 180g 60 mm, 2 1/4" total length |  | 2
| 32.55 | "DeBakey" mini bulldog clamp 55°gew., serrated jaw 20 mm 57 mm, 2 1/4" total length |  | 2
| 32.56 | "DeBakey" bulldog clamp straight, surrated jaw 20 mm 80 mm, 3 1/8" total length |  | 2
| 32.57 | "DeBakey" bulldog clamp curved, serrated jaw 27 mm 86 mm, 3 1/4" total length |  | 2
| 32.58 | "DeBakey" mini bulldog clamp cur, serr. jaw 17mm, closing pre. 180g 58 mm, 2 1/4" total length |  | 2
| 32.59 | "DeBakey" bulldog clamp curved, serrated jaw 18 mm 78 mm, 3 1/8" total length |  | 2
| 32.60 | "DeBakey" bulldog clamp curved, serrated jaw 27 mm 86 mm, 3 1/4" total length |  | 2
| 32.61 | "Müller" vascular clip, DeBakey mouth straight, closing pressure 350 g 95 mm, 3 3/4" total length |  | 2
| 32.62 | "Müller" vascular clip curved, toothing DeBakey 120 mm, 4 3/4" total length |  | 2
| 32.63 | "Müller" vascular clip, DeBakey mouth curved tip, closing pressure 350g 115 mm, 4 1/2" total length |  | 2
| 32.64 | "Leland-Jones" atr. peripheral-clamp branche backw. cur., 60° ang, tip 2,3 mm 185 mm, 7 1/4" total length |  | 2
| 32.65 | "DeBakey-Derra" atr. baby anast. clamp large 170 mm, 6 3/4" total length |  | 1
| 32.66 | "DeBakey" atr. baby vas. multipurpose cl 60° angle 160 mm, 6 1/4" total length |  | 1
| 32.67 | "Salzburger" spatula for carotid & delicate vessels 155 mm, 6 1/8" total length |  | 1
| 32.68 | Tunneler instrument set consisting of GF6043, GF6041 and GF6042 750 mm, 29 1/2" total length |  | 1
| 32.69 | "Allis" atr. intestinal grasping forceps straight, tip 6,5 mm, DeBakey teeth 155 mm, 6 1/8" total length |  | 2
| 32.70 | "Pelkmann" uterine dressing forceps curved 260 mm, 10 1/4" total length |  | 1
| 32.71 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 32.72 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 32.73 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 32.74 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 32.75 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 32.76 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
33 |  | THORACOTOMY SET (LUNG) |  | 1 |  | 0
| 33.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 33.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 12
| 33.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 8
| 33.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 2
| 33.5 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 210 mm, 8 1/4" total length |  | 1
| 33.6 | "Titanit" Metzenbaum scissors curved, 180 mm 7 1/8" total length |  | 1
| 33.7 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 33.8 | Titanit "Metzenbaum-Fino" diss.scissors curved 230 mm 9 1/8" total length |  | 1
| 33.9 | "Titanit" Metzenbaum-Fino scissors curved, 260 mm 10 1/4" total length |  | 1
| 33.10 | "Titanit" Metzenbaum-Fino scissors bl./bl., curved, 280 mm 11 1/8" total length |  | 1
| 33.11 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 33.12 | Power TC ligature scissors, curved, wavecut, 230 mm 9 1/8" total length |  | 1
| 33.13 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 33.14 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 33.15 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 33.16 | Surgical forceps str., tip 1,8 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 33.17 | Surgical forceps str., tip 2,2 mm, 1x2 teeth, flat handle 250 mm, 10" total length |  | 2
| 33.18 | Surgical forceps str., tip 2,5 mm, 1x2 teeth, flat handle 300 mm, 12" total length |  | 2
| 33.19 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 33.20 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 33.21 | "DeBakey" atr. forceps, straight safety-pin with press-stopper, tip 2,0 m 240 mm, 9 1/2" total length |  | 2
| 33.22 | "DeBakey" atr. forceps, str. Safety-Pin with Press-Stopper, tip 2,8mm 240 mm, 9 1/2" total length |  | 2
| 33.23 | "DeBakey" atr. forceps, str. Safety-Pin with Press-Stopper,tip 2,8mm 300 mm, 12" total length |  | 2
| 33.24 | "Halsted-Mosquito" haemostatic forceps curved, tip 1,0 mm, serrated jaw 125 mm, 5" total length |  | 6
| 33.25 | "Kocher" haemostatic forceps straight, tip 2,0 mm, 1x2 teeth,delicate 140 mm, 5 1/2" total length |  | 6
| 33.26 | "Kocher-Ochsner" haemostatic forceps str., tip 3,4mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 2
| 33.27 | "Kocher-Ochsner" haemostatic forceps str.,tip 3,8mm,1x2 teeth, serrated jaw 240 mm, 9 1/2" total length |  | 2
| 33.28 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 4
| 33.29 | "Mikulicz" peritoneum forceps curv.,tip 2,7 mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 4
| 33.30 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,0 mm, serrated jaw 210 mm, 8 1/4" total length |  | 1
| 33.31 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,3 mm, serrated jaw 220 mm, 8 3/4" total length |  | 1
| 33.32 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,5 mm, serrated jaw 230 mm, 9 1/8" total length |  | 1
| 33.33 | "Overholt-Geissendörfer" dissecting for. curved, tip 1,9 mm, serrated jaw 225 mm, 9" total length |  | 1
| 33.34 | "Overholt-Geissendörfer" dissecting for. curved, tip 1,7 mm, serrated jaw 260 mm, 10 1/4" total length |  | 1
| 33.35 | "Overholt-Geissendoerfer" dissecting- forceps, fig.3, 275 mm, 11" |  | 1
| 33.36 | "Rumel" dissecting forceps fig.4 240 mm, 9 1/2" total length |  | 1
| 33.37 | "Rumel" dissecting forceps curved, tip 3,2 mm, serrated jaw 240 mm, 9 1/2" total length |  | 1
| 33.38 | "Zenker" dissecting- and ligature forcep lightly curved 300 mm, 12" total length |  | 1
| 33.39 | atraumatic ovum forceps straight, DeBakey serration 250 mm, 10" total length |  | 6
| 33.40 | "DeBakey-Satinsky" atr vascular clamp branche backw. cur., doub. angled 220 mm, 8 3/4" total length |  | 1
| 33.41 | "DeBakey-Satinsky" atr vascular clamp fig.2, branche backw. cur., doub. ang. 240 mm, 9 1/2" total length |  | 1
| 33.42 | "DeBakey-Satinsky" atr vascular clamp fig.3, branche backw. cur., doub. ang. 250 mm, 10" total length |  | 1
| 33.43 | "DeBakey" atraumatic coarctation clamp branche 15° angled, tip 3,5 mm 240 mm, 9 1/2" total length |  | 1
| 33.44 | "DeBakey" atraumatic vascular clamp branche backw. curved, 60° angled 265 mm, 10 1/2" total length |  | 1
| 33.45 | "DeBakey" atraumatic multi purpose clamp 90° angled, jaws 74 mm 250 mm, 10" total length |  | 1
| 33.46 | organ-/tissue grasping forceps special serration with safety grip 230 mm, 9 1/8" total length |  | 1
| 33.47 | Chest tube forceps curved, tip 3,3 mm, without ratchet 355 mm, 14 1/8" total length |  | 1
| 33.48 | "Resano" atr. forceps, cross-action titanium, transmitted, serrated tip 300 mm, 12" total length |  | 1
| 33.49 | "Allis" atr. intestinal grasping forceps mouth 6,5 mm 220 mm, 8 3/4" total length |  | 2
| 33.50 | "Collin" grasping forceps fenestrated, 1x2 teeth 200 mm, 8" total length |  | 2
| 33.51 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 33.52 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 180 mm, 7 1/8" total length |  | 1
| 33.53 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 255 mm, 10 1/8" total length |  | 1
| 33.54 | "DeBakey" power TC needle holder straight, mouth 1,8 mm, serrated jaw 305 mm, 12 1/8" total length |  | 1
| 33.55 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 2
| 33.56 | Power TC needle holder and wire forceps 15,5 mouth length 175 mm, 7" total length |  | 1
| 33.57 | Power TC wire cutting pliers geb, wire: soft up to 2,0 mm, hard 1,7mm 180 mm, 7 1/8" total length |  | 1
| 33.58 | "Kocher" retractor 4-teeth, sharp, blade 15x20 mm 220 mm, 8 3/4" total length |  | 2
| 33.59 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 33.60 | "Langenbeck" retractor blunt, blade 40x11 mm 210 mm, 8 1/4" total length |  | 2
| 33.61 | "Sauerbruch" retractor blunt, blade 55x22 mm 225 mm, 9" total length |  | 2
| 33.62 | "Sauerbruch" retractor blunt, blade 72x19 mm 225 mm, 9" total length |  | 2
| 33.63 | "Coryllos" retractor , fig.3, 115X65mm 225mm, 9" total length |  | 2
| 33.64 | "Cushing" retractor 15X22 mm 250 mm total length |  | 1
| 33.65 | "Cushing" retractor 15X22 mm 250 mm total length |  | 1
| 33.66 | "DeBakey" sternum-/rib retractor complete, 3 pairs GF3210, GF3220, GF3230 |  | 1
| 33.67 | "Allison" lung spatula 30° curved, blade 54 mm wide 320 mm, 12 3/4" total length |  | 1
| 33.68 | Lung spatula curved, blade 137 mm wide 300 mm, 12" total length |  | 1
| 33.69 | "Harrington" retractor, figure 1 flexible, blunt, blade 123x45 mm 320 mm, 12 3/4" total length |  | 1
| 33.70 | "Brunner" rib shears backwards curved, open to the left 340 mm, 13 1/2" total length |  | 1
| 33.71 | "Doyen" raspatory, for adults curved to left, 30x25 mm 175 mm, 7" total length |  | 1
| 33.72 | "Doyen" raspatory, for adults curved to right, 30x25 mm 175 mm, 7" total length |  | 1
| 33.73 | "Farabeuf" raspatory curved, 12,5 mm wide 140 mm, 5 1/2" total length |  | 1
| 33.74 | "Semb" raspatory 12,0 mm wide, sharp 230 mm, 9 1/8" total length |  | 1
| 33.75 | "Semb" rib raspatory 13,0 mm wide 230 mm, 9 1/8" total length |  | 1
| 33.76 | "Semb" rib/bone holding forceps 8,0 mm wide 190 mm, 7 1/2" total length |  | 1
| 33.77 | "Stille-Liston" bone cutter 30° angeld 240 mm, 9 1/2" total length |  | 1
| 33.78 | "Luer-Stille" bone rongeur curved, 9,5x15 mm jaw 225 mm, 9" total length |  | 1
| 33.79 | "Bailey-Gibbon" rib contractor 215 mm, 8 1/2" total length |  | 1
| 33.80 | "Lebsche" sternum chisel 250 mm, 10" total length |  | 1
| 33.81 | Mallet head-Ø 30 mm, 350 g 245 mm, 9 3/4" total length |  | 1
| 33.82 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 33.83 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 33.84 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 33.85 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 33.86 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 33.87 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
34 |  | CHANGE OF BANDAGE SET |  | 5 |  | 0
| 34.1 | Anatomic tissue forceps str., tip 2,2 mm, serrated, flat handle 145 mm, 5 3/4" total length |  | 1
| 34.2 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 1
| 34.3 | "Lister" bandage- and cloth scissors ang., 1 blade with button 180 mm, 7 1/8" total length |  | 1
| 34.4 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 34.5 | Myrtle leaf probe tip Ø 2,0 mm 145 mm, 5 3/4" total length |  | 1
| 34.6 | "Pean" haemostatic forceps curved, tip 2,5 mm, serrated jaw 140 mm, 5 1/2" total length |  | 1
35 |  | CURETTAGE SET |  | 2 |  | 0
| 35.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 35.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 4
| 35.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 2
| 35.4 | "Kristeller" set vaginal specula 90x36 mm, upper & lower blade 175 mm, 7" total length |  | 1
| 35.5 | "Kristeller" set vaginal specula 110x20 mm, upper & lower blade 220 mm, 8 3/4" total length |  | 1
| 35.6 | "Doyen" vaginal specula, figure 2 blade 90x45 mm 240 mm, 9 1/2" total length |  | 1
| 35.7 | Scalpel handle No.7, mirrow finsh. for blades fig. 10-15, 40, 42 160 mm, 6 1/4" total length |  | 1
| 35.8 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 35.9 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 2
| 35.10 | "Schröder" uterine tenaculum forceps sharp 250 mm, 10" total length |  | 2
| 35.11 | Surgical forceps, standard str., tip 3,0 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 35.12 | "Sims" uterine probe graduated, malleable, Ø 4 mm 330 mm, 13 1/8" total length |  | 1
| 35.13 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 1
| 35.14 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 1
| 35.15 | Hegar uterine dilator Ø3,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.16 | Hegar uterine dilator Ø3,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.17 | Hegar uterine dilator Ø4,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.18 | Hegar uterine dilator Ø4,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.19 | Hegar uterine dilator Ø5,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.20 | Hegar uterine dilator Ø5,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.21 | Hegar uterine dilator Ø6,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.22 | Hegar uterine dilator Ø6,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.23 | Hegar uterine dilator Ø7,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.24 | Hegar uterine dilator Ø7,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.25 | Hegar uterine dilator Ø8,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.26 | Hegar uterine dilator Ø8,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.27 | Hegar uterine dilator Ø9,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.28 | Hegar uterine dilator Ø9,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.29 | Hegar uterine dilator Ø10,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.30 | Hegar uterine dilator Ø10,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.31 | Hegar uterine dilator Ø11,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.32 | Hegar uterine dilator Ø11,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.33 | Hegar uterine dilator Ø12,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.34 | Hegar uterine dilator Ø12,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.35 | Hegar uterine dilator Ø13,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.36 | Hegar uterine dilator Ø13,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.37 | Hegar uterine dilator Ø14,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.38 | Hegar uterine dilator Ø14,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.39 | Hegar uterine dilator Ø15,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.40 | Hegar uterine dilator Ø15,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.41 | Hegar uterine dilator Ø16,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.42 | Hegar uterine dilator Ø16,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.43 | Hegar uterine dilator Ø17,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.44 | Hegar uterine dilator Ø17,5 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.45 | Hegar uterine dilator Ø18,0 mm, very conical point 185mm, 7 1/4'' total length |  | 1
| 35.46 | "Recamier" uterine curette fig.0, sharp, rigid, 6 mm wide 300 mm, 12" total length |  | 1
| 35.47 | "Recamier" uterine curette fig.2, sharp, rigid, 8 mm wide 300 mm, 12" total length |  | 1
| 35.48 | "Recamier" uterine curette fig.4, sharp, rigid, 11 mm wide 300 mm, 12" total length |  | 1
| 35.49 | "Recamier" uterine curette fig.6, sharp, rigid, 14 mm wide 300 mm, 12" total length |  | 1
| 35.50 | "Recamier" uterine curette fig.2, blunt, rigid, 8 mm wide 300 mm, 12" total length |  | 1
| 35.51 | "Recamier" uterine curette fig.4, blunt, rigid, 11 mm wide 300 mm, 12" total length |  | 1
| 35.52 | "Recamier" uterine curette fig.6, blunt, rigid, 14 mm wide 300 mm, 12" total length |  | 1
| 35.53 | "Recamier" uterine curette fig.8, blunt, rigid, 16 mm wide 300 mm, 12" total length |  | 1
| 35.54 | "Winter" placenta-/ovum forceps fig. 1, curved 290 mm, 11 1/2" total length |  | 1
| 35.55 | "Winter" placenta-/ovum forceps fig. 2, curved 290 mm, 11 1/2" total length |  | 1
| 35.56 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 35.57 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 35.58 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 35.59 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
36 |  | GYNECOLOGY SET(ABDOMINAL HYSTERECTOMY) |  | 1 |  | 0
| 36.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 36.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 4
| 36.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 4
| 36.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 2
| 36.5 | Scalpel handle No.3, str. ,mirrow finsh. for blades fig. 10-15, 40, 42 210 mm, 8 1/4" total length |  | 1
| 36.6 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 36.7 | Titanit "Metzenbaum-Fino" diss.scissors curved 230 mm 9 1/8" total length |  | 1
| 36.8 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 36.9 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 36.10 | Power TC episiotomy scissors curved 200 mm, 8" total length |  | 1
| 36.11 | "Titanit" Sims uterus scissors bl./bl., curved, 200 mm 8" total length |  | 1
| 36.12 | "Titanit" Wertheim dissecting scissors curved, 230 mm 9 1/8" total length |  | 1
| 36.13 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 36.14 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 36.15 | Surgical forceps, standard str., tip 3,0 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 36.16 | Surgical forceps, standard str., tip 3,4 mm, 1x2 teeth, flat handle 250 mm, 10" total length |  | 2
| 36.17 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 36.18 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 36.19 | "DeBakey" atr. forceps, straight safety-pin with press-stopper, tip 2,0 m 240 mm, 9 1/2" total length |  | 2
| 36.20 | "Rochester-Ochsner" haemostatic forceps str., tip 2,4mm,1x2 teeth, serrated jaw 140 mm, 5 1/2" total length |  | 6
| 36.21 | "Pean" haemostatic forceps straight, tip 2,5 mm, serrated jaw 140 mm, 5 1/2" total length |  | 6
| 36.22 | "Kocher-Ochsner" haemostatic forceps str., tip 3,4mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 4
| 36.23 | "Kocher-Ochsner" haemostatic forceps str.,tip 3,8mm,1x2 teeth, serrated jaw 240 mm, 9 1/2" total length |  | 2
| 36.24 | "Wertheim" hysterectomy forceps curv.,tip 4,4 mm,1x2 teeth, serrated jaw 230 mm, 9 1/8" total length |  | 4
| 36.25 | "Mikulicz" peritoneum forceps curv.,tip 2,7 mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 6
| 36.26 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 4
| 36.27 | "Schröder" uterine tenaculum forceps sharp 250 mm, 10" total length |  | 2
| 36.28 | "Schröder" uterine tenaculum forceps 3x3 teeth, straight, 11 mm wide 270 mm, 10 3/4" total length |  | 1
| 36.29 | "Foerster" sponge-/dressing forceps straight, tip 13,0 mm, surrated jaw 245 mm, 9 3/4" total length |  | 1
| 36.30 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,3 mm, serrated jaw 220 mm, 8 3/4" total length |  | 2
| 36.31 | "Overholt-Geissendörfer" dissecting for. curved, tip 2,5 mm, serrated jaw 230 mm, 9 1/8" total length |  | 2
| 36.32 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 36.33 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 1
| 36.34 | Power TC needle holder a. Hegar strong model 205 mm 8 1/8" total length |  | 1
| 36.35 | Power TC needle holder a. Hegar strong model 240 mm 9 1/2" total length |  | 1
| 36.36 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 36.37 | "Volkmann" retractor 4 teeth, 8,5x19 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 36.38 | "Kocher" retractor blunt, blade 80x60 mm 250 mm, 10" total length |  | 1
| 36.39 | "Tübinger model" vaginal retractor blunt, 1 blade 95x20 mm 265 mm, 10 1/2" total length |  | 1
| 36.40 | "Fritsch" abdominal retractor fig. 4, blunt, blade 45x75 mm 250 mm, 10" total length |  | 2
| 36.41 | "Mikulicz" abdominal retractor blunt, blade 120x50 mm 250 mm, 10" total length |  | 2
| 36.42 | "Haberer" abdominal/intestinal spatula width of endings 40 mm a. 50 mm, malleab 280 mm, 11 1/8" total length |  | 1
| 36.43 | Double ended probe tip Ø 2,0 mm 300 mm, 12" total length |  | 1
| 36.44 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 36.45 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 36.46 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 36.47 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 36.48 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
37 |  | CAESARIAN SECTION SET |  | 2 |  | 0
| 37.1 | "Maier" sponge-/dressing forceps with lock, curved 265 mm, 10 1/2" total length |  | 2
| 37.2 | "Backhaus" towel clamp curved, tip 1,5 mm, sharp 110 mm, 4 1/4" total length |  | 6
| 37.3 | Towel clamp for papertowel buttoned tip 115 mm, 4 1/2" total length |  | 6
| 37.4 | Scalpel handle No.4, short for blades fig. 18-36 135 mm, 5 1/4" total length |  | 2
| 37.5 | "Titanit" Metzenbaum scissors curved, 200 mm 8" total length |  | 1
| 37.6 | Titanit "Metzenbaum-Fino" diss.scissors curved 230 mm 9 1/8" total length |  | 1
| 37.7 | Power TC ligature scissors, curved, wavecut, 180 mm 7 1/8" total length |  | 1
| 37.8 | Titanit scissors a. Lexer, curved, 165 mm, 6 1/2" total length |  | 1
| 37.9 | "Titanit" Sims uterus scissors bl./bl., curved, 200 mm 8" total length |  | 1
| 37.10 | Surgical scissors sh./bl., str., tip 5,0 mm 145 mm, 5 3/4" total length |  | 1
| 37.11 | Surgical forceps str., tip 1,25mm, 1x2 teeth, flat handle 145 mm, 5 3/4" total length |  | 2
| 37.12 | Surgical forceps, standard str., tip 3,0 mm, 1x2 teeth, flat handle 200 mm, 8" total length |  | 2
| 37.13 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 155 mm, 6 1/8" total length |  | 2
| 37.14 | "DeBakey" atr. forceps, straight Safety-Pin, tip 2,0 mm 200 mm, 8" total length |  | 2
| 37.15 | "Rochester-Ochsner" haemostatic forceps str., tip 2,4mm,1x2 teeth, serrated jaw 140 mm, 5 1/2" total length |  | 6
| 37.16 | "Pean" haemostatic forceps straight, tip 2,5 mm, serrated jaw 140 mm, 5 1/2" total length |  | 6
| 37.17 | "Kocher-Ochsner" haemostatic forceps str., tip 3,4mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 6
| 37.18 | "Rochester-Pean" haemostatic forceps straight, tip 3,2 mm, serrated jaw 200 mm, 8" total length |  | 6
| 37.19 | "Maier" sponge-/dressing forceps with lock, straight 265 mm, 10 1/2" total length |  | 4
| 37.20 | "Mikulicz" peritoneum forceps curv.,tip 2,7 mm,1x2 teeth, serrated jaw 200 mm, 8" total length |  | 6
| 37.21 | "Faure" peritoneum forceps curv.,tip 3,5 mm,1x2 teeth, serrated jaw 210 mm, 8 1/8" total length |  | 2
| 37.22 | "Heywood-Smith" hemorhoidal forceps surrated jaw 215 mm, 8 1/2" total length |  | 2
| 37.23 | "Recamier" uterine curette fig. 15, blunt, rigid, tip 35 mm 360 mm, 14 1/4" total length |  | 1
| 37.24 | "Crile-Wood" power TC needle holder straight, tip 1,8 mm, needle size 4-6 mm 150 mm, 6" total length |  | 1
| 37.25 | "Hegar-Mayo" power TC needle holder straight, tip 2,8 mm, needle size 1-4 mm 205 mm, 8 1/8" total length |  | 1
| 37.26 | Power TC needle holder a. Hegar strong model 205 mm 8 1/8" total length |  | 1
| 37.27 | Power TC needle holder a. Hegar strong model 240 mm 9 1/2" total length |  | 1
| 37.28 | "Roux" retractor, double ended 1 set/3 pieces (WH0250/WH0260/WH0270) 165 mm, 6 1/2" total length |  | 1
| 37.29 | "Volkmann" retractor 6 teeth, 8,5x29 mm, semi-sharp 220 mm, 8 3/4" total length |  | 2
| 37.30 | "Haberer" abdominal/intestinal spatula width of endings 40 mm a. 50 mm, malleab 280 mm, 11 1/8" total length |  | 1
| 37.31 | "Fritsch" abdominal retractor fig. 5, blunt, blade 64x85 mm 250 mm, 10" total length |  | 2
| 37.32 | Double ended probe tip Ø 2,0 mm 300 mm, 12" total length |  | 1
| 37.33 | Laboratory bowl, heavy-weight model B4, capacity 0,16 L, Ø 85 mm 40 mm height |  | 1
| 37.34 | Laboratory bowl, heavy-weight model B6, capacity 0,4 L, Ø 110 mm 55 mm height |  | 1
| 37.35 | Kidney tray stainless steel capacity 0,50 L, height 39 mm 250 mm, 10" total length |  | 1
| 37.36 | Cont. 1/1 bottom non-perforated 600x272x138mm |  | 1
| 37.37 | Cont. 1/1 flex lid system blue, lid pps biostop system 600x285 mm |  | 1
| 37.38 | Wire basket 1/1, fall handle, 480x250x70mm |  | 1
38 |  | SUPPLEMENT SET FOR VAGINAL HYSTERECTOMY |  | 1 |  | 0
| 38.1 | "Doyen" vaginal specula, figure 2 blade 90x60 mm 240 mm, 9 1/2" total length |  | 1
| 38.2 | "Doyen" vaginal specula, figure 3 blade 120x60 mm 240 mm, 9 1/2" total length |  | 1
| 38.3 | Vaginal specula "Breisky", 130X30 mm, 320 mm, 12 3/4" total length |  | 1
| 38.4 | Breisky vaginal specula 130X40 mm 320 mm, 12 3/4" total length |  | 1
| 38.5 | "DeBakey" atr. forceps, str. Safety-Pin with Press-Stopper,tip 3,5mm 300 mm, 12" total length |  | 2
| 38.6 | Surgical forceps, standard str., tip 4,0 mm, 1x2 teeth, flat handle 300 mm, 12" total length |  | 2
| 38.7 | Power TC needle holder a. Bozemann s-curved, tip 3,0 mm 300 mm 12" total length |  | 1
| 38.8 | "Masson" power TC needle holder straight, tip 2,8 mm, needle size 2-4 mm 265 mm, 10 1/2" total length |  | 1
| 38.9 | "Zenker" dissecting- and ligature forcep lightly curved 300 mm, 12" total length |  | 1
| 38.10 | "Zenker" dissecting- and ligature forcep strongly curved, 290 mm, 11 1/2" total length |  | 1
| 38.11 | Uterine tenaculum forceps Aesculap-Pratt 4x4 teeth, 260 mm, 10 1/4" total length width 18mm |  | 1
| 38.12 | "Doyen" ovum forceps straight, fenestrated 185 mm, 7 1/4" total length |  | 1
| 38.13 | "Heywood-Smith" polypus forceps straight 250 mm, 10" total length |  | 1
39 |  | Instruments set for Laparoscopy |  | 2 |  | 0
| 39.1 | Trocar sleeve, capacity 5.5 mm, length 100 mm |  | 3
| 39.2 | Trocar 5.5mm, pyramidal tip, length 100 mm |  | 3
| 39.3 | Trocar sleeve, capacity 10 mm, length 100 mm, with insufflation tap |  | 2
| 39.4 | Trocar 10 mm, pyramidal tip, length 100 mm |  | 2
| 39.5 | Trocar sleeve, capacity 5.5 mm, length 150 mm |  | 2
| 39.6 | Trocar 5.5 mm, pyramidal tip, length 150 mm |  | 2
| 39.7 | Trocar sleeve, capacity 10 mm, length 150 mm, with insufflation tap |  | 2
| 39.8 | Trocar 10 mm, pyramidal tip, length 150 mm |  | 2
| 39.9 | Reducing adapter 12.5 /10 mm, 12.5 / 5mm, 10 / 5 mm (pack of 5) |  | 1
| 39.10 | Monopolar hook scissors ", Ø 5 mm, working length 330 mm |  | 1
| 39.11 | Monopolar scissors "Metzenbaum", Ø 5 mm, working length 330 mm, curved |  | 1
| 39.12 | Monopolar grasping and dissecting forceps "Maryland", Ø 5 mm, length 330 mm |  | 1
| 39.13 | Universal grasping forceps, Ø 5 mm, length 330 mm, monopolar |  | 1
| 39.14 | Monopolar atraumatic grasping forceps, with openings, Ø 5mm, working length 330 mm |  | 2
| 39.15 | Atraumatic grasping forceps, Ø 5 mm, jaws with toothing, length of jaws 26 mm, working length 330 mm |  | 1
| 39.16 | Monopolar biopsy forceps with needles, Ø 5 mm, double jaw action, length 330 mm |  | 1
| 39.17 | Clip applicator, reusable |  | 1
| 39.18 | Suction and irrigation tube, Ø 5 mm, length ≈300mm |  | 1
| 39.19 | Suction and irrigation tube, Ø 5mm, length ≈400mm |  | 1
| 39.20 | Injection cannula with Luer connector, Ø 5 mm |  | 2
| 39.21 | Monopolar hook electrode, Ø 5 mm, working length ≈340 mm |  | 1
| 39.22 | Monopolar hook electrode, Ø 5 mm, working length ≈420 mm |  | 1
| 39.23 | Monopolar atraumatic grasping forceps, with openings Ø 5 mm, working length 450 mm |  | 1
| 39.24 | Monopolar grasping and dissecting forceps, with openings Ø 5 mm, single jaw action, working length 450 mm |  | 1
| 39.25 | HF Monopolar connecting cable |  | 2
| 39.26 | Bipolar scissors "Metzenbaum”, Ø 5 mm, working length 330 mm, curved |  | 1
| 39.27 | Bipolar dissecting forceps "Maryland", Ø 5 mm, length 330 mm |  | 1
| 39.28 | Insert of bipolar grasping forceps, fenestrated, Ø 5, length 330 mm |  | 1
| 39.29 | HF Bipolar connecting cable |  | 1
| 39.30 | Grasping forceps for extraction, single jaw action, 2/3 teeth, Ø 10 mm, length 330 mm |  | 1
| 39.31 | Modular needle holder with carbide insert, straight, Ø 5 mm, length 330 mm, axial handle |  | 1
| 39.32 | Insert of modular needle holder with carbide insert, curved left, Ø 5 mm, length 330 mm |  | 1
| 39.33 | Insert of modular needle holder with function of automatic needle alignment, Ø 5 mm, length 330 mm |  | 1
| 39.34 | Veres cannula, length 120 mm |  | 1
| 39.35 | Veres cannula, length 150 mm |  | 1
| 39.36 | Box for instruments sterilization with wire basket, za sterilization, keeping and transport of instruments |  | 1
| 39.37 | Spare sealing caps for trocars 5,5 mm (pack od 10) |  | 1
| 39.38 | Spare sealing caps for trocars 10 mm (pack od 10) |  | 1
| 39.39 | Box for sterilization of telescope |  | 1
40 |  | Instruments set for Laparoscopy-colorectal Surgery |  | 1 |  | 0
| 40.1 | Trocar sleeve, capacity 5.5 mm, length 100 mm |  | 2
| 40.2 | Trocar 5.5mm, pyramidal tip, length 100 mm |  | 2
| 40.3 | Trocar sleeve, capacity 10 mm, length 100 mm, with insufflation tap |  | 3
| 40.4 | Trocar 10 mm, pyramidal tip, length 100 mm |  | 3
| 40.5 | Trocar sleeve, capacity 12,5 mm, length 100 mm, with insufflation tap |  | 1
| 40.6 | Trocar 12,5 mm, pyramidal tip, length 100mm |  | 1
| 40.7 | Reducing adapter 12.5 /10 mm, 12.5 / 5mm, 10 / 5 mm (pack of 5) |  | 1
| 40.8 | Monopolar scissors "Metzenbaum", Ø 5 mm, working length 330 mm, curved |  | 1
| 40.9 | Atraumatic grasping forceps , Ø 5 mm, jaws with fine shaped fenestrated teeth, length of jaws 26 mm, working length 330 mm |  | 1
| 40.10 | Atraumatic grasper forceps "Babcock", Ø 10 mm, length 330 mm |  | 1
| 40.11 | Atraumatic intestinal forceps, Ø 5 mm, length 330 mm |  | 1
| 40.12 | Grasper forceps" Babcock", fenestrated, single jaws action, Ø 5 mm, length 330mm |  | 1
| 40.13 | Grasping and dissecting forceps "Maryland", Ø 10 mm, length 330 mm |  | 1
| 40.14 | Modular needle holder with carbide insert, straight, Ø 5 mm, length 330 mm, axial handle |  | 1
| 40.15 | Modular needle holder with carbide insert, curved left, Ø 5 mm, length 330 mm, axial handle |  | 1
| 40.16 | Stapler holder 30° anvil grasper, curved, double action, Ø 10 mm, length 380 mm |  | 1
| 40.17 | Monopolar atraumatic grasping forceps, with openings Ø 5 mm, working length 330 mm |  | 1
| 40.18 | Bipolar grasper forceps, with openings, Ø 5mm, length 330 mm |  | 1
| 40.19 | HF Bipolar connecting cable |  | 1
| 40.20 | Suction and irrigation tube, Ø 5 mm, length ≈300mm |  | 1
| 40.21 | Suction and irrigation tube, Ø 5 mm, length ≈400mm |  | 1
| 40.22 | HF Monopolar connecting cable |  | 1
| 40.23 | Clip applicator, reusable |  | 1
| 40.24 | Monopolar hook electrode, Ø 5 mm, working length ≈340 mm |  | 1
| 40.25 | Monopolar hook electrode, Ø 5 mm, working length ≈420 mm |  | 1
| 40.26 | Veres cannula, length 120 mm |  | 1
| 40.27 | Box for instruments sterilization with wire basket, za sterilization, keeping and transport of instruments |  | 1
| 40.28 | Spare sealing caps for trocars 5,5 mm (pack od 10) |  | 1
| 40.29 | Spare sealing caps for trocars 10 mm (pack od 10) |  | 1
| 40.30 | Box for sterilization of telescope |  | 1
41 |  | Instruments set for Laparoscopy in Gynecology |  | 1 |  | 0
| 41.1 | Trocar sleeve, capacity 5.5 mm, length 100 mm |  | 3
| 41.2 | Trocar 5.5mm, pyramidal tip, length 100 mm |  | 3
| 41.3 | Trocar sleeve, capacity 10 mm, length 100 mm, with insufflation tap |  | 2
| 41.4 | Trocar 10 mm, pyramidal tip, length 100 mm |  | 2
| 41.5 | Reducing adapter 12.5 /10 mm, 12.5 / 5mm, 10 / 5 mm (pack of 5) |  | 1
| 41.6 | Monopolar scissors "Metzenbaum", Ø 5 mm, working length 330 mm, curved |  | 2
| 41.7 | Atraumatic grasping forceps , Ø 5 mm, jaws with fine shaped fenestrated teeth, length of jaws 26 mm, working length 330 mm |  | 2
| 41.8 | Grasper forceps" Babcock", fenestrated, single jaws action, Ø 5 mm, length 330mm |  | 1
| 41.9 | Monopolar grasping and dissecting forceps "Maryland", Ø 5 mm, length 330 mm |  | 1
| 41.10 | Monopolar atraumatic grasping forceps, Ø 5mm, working length 330 mm |  | 3
| 41.11 | Grasping forceps for extraction, single jaw action, 2/3 teeth, Ø 10 mm, length 330 mm |  | 1
| 41.12 | Grasping forceps for extraction, single jaw action, 2/3 teeth, Ø 5 mm, length 330 mm |  | 1
| 41.13 | Modular needle holder with carbide insert, straight, Ø 5 mm, length 330 mm, axial handle |  | 1
| 41.14 | Insert of modular needle holder with carbide insert, curved left, Ø 5 mm, length 330 mm |  | 1
| 41.15 | Insert of modular needle holder with function of automatic needle alignment, Ø 5 mm, length 330 mm |  | 1
| 41.16 | Bipolar scissors "Metzenbaum", Ø 5 mm, working length 330 mm, curved |  | 1
| 41.17 | Bipolar grasper forceps, with openings, Ø 5mm, length 330 mm |  | 1
| 41.18 | HF Bipolar connecting cable |  | 2
| 41.19 | Suction and irrigation tube, Ø 5 mm, length ≈300mm |  | 1
| 41.20 | HF Monopolar connecting cable |  | 2
| 41.21 | Veres cannula, length 120 mm |  | 2
| 41.22 | Box for instruments sterilization with wire basket, za sterilization, keeping and transport of instruments |  | 1
| 41.23 | Spare sealing caps for trocars 5,5 mm (pack od 10) |  | 1
| 41.24 | Spare sealing caps for trocars 10 mm (pack od 10) |  | 1
| 41.25 | Box for sterilization of telescope |  | 1
42 |  | FESS instrument set |  | 1 |  | 0
| 42.1 | Straight Forward Telescope 0°, enlarged view, diameter 4 mm, length 18 cm, autoclavable, fiberoptic light transmission incorporated, with plastic containter |  | 1
| 42.2 | Telescope Handle, flat, standard model, length 11 cm, for use with Staright Forward Telescopes 0° with diameter 4 mm and length 18 cm |  | 1
| 42.3 | Suction and Irrigation Sheath 0°, outer diameter 4.8 x 6 mm, working length 14 cm, for use with Telescopes 0° and lens irrigation system |  | 1
| 42.4 | Forward-Oblique Telescope 30°, enlarged view, diameter 4 mm, length 18 cm, autoclavable, fiber optic light transmission incorporated, with plastic containter |  | 1
| 42.5 | Telescope Handle, round, standard model, length 11 cm, for use with Telescope 30°- 120° with diameter 4 mm and length 18 cm |  | 1
| 42.6 | Suction and Irrigation Sheath 30°, outer diameter 4.8 x 6 mm, working length 14 cm, for use with Telescopes 30° and lens irrigation system |  | 1
| 42.7 | Lateral Telescope 70°, enlarged view, diameter 4 mm, length 18 cm, autoclavable, Fiber optic lighttransmission incorporated |  | 1
| 42.8 | Telescope Handle, round, standard model, length 11 cm, for use with Telescope 30°- 120° with diameter 4 mm and length 18 cm |  | 1
| 42.9 | Suction and Irrigation Sheath 70°, outer diameter 4.8 x 6 mm, working length 14 cm, for use with Telescopes 70° and lens irrigation system |  | 1
| 42.10 | Sickle Knife, pointed, length 19 cm |  | 1
| 42.11 | Elevator, double-ended, semisharp and blunt, length 20 cm |  | 1
| 42.12 | BLAKESLEY Nasal Forceps, straight, size 1, with cleaning connector, working length 13 cm |  | 1
| 42.13 | BLAKESLEY Nasal Forceps, 45° upturned, size 1, with cleaning connector, working length 13 cm |  | 1
| 42.14 | BLAKESLEY Nasal Forceps, 90° upturned, size 1, with cleaning connector, working length 13 cm |  | 1
| 42.15 | Bone Punch, detachable, rigid, 90° upbiting, not through- cutting, size 2 mm, working length 17 cm |  | 1
| 42.16 | Antrum Punch, backward cutting, sheath 360° rotatable, with fixing screw, working length 10 cm, Take-apart sheath |  | 1
| 42.17 | Punch, circular cutting, 65° upturned, for frontal sinus recess, diameter 3.5 mm, with cleaning connector, working length 17 cm, including Cleaning Tool |  | 1
| 42.18 | Bipolar Coagulating Forceps, insulated, angular, blunt, for nose, with integrated suction channel for epistaxis, with cut-off hole, length 19 cm, for use with bipolar high frequency cord |  | 1
| 42.19 | Bipolar High Frequency Cord, length 300 cm |  | 1
| 42.20 | Probe, double-ended, maxillary sinus ostium seeker, ball-shaped ends diameter 1.2 and 2 mm, length 19 cm |  | 1
| 42.21 | Frontal ostium seeker curved, double ended |  | 1
| 42.22 | Antrum Curette, oblong, small size, length 19 cm |  | 1
| 42.23 | Frontal Sinus Curette, 55° curved, oval, forward cutting, length 19 cm |  | 1
| 42.24 | Antrum Grasping Forceps, jaws curved downwards, fixed jaw curved 90°, movable jaw backward opening 120°, with cleaning connector, working length 10 cm |  | 1
| 42.25 | Suction Tube, with mandrel and cut-off hole, with distance marking at 5 – 9 cm, 9 Fr., working length 10 cm |  | 1
| 42.26 | Antrum Cannula, LUER- Lock, long curved, outer diameter 3 mm, length 12.5 cm |  | 1
| 42.27 | Antrum Cannula, LUER- Lock, long curved, outer diameter 4 mm, length 12.5 cm |  | 1
| 42.28 | Bipolar Suction Forceps, 15° upturned, with suction channel, for bipolar coagulation in paranasal areas, working length 12.5 cm, for use with Bipolar High Frequency Cords |  | 1
| 42.29 | Bipolar Suction Forceps, 45° upturned, with suction channel, for bipolar coagulation in paranasal areas, working length 12.5 cm, for use with Bipolar High Frequency Cords |  | 1
| 42.30 | Fiber Optic Light Cable, with straight connector, diameter 3.5 mm, length 230 cm |  | 1
| 42.31 | Plastic Container, for Sterilizing and Storage perforated, with transparent lid, with silicone mat, for two-level storage |  | 1
43 |  | Instrument set for Tonsillectomy |  | 1 |  | 0
| 43.1 | Tongue Blade, with central groove, size 1, 30 x 70 mm |  | 1
| 43.2 | Tongue Blade, with central groove, size 2, 34 x 80 mm |  | 1
| 43.3 | Tongue Blade, with central groove, size 3, 39 x 84 mm |  | 1
| 43.4 | Tongue Blade, with central groove, size 4, 42 x 98 mm |  | 1
| 43.5 | Tongue Depressor, length 17 cm |  | 1
| 43.6 | Nasal Cutting Forceps, dismantling, round, size 3, working length 11 cm |  | 1
| 43.7 | Tongue Depressor, length 19 cm |  | 1
| 43.8 | Needle, Tonsil, LUER-lock, straight, long |  | 1
| 43.9 | Septum Cannula, angular, LUER- Lock |  | 1
| 43.10 | BACKHAUS Towel Forceps, length11 cm |  | 6
| 43.11 | Mouth Gag, standard model, with sliding Tooth Hook 744501, for use with Tongue Blades |  | 1
| 43.12 | Mouth Gag, with removable bow for replacement of rubber tube, for use with Tongue Blades |  | 1
| 43.13 | Tonsil Snare, length 29 cm |  | 1
| 43.14 | Wire Snare Loop, stainless, package of 100, for tonsil snare, 749000, 753500 and 753700 |  | 1
| 43.15 | Forceps, length 20 cm |  | 1
| 43.16 | Forceps, length 19 cm |  | 1
| 43.17 | Scissors, length 19.5 cm |  | 1
| 43.18 | Suction Tube, standard model, with large holes, length 29 cm |  | 1
| 43.19 | Elevator, width 11.5 mm, length 23.5 cm, (standard size) |  | 1
| 43.20 | Forceps, slightly curved, serrated, length 19 cm |  | 1
| 43.21 | Forceps, slender, 1 x 2 teeth, length 18 cm |  | 1
| 43.22 | Forceps, serrated, width 4 mm, length 20 cm |  | 1
| 43.23 | Tissue Forceps, 2 x 3 teeth, width 4.5 mm, length 20 cm |  | 1
| 43.24 | ROEDER Loop Applicator, Model E length 27 cm |  | 1
| 43.25 | Loops, Polyglycolic asid, absorbable, 12 sterile packed, size1, fro ROEDER Applikator 757700 |  | 1
| 43.26 | BECKMANN Adenoid Curette, straight, size 0, length 22 cm |  | 1
| 43.27 | BECKMANN Adenoid Curette, straight, size 1, length 22 cm |  | 1
| 43.28 | BECKMANN Adenoid Curette, straight, size 2, length 22 cm |  | 1
| 43.29 | BECKMANN Adenoid Curette, straight, size 3, length 22 cm |  | 1
| 43.30 | BECKMANN Adenoid Curette, straight, size 4, length 22 cm |  | 1
| 43.31 | NAGER Pillar Retractor, length 21 cm |  | 1
| 43.32 | JURACZ Adenoid Forceps, sharp, length 19 cm |  | 1
| 43.33 | HEGAR Needle Holder, tungsten carbide inserts, slender, length 18 cm |  | 1
| 43.34 | Bipolar Forceps, insulated, tips curved, blunt, tips 2 mm, length 19 cm |  | 1
| 43.35 | Plastic Container, for Sterilizing and Storage perforated, with transparent lid, with silicone mat, for two-level storage, external dimensions (w x d x h): 520 mm x 250 mm x 145 mm |  | 1
44 |  | Diagnostic Instrument set for ENT |  | 2 |  | 0
| 44.1 | Ear Speculum, standard model, outer diameter 3 mm, length 3.6 cm |  | 2
| 44.2 | Ear Speculum, standard model, outer diameter 4 mm, length 3.6 cm |  | 5
| 44.3 | Ear Speculum, standard model, outer diameter 5 mm, length 3.6 cm |  | 10
| 44.4 | Ear Speculum, standard model, outer diameter 6 mm, length 3.6 cm |  | 10
| 44.5 | Ear Speculum, standard model, outer diameter 7.5 mm, length 3. 6 cm |  | 2
| 44.6 | Tuning Fork, A1-440, with stand |  | 2
| 44.7 | Cotton Applicator, standard model, triangular serrated, diameter 1.1 mm, working length 13 cm |  | 10
| 44.8 | Cotton Applicator, standard model, triangular, serrated, diameter 1.3 mm, length 15 cm |  | 5
| 44.9 | Curved Spatula, length 14. 5 cm |  | 1
| 44.10 | Ear Hook, without ball end, size 1, length 15.5 cm |  | 5
| 44.11 | Ear Hook, with ball end, size 2, length 15.5 cm |  | 2
| 44.12 | Ear Curette, blunt, size 0, length 14.5 cm |  | 2
| 44.13 | Ear Curette, blunt, size 1, length 14.5 cm |  | 2
| 44.14 | Ear Knife, screw handle, bayonet- shaped, working length 8 cm |  | 1
| 44.15 | Ear Forceps, bayonet-shaped, length 14.5 cm |  | 10
| 44.16 | Nasal Dressing Forceps, bayonet-shaped, length 16.5 cm |  | 5
| 44.17 | Ear Dressing Forceps, slender, extra delicate, working length 6 cm |  | 1
| 44.18 | Ear Forceps, detachable, extra delicate, serrated, 1 x 4.5 mm, working length 8 cm |  | 1
| 44.19 | Ear Forceps, dismantling, serrated, working length 8 cm |  | 1
| 44.20 | Ear Forceps, dismantling, working length 8 cm |  | 1
| 44.21 | Suction Handle, with cut- off hole, LUER cone, length 5.5 cm, for use with Suction Tubes |  | 2
| 44.22 | Suction Tube, straight, LUER-Lock, malleable, outer diameter 1.5 mm, length 10 cm |  | 3
| 44.23 | Suction Tube, straight, LUER-Lock, malleable, outer diameter 2 mm, length 10 cm |  | 6
| 44.24 | Suction Tube, straight, LUER-Lock, malleable, outer diameter 2.5 mm, length 10 cm |  | 3
| 44.25 | Dressing Forceps, standard width, length 14.5 cm |  | 1
| 44.26 | Tissue Forceps, 1 x 2 teeth, standard width, length 14.5 cm |  | 1
| 44.27 | Needle Holder, tungsten carbide inserts, length 15 cm |  | 1
| 44.28 | Scissors, straight, sharp/blunt, length 14.5 cm |  | 1
| 44.29 | Nasal Speculum, standard model, blade length 75 mm, length 13.5 cm |  | 1
| 44.30 | Nasal Speculum, for adults, length 13 cm |  | 20
| 44.31 | Nasal Speculum, for children, length 13 cm |  | 5
| 44.32 | Suction Tube, conical, LUER, outer diameter 3 mm, working length 11 cm |  | 2
| 44.33 | Suction Tube, conical, LUER, outer diameter 3.5 mm, working length 11 cm |  | 2
| 44.34 | Suction Tube, conical, LUER, outer diameter 4.5 mm, working length 13 cm |  | 1
| 44.35 | Nasal Dressing Forceps, light model, working length 11 cm |  | 1
| 44.36 | LUER-Lock Tube Connector, male, tube diameter 6 mm |  | 1
| 44.37 | Antrum Cannula, LUER- Lock, long curved, outer diameter 2.5 mm, length 12.5 cm |  | 2
| 44.38 | Scissors, curved, length 14 cm |  | 1
| 44.39 | Surgical Handle, Fig. 3, length 12.5 cm, for Blades 208010–15, 208210–15 |  | 1
| 44.40 | Tongue Depressor, length 19 cm |  | 20
| 44.41 | Forceps, serrated, width 4 mm, length 20 cm |  | 1
| 44.42 | Postnasal Applicator, length 22 cm |  | 1
| 44.43 | Laryngeal Applicator, length 27 cm |  | 1
| 44.44 | Plastic Container, for Sterilizing and Storage perforated, with transparent lid, with silicone mat, for two-level storage, external dimensions (w x d x h): 520 mm x 250 mm x 145 mm |  | 2
45 |  | Basic set for Ear Micro Surgery |  | 1 |  | 0
| 45.1 | Nasal Speculum, for adults, length 13 cm |  | 1
| 45.2 | Raspatory, width 8 mm, length 18 cm |  | 1
| 45.3 | Raspatory, width 3 mm, length 19 cm |  | 1
| 45.4 | Tissue Forceps, 1 x 2 teeth, length 11.5 cm |  | 1
| 45.5 | Forceps, serrated, length 15 cm |  | 1
| 45.6 | Scissors, curved, sharp/blunt, length 12 cm |  | 1
| 45.7 | Retractor, 2 x 2 prongs, length 11 cm |  | 2
| 45.8 | Retractor, 2 prongs on right, full blade on left, length 11 cm |  | 1
| 45.9 | Retractor, 2 prongs on left, full blade on right, length 11 cm |  | 1
| 45.10 | Retractor, 4 prongs, sharp, length 13 cm |  | 1
| 45.11 | Scissors, curved, sharp/blunt, length 14.5 cm |  | 1
| 45.12 | Artery Forceps, straight, length 14 cm |  | 1
| 45.13 | Needle Holder, tungsten carbide inserts, length 13 cm |  | 1
| 45.14 | “Mosquito” Artery Forceps, straight, length 12.5 cm |  | 1
| 45.15 | “Mosquito” Artery Forceps, curved, length 12.5 cm |  | 2
| 45.16 | Backhaus Towel Forceps, length 9 cm |  | 6
| 45.17 | Wire Cutter, serrated, length 10 cm |  | 1
| 45.18 | Instrument Guard, teflon, autoclavable, for micro forceps and micro scissors with working length 8 cm, length 11.5 cm |  | 6
| 45.19 | Ear Speculum, standard model, outer diameter 5 mm, length 3.6 cm |  | 1
| 45.20 | Ear Speculum, standard model, outer diameter 6 mm, length 3.6 cm |  | 1
| 45.21 | Ear Speculum, standard model, outer diameter 7.5 mm, length 3. 6 cm |  | 1
| 45.22 | Ear Forceps, detachable, extra delicate, serrated, 1 x 4.5 mm, working length 8 cm |  | 1
| 45.23 | Ear Forceps, detachable, extra delicate, pointed, serrated, 1 x 4.5 mm, working length 8 cm |  | 1
| 45.24 | Ear Forceps, dismantling, serrated, working length 8 cm |  | 1
| 45.25 | Ear Forceps, detachable, extra delicate, oval cupped jaws, 0.9 mm, working length 8 cm |  | 1
| 45.26 | Ear Forceps, detachable, round cupped jaws, diameter 2 mm, working length 8 cm |  | 1
| 45.27 | Scissors, detachable, delicate standard model, blade length 8 mm, working length 8 cm |  | 1
| 45.28 | Malleus Nipper, detachable, upbiting, working length 8 cm |  | 1
| 45.29 | Wire Crimper, detachable, straight, working length 8 cm |  | 1
| 45.30 | Suction Tube, angular, LUER-Lock, outer diameter 0.7 mm, working length 6 cm |  | 4
| 45.31 | Suction Tube, angular, LUER-Lock, outer diameter 1 mm, working length 6 cm |  | 4
| 45.32 | Suction Tube, angular, LUER-Lock, outer diameter 1.3 mm, working length 6 cm |  | 4
| 45.33 | Suction Tube, angular, LUER-Lock, outer diameter 1.5 mm, working length 6 cm |  | 4
| 45.34 | Suction Tube, angular, LUER-Lock, outer diameter 2 mm, working length 6 cm |  | 4
| 45.35 | Suction Handle, with cut-off hole, LUER cone, length 10 cm, for use with Suction Tubes 204005 - 204025, 204305 - 204330 |  | 1
| 45.36 | Suction Tube, with grip plate, cut-off hole and stylet, LUER, 5 Fr., length 18 cm |  | 1
| 45.37 | Suction Tube, with grip plate, cut-off hole and stylet, LUER, 7 Fr., length 18 cm |  | 1
| 45.38 | Suction Tube, with grip plate, cut-off hole and stylet, LUER, 9 Fr., length 18 cm |  | 1
| 45.39 | Ear Hook, without ball end, size 1, length 15.5 cm |  | 1
| 45.40 | Seeker, with ball end, angled 45°, size 3, length 15.5 cm |  | 1
| 45.41 | Ear Curette, sharp, size 00, length 14.5 cm |  | 1
| 45.42 | Buck Ear Curette, sharp, size 0, length 14.5 cm |  | 1
| 45.43 | Hook, retrograde, 2 mm, length 15.5 cm |  | 1
| 45.44 | Curette, large, spoon sizes 2 x 3.2 mm and 1.6 x 2.6 mm, length 15 cm |  | 2
| 45.45 | Curette, small, spoon sizes 1 x 1.6 mm and 1.3 x 2 mm, length 15 cm |  | 2
| 45.46 | Hook, reinforced, retrospective, 0.8 mm, length 16 cm |  | 1
| 45.47 | Pick, 45°, size 0.5 mm, length 16 cm |  | 1
| 45.48 | Knife. Footplate, 16 cm |  | 1
| 45.49 | Footplate Hook, 0.2 mm, length 16 cm |  | 2
| 45.50 | Footplate Hook, 0.6 mm, length 16 cm |  | 2
| 45.51 | Manual Perforator, diameter 0.3 mm, length 16 cm |  | 1
| 45.52 | Guiding Fork, length 15.5 cm |  | 1
| 45.53 | Pick, 45°, size 1 mm, length 16 cm |  | 1
| 45.54 | Round Knife, vertical, standard size 3.5 x 2.5 mm, length 16 cm |  | 2
| 45.55 | Round Knife, 45°, diameter 2 mm, length 16 cm |  | 2
| 45.56 | Sickle Knife, double- edged, standard model, slightly curved, length 16 cm |  | 2
| 45.57 | Needle, slight curve, length 16.5 cm |  | 2
| 45.58 | Pick, 90°, size 0.4 mm, length 16 cm |  | 1
| 45.59 | Pick, 90°, size 0.6 mm, length 16 cm |  | 1
| 45.60 | Pick, 90°, size 0.8 mm, length 16 cm |  | 1
| 45.61 | Pick, 90°, size 1 mm, length 16 cm |  | 1
| 45.62 | Bipolar Forceps, insulated, tips curved, blunt, tips 1 mm, length 16 cm |  | 1
46 |  | Instrument sets Laryngomicroscopy (LMS) |  | 1 |  | 0
| 46.1 | Operating Laryngoscope, for adults, large, length 17 cm
| 46.2 | Operating Laryngoscope, for adults, medium, length 17 cm (most commonly used model)
| 46.3 | Operating Laryngoscope, for adults, small, length 18 cm (for particularly difficult anatomical circumstances)
| 46.4 | Fiber Optic Light Carrier, for distal illumination, working length 14 cm, for use with Laryngoscopes
| 46.5 | Laryngoscope Holder and Chest Support consisting of: Laryngoscope Holder, with adjustment wheel, Support Rod, movable, with metal ring, diameter 9 cm, length 34 cm
| 46.6 | Teeth Protector, silicone, autoclavable, can also be used as inlay in metal teeth protector |  | 5
| 46.7 | Needle Holder, delicate, straight, serrated jaws, size 1.8 x 3.5 mm, sheath conically reinforced from distal to proximal end, with ratchet, with cleaning connector, working length 23 cm
| 46.8 | Laryngeal Biopsy Forceps, heavy model with oval jaws, size 3 x 4 mm, single action jaws, with cleaning connector, working length 23 cm
| 46.9 | Scissors, straight, with cleaning connector, working length 23 cm
| 46.10 | Grasping Forceps, without ratchet, serrated, straight, with cleaning connector, straight, working length 23 cm
| 46.11 | Scissors, angled 45°, with cleaning connector, working length 23 cm
| 46.12 | Scissors, curved to right, with cleaning connector, working length 23 cm |  | 1
| 46.13 | Scissors, curved to left, with cleaning connector, working length 23 cm |  | 1
| 46.14 | Forceps, straight, with 2 mm cupped jaws, with cleaning connector, working length 23 cm |  | 1
| 46.15 | Forceps, curved upwards, with 2 mm cupped jaws, with cleaning connector, working length 23 cm |  | 1
| 46.16 | Forceps, curved to right, with 2 mm cupped jaws, with cleaning connector, working length 23 cm |  | 1
| 46.17 | Forceps, curved to left, with 2 mm cupped jaws, with cleaning connector, working length 23 cm |  | 1
| 46.18 | Handle |  | 1
| 46.19 | Hook, blunt, with probe end, working length 23 cm |  | 1
| 46.20 | Hook, sharp, with probe end, working length 23 cm |  | 1
| 46.21 | Knot Tier, working length 23 cm |  | 1
| 46.22 | Ligature Needle, curved to left, working length 23 cm |  | 1
| 46.23 | Ligature Needle, curved to right, working length 23 cm |  | 1
| 46.24 | Hook, blunt, angled 90º, with probe end, length 5 mm, working length 23 cm |  | 1
| 46.25 | Ligature Needle, curved to right, working length 23 cm |  | 1
| 46.26 | Knife, oval, straight, working length 23 cm |  | 1
| 46.27 | Suction Raspatory, working length 23 cm |  | 1
| 46.28 | Suction Tube, outer diameter 2.5 mm, working length 23 cm |  | 1
| 46.29 | Injection Needle, LUER-Lock, straight, working length 23 cm |  | 1
| 46.30 | Suction and Coagulation Cannula, outer diameter 3 mm, working length 26 cm, for use with unipolar High Frequency Cords |  | 1
| 46.31 | Unipolar High Frequency Cord, with 5 mm plug, length 300 cm |  | 1
| 46.32 | Plastic Container, for Sterilizing and Storage perforated, with transparent lid, with silicone mat, for two-level storage, external dimensions (w x d x h): 520 mm x 250 mm x 145 mm |  | 1
47 |  | Oftalmology instruments for cataract injurries, shutters and vitreoretinal |  | 1 |  | 0
| 47.1 | WESTCOTT SCISSORS, blunt tips, curved, standard blades, overall length 120 mm |  | 2
| 47.2 | SAUTTER, CORNEAL SCISSORS, light curve, fine blunt tips,10 mm blades, 14 mm from tip to pivot, overall length 100 mm |  | 2
| 47.3 | LIEBERMANN-TENNANT, EYE SPECULUM, adjustable, fenestrated for temporal placement, 15 mm open blades,overall length 80 mm |  | 2
| 47.4 | KERSHNER, EYE SPECULUM, adjustable, reversible, for nasal and temporal placement, 15 mm fenestrated blades, overall length 75 mm |  | 2
| 47.5 | CAPSULORHEXIS FORCEPS, strong, with heel stop to protect the iris, interlocking micro serrated tips, 12 mm straight shafts, overall length 103 mm |  | 2
| 47.6 | TUEBINGEN MODEL, CAPSULOTOMY SCISSORS, straight, sharp tips, 9 mm blades, 1.8 mm pivot width, overall length 88 mm |  | 2
| 47.7 | TUEBINGEN MODEL, CAPSULOTOMY SCISSORS, curved, sharp tips, 9 mm blades, 1.8 mm pivot width, overall length 88 mm |  | 2
| 47.8 | BARRAQUER, IRIS SCISSORS, 7 mm blades, blunt tips, overall length 56 mm |  | 2
| 47.9 | THOMAS NEUHANN, NUCLEUS DIVIDING HOOK, for vertical chopping techniques, 0.5 mm wide fine tip, overall length 108 mm |  | 2
| 47.10 | SNELLEN, LENS LOOP, smooth, curved, 3.5 mm wide, overall length 135 mm |  | 2
| 47.11 | BRAUWEILER, ASPIRATION CANNULA HANDPIECE, male / male connectors, 80 mm |  | 2
| 47.12 | BRAUWEILER, ASPIRATION HANDPIECE, angled at 13 mm, 0.3 mm beveled port, textured tip, 21 gauge / 0.8 mm |  | 2
| 47.13 | BRAUWEILER, IRRIGATION CANNULA HANDPIECE, male / female connectors, 78 mm |  | 2
| 47.14 | BRAUWEILER, IRRIGATION HANDPIECE, angled at 12 mm, front opening, 21 gauge / 0.8 mm |  | 2
| 47.15 | BRAUWEILER, ASPIRATION CANNULA HANDPIECE, male / male connectors, 80 mm |  | 2
| 47.16 | SHEPARD IMPLANTATION FORCEPS, curved, 1.4 mm disc-shaped tips, delicate lens holding shafts, overall length 120 mm |  | 2
| 47.17 | SINSKEY, LENS MANIPULATION HOOK, for 0.2 mm IOL hole, 10 mm angled shaft, overall length 103 mm, round handle |  | 2
| 47.18 | BREMEN MODEL, LENS MANIPULATOR, vertical Y-hook, 0.15 mm diameter tip, overall length 116 mm |  | 2
| 47.19 | LESTER, LENS LOOP MANIPULATOR, vaulted, 0.25 mm diameter tip, overall length 102 mm |  | 2
| 47.20 | DARDENNE, IRIS HOOK AND LENS MANIPULATOR, push-pull tip, overall length 105 mm |  | 2
| 47.21 | BINKHORST SYNECHIAE SPATULA, 10 mm long, semi-sharp, extra delicate, overall length 129 mm |  | 2
| 47.22 | RENTSCH SPATULA, 12 mm long spatula, tapered from 0.6 mm to 0.3 mm, overall length 118 mm |  | 2
| 47.23 | EYE SCISSORS, delicate, sharp tips, 31 mm from tip to pivot, overall length 105 mm, straight |  | 2
| 47.24 | CASTROVIEJO SUTURING FORCEPS, new model, 1 x 2 teeth, 6 mm tying platform, overall length 109 mm, 0.9 mm |  | 2
| 47.25 | BARRAQUER NEEDLE HOLDER, 10 mm delicate jaws, 0.8 x 0.55 mm tip, overall length 120 mm, curved, without lock |  | 2
| 47.26 | TÜBINGEN MODEL TYING FORCEPS, straight, smooth jaws, 6 mm tying platform, overall length 101 mm |  | 2
| 47.27 | ROHRSCHNEIDER NEEDLE HOLDER, straight, 13 mm heavy jaws, 1.3 x 1.5 mm tips, overall length 100 mm |  | 2
| 47.28 | SCALPEL HANDLE |  | 2
| 47.29 | DESMARRES CHALAZION FORCEPS, oval solid and fenestrated plates, 20 mm, overall length 88 mm |  | 2
| 47.30 | DESMARRES LID RETRACTOR, 14 mm wide, overall length 135 mm |  | 2
| 47.31 | GOLDSTEIN LACRIMAL SAC RETRACTOR, 3 x 3 prongs, adjustable, overall length 30 mm |  | 2
| 47.32 | HEBRA, CHALAZION CURETTE, small, 2.5 mm x 4.0 mm oval scoop, overall length 133 mm |  | 2
| 47.33 | HEBRA CHALAZION, medium, 3 mm x 5 mm oval scoop, overall length 133 mm |  | 2
| 47.34 | EYE SCISSORS, curved, sharp tips, overall length: tip to pivot: 90 mm 25 mm |  | 2
| 47.35 | ELSCHNIG SUPERIOR RECTUS FORCEPS, standard, straight, 1 x 2 teeth, overall length 110 mm |  | 2
| 47.36 | TISSUE FORCEPS, 1 x 2 teeth, overall length 100 mm, 1.0 mm |  | 2
| 47.37 | HALSTEAD MOSQUITO FORCEPS, serrated, 25 mm jaws, overall length 125 mm, straight |  | 2
| 47.38 | BARRAQUER CILIA FORCEPS, lightly angled, 6.5 mm platform, overall length 109 mm |  | 2
| 47.39 | BARRAQUER FIXATION FORCEPS, 3 sets of 1 x 2 teeth, 11 mm fixation platform, overall length 100 mm |  | 2
| 47.40 | BERKE PTOSIS FORCEPS, 22 mm long serrated tips, slide lock mechanism, overall length 110 mm |  | 2
| 47.41 | KUHNT CORNEAL SCARIFIER, rounded tip, overall length 123 mm, fig. 1, 5.5 mm x 2.5 mm blade |  | 2
| 47.42 | CASTROVIEJO NEEDLE HOLDER, 10.5 mm delicate jaws, 1.5 x 0.8 mm tips, overall length 140 mm, straight, without lock |  | 2
| 47.43 | SINGLE-USE IRIS RETRACTOR, for mechanical dilation of the pupil, flexible, 4 iris retractors per set, 2 sets per box, sterile |  | 8
| 47.44 | SINGLE-USE BACKFLUSH CANNULA blunt tip,5 pcs. per box, sterile, 23 gauge / 0.6 mm |  | 2
| 47.45 | VITREORETINAL FORCEPS, 20 gauge / 0.9 mm, 6 mm diameter handle, titanium, blue-colored, overall length 155 mm |  | 2
| 47.46 | VITREORETINAL FORCEPS, 23 gauge / 0.6 mm, 6 mm diameter handle, titanium, blue-colored, overall length 155 mm, straight endgripping, extra delicate 0.2 mm x 0.05 mm |  | 2
| 47.47 | VITREORETINAL FORCEPS, 23 gauge / 0.6 mm, 6 mm diameter handle, titanium, blue-colored, overall length 155 mm |  | 2
| 47.48 | VITREORETINAL FORCEPS, 23 gauge / 0.6 mm, 6 mm diameter handle, titanium, blue-colored, overall length 155 mm, straight, serrated “crocodile“ tips |  | 2
| 47.49 | FIXATION, HOLDING AND MEASURING INSTRUMENT, for trocar system, radius 3.5 mm |  | 2
48 |  | Operating table for general Surgery |  | 1 |  | 0
| 48.1 | Mobile operating table with electrohydraulic controls
| 48.2 | Operating table top subdivided into five sections : head rest, upper back plate, lower back plate, seat plate, leg plates with separations
| 48.3 | Height without padding 600 mm - 950 mm
| 48.4 | Trendelenburg : + 25°/ - 25°
| 48.5 | Lateral tilt : 15°
| 48.6 | Lower back plate : + 75° / - 40°
| 48.7 | Leg plate - 90° (down)
| 48.8 | Flex position 220°
| 48.9 | Reflex position 140°
| 48.10 | Longitudinal shift 300 mm
| 48.11 | Hand control with display with status notifications and current adjustment angles
| 48.12 | Stable base construction with four double swivel castors
| 48.13 | Base brake mechanism with foot pedal
| 48.14 | 0-position with one control
| 48.15 | Overall load minimum 350 kg
| 48.16 | Entire table top is designed without crossbars to permit radiography
| 48.17 | Batteries with capacity for one working week
| 48.18 | Electronic monitoring of the battery charge level with optical indicator and acoustic signal
| 48.19 | Antidecubital SFC padding 80mm
| 48.20 | Possibility to connect with OR integration
|  | Accessories for operating table: |  | 1
| 48.21 | Head rest |  | 1
| 48.22 | Leg rest, 2-section |  | 1
| 48.23 | Corded hand control |  | 1
| 48.24 | Arm support |  | 2
| 48.25 | Body strap |  | 1
| 48.26 | Anaesthesia screen |  | 1
| 48.27 | Extension back plate |  | 1
| 48.28 | Radial setting clamp |  | 1
49 |  | Operating table for colorectal surgery |  | 1 |  | 0
| 49.1 | Mobile operating table with electrohydraulic controls
| 49.2 | Operating table top subdivided into five sections : head rest, upper back plate, lower back plate, seat plate, leg plates with separations
| 49.3 | Height without padding 600 mm - 950 mm
| 49.4 | Trendelenburg : + 25°/ - 25°
| 49.5 | Lateral tilt : 15°
| 49.6 | Lower back plate : + 75° / - 40°
| 49.7 | Leg plate - 90° (down)
| 49.8 | Flex position 220°
| 49.9 | Reflex position 140°
| 49.10 | Longitudinal shift 300 mm
| 49.11 | Hand control with display with status notifications and current adjustment angles
| 49.12 | Stable base construction with four double swivel castors
| 49.13 | Base brake mechanism with foot pedal
| 49.14 | 0-position with one control
| 49.15 | Overall load minimum 350 kg
| 49.16 | Entire table top is designed without crossbars to permit radiography
| 49.17 | Batteries with capacity for one working week
| 49.18 | Electronic monitoring of the battery charge level with optical indicator and acoustic signal
| 49.19 | Antidecubital SFC padding 80mm
| 49.20 | Possibility to connect with OR integration
Accessories for operating table:
| 49.21 | Head rest |  | 1
| 49.22 | Leg rest, 2-section |  | 1
| 49.23 | Corded hand control |  | 1
| 49.24 | Arm support |  | 2
| 49.25 | Body strap |  | 1
| 49.26 | Anaesthesia screen |  | 1
| 49.27 | Extension back plate |  | 1
| 49.28 | Radial setting clamp |  | 1
| 49.29 | Leg holder with one-hand operation, pair |  | 1
50 |  | Operating table for ENT |  | 1 |  | 0
| 50.1 | Mobile operating table with electrohydraulic controls
| 50.2 | Operating table top subdivided into five sections : head rest, upper back plate, lower back plate, seat plate, leg plates with separations
| 50.3 | Height without padding 600 mm - 950 mm
| 50.4 | Trendelenburg : + 25°/ - 25°
| 50.5 | Lateral tilt : 15°
| 50.6 | Lower back plate : + 75° / - 40°
| 50.7 | Leg plate - 90° (down)
| 50.8 | Flex position 220°
| 50.9 | Reflex position 140°
| 50.10 | Longitudinal shift 300 mm
| 50.11 | Hand control with display with status notifications and current adjustment angles
| 50.12 | Stable base construction with four double swivel castors
| 50.13 | Base brake mechanism with foot pedal
| 50.14 | 0-position with one control
| 50.15 | Overall load minimum 350 kg
| 50.16 | Entire table top is designed without crossbars to permit radiography
| 50.17 | Batteries with capacity for one working week
| 50.18 | Electronic monitoring of the battery charge level with optical indicator and acoustic signal
| 50.19 | Antidecubital SFC padding 80mm
| 50.20 | Possibility to connect with OR integration
Accessories for operating table:
| 50.21 | Head rest |  | 1
| 50.22 | Leg rest, 2-section |  | 1
| 50.23 | Corded hand control |  | 1
| 50.24 | Arm support |  | 2
| 50.25 | Body strap |  | 1
| 50.26 | Anaesthesia screen |  | 1
| 50.27 | Extension back plate |  | 1
| 50.28 | Radial setting clamp |  | 1
| 50.29 | Connection bracket |  | 1
| 50.30 | Adjustable head rest adapter |  | 1
| 50.31 | Horseshoe head rest - two-section |  | 1
51 |  | Operating table for oflalmology |  | 1 |  | 0
| 51.1 | Mobile operating table with electrohydraulic controls
| 51.2 | Operating table top subdivided into five sections : head rest, upper back plate, lower back plate, seat plate, leg plates with separations
| 51.3 | Height without padding 600 mm - 950 mm
| 51.4 | Trendelenburg : + 25°/ - 25°
| 51.5 | Lateral tilt : 15°
| 51.6 | Lower back plate : + 75° / - 40°
| 51.7 | Leg plate - 90° (down)
| 51.8 | Flex position 220°
| 51.9 | Reflex position 140°
| 51.10 | Longitudinal shift 300 mm
| 51.11 | Hand control with display with status notifications and current adjustment angles
| 51.12 | Stable base construction with four double swivel castors
| 51.13 | Base brake mechanism with foot pedal
| 51.14 | 0-position with one control
| 51.15 | Overall load minimum 350 kg
| 51.16 | Entire table top is designed without crossbars to permit radiography
| 51.17 | Batteries with capacity for one working week
| 51.18 | Electronic monitoring of the battery charge level with optical indicator and acoustic signal
| 51.19 | Antidecubital SFC padding 80mm
| 51.20 | Possibility to connect with OR integration
Accessories for operating table:
| 51.21 | Head rest |  | 1
| 51.22 | Leg rest, 2-section |  | 1
| 51.23 | Corded hand control |  | 1
| 51.24 | Arm support |  | 2
| 51.25 | Body strap |  | 1
| 51.26 | Anaesthesia screen |  | 1
| 51.27 | Extension back plate |  | 1
| 51.28 | Radial setting clamp |  | 1
| 51.29 | Trapezoidal shaped head plate |  | 1
| 51.30 | Horseshoe head rest with one-handed adjustment |  | 1
52 |  | Operating table for Urology |  | 1 |  | 0
| 52.1 | Mobile operating table with electrohydraulic controls
| 52.2 | Operating table top subdivided into five sections : head rest, upper back plate, lower back plate, seat plate, leg plates with separations
| 52.3 | Height without padding 600 mm - 950 mm
| 52.4 | Trendelenburg : + 25°/ - 25°
| 52.5 | Lateral tilt : 15°
| 52.6 | Lower back plate : + 75° / - 40°
| 52.7 | Leg plate - 90° (down)
| 52.8 | Flex position 220°
| 52.9 | Reflex position 140°
| 52.10 | Longitudinal shift 300 mm
| 52.11 | Hand control with display with status notifications and current adjustment angles
| 52.12 | Stable base construction with four double swivel castors
| 52.13 | Base brake mechanism with foot pedal
| 52.14 | 0-position with one control
| 52.15 | Overall load minimum 350 kg
| 52.16 | Entire table top is designed without crossbars to permit radiography
| 52.17 | Batteries with capacity for one working week
| 52.18 | Electronic monitoring of the battery charge level with optical indicator and acoustic signal
| 52.19 | Antidecubital SFC padding 80mm
| 52.20 | Possibility to connect with OR integration
Accessories for operating table:
| 52.21 | Head rest - |  | 1
| 52.22 | Leg rest, 2-section |  | 1
| 52.23 | Corded hand control |  | 1
| 52.24 | Arm support |  | 2
| 52.25 | Body strap |  | 1
| 52.26 | Anaesthesia screen |  | 1
| 52.27 | Extension back plate |  | 1
| 52.28 | Extension seat plate |  | 1
| 52.29 | Radial setting clamp |  | 3
| 52.30 | Gopel knee holder |  | 2
| 52.31 | TUR rinsing set |  | 1
53 |  | Operating table for Gynecology |  | 1 |  | 0
| 53.1 | Mobile operating table with electrohydraulic controls
| 53.2 | Operating table top subdivided into five sections : head rest, upper back plate, lower back plate, seat plate, leg plates with separations
| 53.3 | Height without padding 600 mm - 950 mm
| 53.4 | Trendelenburg : + 25°/ - 25°
| 53.5 | Lateral tilt : 15°
| 53.6 | Lower back plate : + 75° / - 40°
| 53.7 | Leg plate - 90° (down)
| 53.8 | Flex position 220°
| 53.9 | Reflex position 140°
| 53.10 | Longitudinal shift 300 mm
| 53.11 | Hand control with display with status notifications and current adjustment angles
| 53.12 | Stable base construction with four double swivel castors
| 53.13 | Base brake mechanism with foot pedal
| 53.14 | 0-position with one control
| 53.15 | Overall load minimum 350 kg
| 53.16 | Entire table top is designed without crossbars to permit radiography
| 53.17 | Batteries with capacity for one working week
| 53.18 | Electronic monitoring of the battery charge level with optical indicator and acoustic signal
| 53.19 | Antidecubital SFC padding 80mm
| 53.20 | Possibility to connect with OR integration
Accessories for operating table:
| 53.21 | Head rest - |  | 1
| 53.22 | Leg rest, 2-section |  | 1
| 53.23 | Corded hand control |  | 1
| 53.24 | Arm support |  | 2
| 53.25 | Body strap |  | 1
| 53.26 | Anaesthesia screen |  | 1
| 53.27 | Extension back plate |  | 1
| 53.28 | Extension seat plate |  | 1
| 53.29 | Radial setting clamp |  | 3
| 53.30 | Gopel knee holder |  | 2
| 53.31 | TUR rinsing set |  | 1
54 |  | Operating table for orthopedy |  | 1 |  | 0
| 54.1 | Mobile operating table with electrohydraulic controls
| 54.2 | Operating table top subdivided into five sections : head rest, upper back plate, lower back plate, seat plate, leg plates with separations
| 54.3 | Height without padding 600 mm - 950 mm
| 54.4 | Trendelenburg : + 25°/ - 25°
| 54.5 | Lateral tilt : 15°
| 54.6 | Lower back plate : + 75° / - 40°
| 54.7 | Leg plate - 90° (down)
| 54.8 | Flex position 220°
| 54.9 | Reflex position 140°
| 54.10 | Longitudinal shift 3100 mm
| 54.11 | Hand control with display with status notifications and current adjustment angles
| 54.12 | Stable base construction with four double swivel castors
| 54.13 | Base brake mechanism with foot pedal
| 54.14 | 0-position with one control
| 54.15 | Overall load minimum 350 kg
| 54.16 | Entire table top is designed without crossbars to permit radiography
| 54.17 | Batteries with capacity for one working week
| 54.18 | Electronic monitoring of the battery charge level with optical indicator and acoustic signal
| 54.19 | Antidecubital SFC padding 80mm
| 54.20 | Possibility to connect with OR integration
Accessories for operating table:
| 54.21 | Head rest |  | 1
| 54.22 | Leg rest, 2-section |  | 1
| 54.23 | Corded hand control |  | 1
| 54.24 | Arm support |  | 2
| 54.25 | Body strap |  | 1
| 54.26 | Anaesthesia screen |  | 1
| 54.27 | Extension back plate |  | 1
| 54.28 | Radial setting clamp |  | 4
| 54.29 | Fixture for body support |  | 2
| 54.30 | Lateral support |  | 2
| 54.31 | Extension and traction device for orthopedy (1 trolley with bascet, 1 pair of adapters, 1 short telescopic bar, 1 long telescopic bar, 1 screw tension device, 1 foot plate mount, 2 support bars, 2 radial setting clamp, 1 countertraction post, 1 padded roll for countertraction post, 1 side rail extension) - 1 pcs. |  | 1
| 54.32 | Knee positioning device |  | 1
55 |  | CEILING LED OPERATING LIGHT WITH SATELITE |  | 6 |  | 0
| 55.1 | Main light and satelite on the same column
| 55.2 | Operating light , ceiling model, LED
| 55.3 | Illumination of the main light minimum 160.000 lx
| 55.4 | Illumination of the satelite minimum 160.000 lx
| 55.5 | Field diameter 20 - 25 cm
| 55.6 | Color rendering index : 95 Ra
| 55.7 | LED life time : minimum 60.000 hours
| 55.8 | Color Temperature (K) : 3.900 K
| 55.9 | Ambient light
| 55.10 | Dimming range : 10-100%
| 55.11 | Irradiance < 500 (W/m2)
| 55.12 | Depth of illumination at 60% : 50 cm
| 55.13 | Posibility to mounting Full HD wired and wireless cameras
| 55.14 | Automatic illumination Management (AIM) or similar
| 55.15 | Flux Stability Program (FSP) or similar
| 55.16 | Luminance Management Device (LMD) or similar
56 |  | MOBILE LED EXAMINATION LIGHT |  | 7 |  | 0
| 56.1 | Color temperature : 4,500 K
| 56.2 | Color rendering index (Ra) : 95
| 56.3 | Ilumination at 1 m : 40,000 lx
| 56.4 | Field diameter : 24 cm
| 56.5 | The lamp base is stable and with wheels with a brake
| 56.6 | An on / off switch is present on the lamp head
| 56.7 | Service life : min. 60.000 h