## 8 - Biosystems
Bidder: |  | Offered price:
LOT 8 - Biosystems
ID | Equipment name | QTY | Model, Manufacturer and Country of Origin | Total Price per line item
Biosystems
8.1 | Upright light microscope with microscopy digital camera and software | 1
8.2 | Advanced Stereomicroscope | 1
8.3 | Stereomicroscope | 1
8.4 | Antivibration table 1 | 1
8.5 | Antivibration table 2 | 3
8.6 | Damping plate | 1
8.7 | Small mammal traps | 100
8.8 | Microbiology incubator | 1
8.9 | Benchtop biosafety cabinet | 1
8.1 | Inverted (Fluo) microscope with phase contrast | 1
8.11 | Microplate reader with washer | 1
8.12 | Set of pipettors | 1
8.13 | Lab Waste Vacuum Aspiration System | 1
8.14 | Portable mini real-time PCR | 1
8.15 | PCR machine | 1
8.16 | Electrophoresis system | 1
Total Price per Lot

## 8.1-Upright light microscope
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.1 | UPRIGHT LIGHT MICROSCOPE WITH MICROSCOPY DIGITAL CAMERA AND SOFTWARE |  | 1 |  | 0
| 1 | MAIN BODY: • Field Diaphragm • Focusing stroke: min. 27mm • Refocusing function, by fixing the upper limit of the Z-focus • Minimum 3 build-in filters (e.g. NCB11, ND8, ND32) and possibility to add and/or exchange additional filters • Coarse and Fine Focus • Button for capturing images at the microscope, direct from the camera from same brand as the microscope • Possibility for upgrades: Brightfield, Phase Contrast, Nomarski DIC, Simple polarization, Fluorescence, Upgrade of motorized items at the microscope (Nosepiece, Stage, Fluorescence), Laser Confocal Microscopy
| 2 | ILLUMINATION: • Halogen 12V/100W • Built-in lens for optimal uniform illumination in the viewfiled periphery has to be provided at any magnification
| 3 | LEFT HAND ROTABLE CERAMIC STAGE: • Centerable, rotation angle 202° • Cross travel: min. 78(x) x 54(Y) • Holder for 1 specimen • Handle height and the torque of X-axis and Y-axis must be adjustable
| 4 | ERGONOMIC PHOTO TUBE: • Inclination angle: min. 10-30° at same time tube extension up to 40mm • Simple upgrade-possibility to add a microscopy digital camera • Beam split: 100/0 50/50 • Adapter for digital camera with 0.6x magnification
| 5 | EYE PIECES: • 10x22mm Field of View, with diopter adjustment for both eyepieces (2 pieces)
| 6 | NOSEPIECE: • with place for 6 objectives and usable for DIC
| 8 | SET OF OBJECTIVES: • Plan Fluor 4x or 5x with min. N.A 0.13 and min. W.D 17.1mm • Plan Fluor 10x with min. N.A 0.30 and min. W.D 16.0 mm • Plan Fluor 20x with min. N.A 0.50 and min. W.D 2.1mm, spring-loaded • Plan Apochromat 40x dry with min. N.A 0.95, min. W.D 0.21 mm, Cover glass correction from minimum 0.11-0.23 mm, Spring-loaded • Plan Apochromat 60x or 63x dry with min. N.A 0.95, min. W.D 0.15mm, Cover glass correction from minimum 0.11-0.23 mm, Spring-loaded • Plan Fluor 100x for using with immersion oil with min. N.A 1.30 and min. W.D 0.16mm, Spring-loaded and stopper
| 9 | ACCESSORIES: • Vinyl Cover for Microscope • 5pcs. Halogen Bulbs 12V/100W • Immersion oil • Instruction Manual
| 10 | MICROSCOPY DIGITAL CAMERA AND SOFTWARE: • at least 5.9 megapixel • CMOS image sensor, Diagonale min. 8.48mm, Sensor size min. 6.91x4.92mm • Capture of high-resolution images of up to 2880 x 2048 pixels • Size of sensor 1/1.8 inch" • Image Analysis software with possibility for stitching images in Z axis to get a full image for these different Z-levels, following features must be as minimum included: Time lapse, AVI Acquisition, Custom Aquisition, Annotation, ND Viewer, AVI creation, Snapshot creation for publication, Image Arithmetic, Binary Layers, Volume View with AVI creation, Volume View: 3D ND Crop, Volume Rendering & Movie Maker, Tiling (Montage) View, Max / Min Projections, Manual Channel Alignment, Large Image Stitching, Scan Large Image and Z, Grab Large Image Free Shape, Macro, Auto-Measurent, Custom Measurement Feature, Object Catalog, Live Object Inspection Tool, Region of Interest Tools & Statistics, Segmentation, Pixel Classifier, Report Generator, Vector Layer, Multi-Dimensional File Format, Multi-User Environment
| 11 | WORK STATION: • Personal Computer work station (OS: Windows 7 or 10; CPU: 3GHz or higher; RAM 32GB minimum, HDD: 4TB minimum; VIDEO: Microsoft's Direct X v11 or higher; GPU processing: minimum Radeon 390x series or nVidia 980 series cards, keyboard, mouse) • 20’’ TFT Monitor
| 12 | Installation and training included
| 13 | Warranty: 24 months
| 14 | Availability of the spare parts for 7 years
| 15 | Delivery time: 8 weeks
| 16 | Response time for failure notification – 24 (twenty-four) hours
| 17 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.2-Advanced Stereomicroscope
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.2 | Advanced Stereomicroscope |  | 1 |  | 0
| 1 | MAIN BODY: Zooming Body • Optical System: Parallel-optics type (zooming type) • Minimum zoom ratio: 12.7 : 1 • Zoom range: 0.63-8x • Total possible magnification: 3.15-540x • Possibility for intelligent functions for automatic recognizing objectives & zoom-magnification.
| 2 | FOCUSING: • Coarse and fine focusing
| 3 | NOSEPIECE FOR OBJECTIVES: • Nosepiece for two objectives. • In case of changing objectives the software must recognize automatically which objective is in the light path.
| 4 | OBJECTIVES: • Plan Apo objective with 1x magnification, Working Distance of minimum 70mm • Plan objective 1.5x magnification, Working Distance of minimum 44mm
| 5 | STAND (DIASCOPIC LIGHT): • LED Illumination with build-in lens for a contrast method similar to a relief contrast. • The stand must be a compact and slim-type based for enhances operation efficiency and ergonomy.
| 6 | TUBE & DRAWING TUBE: • Trinocular Tilting Tube with Beam Split 100:0 and 50:50 • Eye pieces with 10x and 22mm field of view and dioptre adjustment for both eyepieces • A drawing tube must be mounted
| 8 | EPISCOPIC ILLUMINATION: • Cold light with ring-light and swan neck light guide with two arms
| 9 | STAGE: • Stage for X and Y Minimum travel range: 150mm x 65mm
| 10 | MICROSCOPY DIGITAL CAMERA AND SOFTWARE: • Microscopy Digital Camera with at least 5.9 megapixel. • CMOS image sensor, Diagonal min. 8.48mm, Sensor size min. 6.91x4.92mm • Capture of high-resolution images of up to 2880 x 2048 pixels • Size of sensor 1/1.8 inch
| 11 | ANALYSIS SOFTWARE: Image Analysis software with possibility for stitching images in Z axis to get a full image for these different Z-levels, following features must be as minimum included: Time lapse, AVI Acquisition, Custom Aquisition, GUI Single Window, Annotation, ND Viewer, AVI creation, Snapshot creation for publication, Tiling (Montage) View, Manual Channel Alignment, Large Image Stitching, Grab Large Image Free Shape, Macro, Pixel Classifier, Report Generator, Vector Layer, Multi-Dimensional File Format, Multi-User Environment The software can run under Windows 7 or 10
| 12 | EXTERNAL STABLE STAND: • External stable stand for using for bigger specimens • Vertical Cross Travel: minimum 229mm • Horizontal Cross Travel: minimum 272mm
| 13 | ACCESSORIES: • Vinyl Cover for Microscope • Instruction Manual
| 14 | POSSIBLE CONTRAST METHODS OR UPGRADES: • With Transmitted light (Dia): - Polarization - Darkfield - Iris diaphragm • With Reflected light (Epi): - Fluorescence - Polarization - Darkfield - Focusing optic at the illumination for having a spot-light - Coaxial Illumination - LED Illumination observations e.g. with Ringlight - Teaching Head
| 15 | Installation and training included
| 16 | Warranty: 24 months
| 17 | Availability of the spare parts for 7 years
| 18 | Delivery time: 8 weeks
| 19 | Response time for failure notification – 24 (twenty-four) hours
| 20 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.3-Stereomicroscope
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.3 | Stereomicroscope |  | 1 |  | 0
| 1 | MAIN BODY: Zooming Body • Optical System: Parallel-optics type (zooming type) • Minimum zoom ratio: 7,5 : 1 • Zoom range: 0.67-5x • Total possible magnification: 3.35-300x • Tube: Fixed tube, Eyepiece inclination 45° with Interpupillary distance adjustment: 52-75mm
| 2 | OBJECTIVES: • Build-in objective with 1x magnification with Working Distance of minimum 115mm • Additional an objective 2x with Working Distance of 43,5mm
| 3 | STAND (DIASCOPIC & EPISCOPIC LIGHT): • LED Illumination, both episcopic and diascopic observatinos must be useable simultaneously (2 pieces)
| 4 | TUBE: • Trinocular Tilting Tube • Inclination angle: 45° • Interpupillary Distance: min. 52-75mm • Eye pieces with 10x and 22mm field of view and dioptre adjustment for both eyepieces
| 5 | MICROSCOPY DIGITAL CAMERA AND SOFTWARE: • Microscopy Digital Camera (CMOS) with at least 5 megapixel (min. 2592x1944 pixels). • Image analysis software for capturing, basic measurement and archiving images. • Operating system for Windows 7 or 10
| 6 | ACCESSORIES • Vinyl Cover for Microscope • Instruction Manual
| 7 | Installation and training included
| 8 | Warranty: 24 months
| 9 | Availability of the spare parts for 7 years
| 10 | Delivery time: 8 weeks
| 11 | Response time for failure notification – 24 (twenty-four) hours
| 12 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.4-Antivibration table 1
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.4 | Pneumatic Vibration Isolation Table size 980x660 mm |  | 1 |  | 0
| 1 | TOP PLATE SIZE 980X660X107 mm • with an internal structure of trussed rectangular crossbar structure (patented) made of welded carbon steel • covered by dull stainless steel surface with no magnetic capability and metric M6 threaded holes drilled on the top plate in 25 x 25 mm pattern • flatness error: +/- 0.25 mm per meter as maximum, along any straight line, anywhere on the top plate • horizontality error of the top plate on a working vibration isolation table: +/- 1.5 mm as maximum, measured between two opposite corners of the top plate
| 2 | LEG SYSTEM • vibration attenuators: 2nd-order, linear phase pneumatic filters built into the cylinders • maximum mass load of the legs: 4 kN, including the top plate • height of the leg system 82 cm (adjustable +/- 2 cm)
| 3 | QUIET AIR COMPRESSOR (gas supply) • Low (switch on) pressure: 2.0 Bar High (switch off) pressure: 2.5 Bar • Supply voltage: 230 VAC • Current consumption in working periods: 1.6 A • Peak current in the starting periods (approximately 2 seconds): 12 A • Noise level: 40 dB(A)
| 4 | Installation and training included
| 5 | Warranty: 3 years of full warranty on every part and accessory
| 6 | Availability of the spare parts for 7 years
| 7 | Delivery time: 8 weeks
| 8 | Response time for failure notification – 24 (twenty-four) hours
| 9 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.5-Antivibration table 2
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | QTY | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.5 | Pneumatic Vibration Isolation Table size 660x660 mm | 3 |  | 1 |  | 0
| 1 | TOP PLATE SIZE 660X660X107 mm • with an internal structure of trussed rectangular crossbar structure (patented) made of welded carbon steel • covered by dull stainless steel surface with no magnetic capability and metric M6 threaded holes drilled on the top plate in 25 x 25 mm pattern • flatness error: +/- 0.25 mm per meter as maximum, along any straight line, anywhere on the top plate • horizontality error of the top plate on a working vibration isolation table: +/- 1.5 mm as maximum, measured between two opposite corners of the top plate | 1
| 2 | LEG SYSTEM • vibration attenuators: 2nd-order, linear phase pneumatic filters built into the cylinders • maximum mass load of the legs: 4 kN, including the top plate • height of the leg system 82 cm (adjustable +/- 2 cm) | 1
| 3 | QUIET AIR COMPRESSOR (gas supply) • Low (switch on) pressure: 2.0 Bar High (switch off) pressure: 2.5 Bar • Supply voltage: 230 VAC • Current consumption in working periods: 1.6 A • Peak current in the starting periods (approximately 2 seconds): 12 A • Noise level: 40 dB(A) | 1
| 4 | Installation and training included
| 5 | Warranty: 3 years of full warranty on every part and accessory
| 6 | Availability of the spare parts for 7 years
| 7 | Delivery time: 8 weeks
| 8 | Response time for failure notification – 24 (twenty-four) hours
| 9 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.6-Damping Plate
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.6 | Elastomer-based Damping Plate |  | 1 |  | 0
| 1 | • 600x400 mm painted carbon steel top plate, magnetic • Height: 90 mm • Mass load: 300 N • vibrations absorbed by 4 pieces of elastomer (silicone rubber) cones • built-in spirit level and adjustable legs for leveling correction even if the load is not concentric
| 2 | Installation and training included
| 3 | Warranty: 3 years of full warranty on every part and accessory
| 4 | Availability of the spare parts for 7 years
| 5 | Delivery time: 8 weeks
| 6 | Response time for failure notification – 24 (twenty-four) hours
| 7 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.7-Small mammal traps
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | QTY | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.7 | Small mammal traps | 100 |  | 1 |  | 0
| 1 | • 2 doors (front and back side) • cage or transparent box possibility to see what is inside without opening it • The door drops automatically behind the small mammal when it crosses the treadle to reach the bait • Plentiful ventilation points to reduce stress to the trapped individual • The trap is easy to bait and set • full instructions are provided • Dimensions not bigger than 50 x 10 x 10 cm | 1
| 2 | Warranty: 1 year of full warranty | 1
| 3 | Delivery time: 8 weeks | 1

## 8.8-Microbiology incubator
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.8 | Microbiology incubator |  | 1 |  | 0
| 1 | Description: Intelligent temperature controller, based on microprocessor with PID parameters controlling function, and digital display make temperature accurate and reliable. The two-layer door makes the temperature fluctuation small. The toughened-glass indoor is convenient to observe in-incubator conditions without affecting in-incubator temperature. The working chamber use mirror stainless steel and the arc roundish, easy to clean.
| 2 | Mains voltage: 220V, 50Hz
| 3 | Temperature control: RT+5~65℃
| 4 | Temperature fluctuation: ± 0.5℃
| 5 | Consumption power: 200 W
| 6 | Power supply: 110/220 V±10%,50/60Hz
| 7 | Time range: 1-9999 min
| 8 | Chamber Dimensions (cm): 32*30 * 31.5
| 9 | Outer Dimension (cm): 61*47* 47
| 10 | Shelves per Chamber: 2
| 11 | Gross weight (kg): 45 kg
| 12 | Installation and training included
| 13 | Warranty: 12 months
| 14 | Availability of the spare parts for 10 years.
| 15 | Delivery time: 8 Weeks
| 16 | Response time for failure notification – 24 (twenty-four) hours
| 17 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.9-Benchtop biosafety cabinet
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.9 | Benchtop biosafety cabinet |  | 1 |  | 0
| 1 | Description: Three protection: operator, sample and environment. Airflow system: 70% air recirculation, 30% air exhaust. A2 Cabinet is suitable for working with microbiological research in the absence of volatile or toxic chemicals and radionuclide. Work zone is made of 304 stainless steel, is surround by negative pressure. Large Digital LCD display is easy to monitor all the safety parameters at a glance and ergonomically sized control panel improves user interface.
| 2 | External Size(W*D*H): 700*650*1230 mm
| 3 | Internal Size(W*D*H): 600*500*540 mm
| 4 | Tested Opening: Safety Height: ≤ 200 mm
| 5 | Max Opening: 370 mm
| 6 | Inflow Velocity: 0.53±0.025 m/s
| 7 | Down Flow Velocity: 0.33±0.025 m/s
| 8 | HEPA Filter: Two, 99.999% efficient at 0.3μm. Filter life indicator.
| 9 | Front Window: Motorized. Two-layer laminated toughened glass ≥ 5mm. Anti UV.
| 10 | Noise: EN12469 ≤ 58 dB / NSF49 ≤ 61 dB
| 11 | UV Lamp: 15W*1; UV timer, UV life indicator, emission of 253.7 nanometers for most efficient decontamination.
| 12 | Fluorescent Lamp: 40W*1
| 13 | Illumination: ≥1000Lux
| 14 | Consumption: 500W
| 15 | Socket: Two, total load of two sockets: 500W
| 16 | Display: LCD display: exhaust filter and downflow filter pressure, filter and UV lamp working time, inflow and downflow velocity, filter life, humidity and temperature, system working time etc.
| 17 | Control System: Microprocessor
| 18 | Airflow System: 70% air recirculation, 30% air exhaust
| 19 | Visual and Audio alarm: Filter replacement, window over height, abnormal air flow velocity.
| 20 | Material: Work Zone - 304 stainless steel; Main Body - Cold-rolled steel with anti-bacteria powder coating.
| 21 | Work Surface Height: 750mm with optional base stand
| 22 | Power Supply: 110/220V±10%, 60/50Hz
| 23 | Standard Accessory: Fluorescent lamp, UV lamp*2, Remote control, Waterproof sockets*2
| 24 | Installation and training included
| 25 | Warranty: 12 months
| 26 | Availability of the spare parts for 10 years.
| 27 | Delivery time: 8 Weeks
| 28 | Response time for failure notification – 24 (twenty-four) hours
| 29 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.10-Inverted (fluo) microscope
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.1 | Inverted (Fluo) microscope with phase contrast |  | 1 |  | 0
| 1 | Viewing Head: trinocular head inclined at 30°C, interpupillary distance 48-75mm
| 2 | Eyepiece: o High point extra wide field eyepiece EW10X/22 o Infinite plan objective - 4X/0.1, WD 18mm - 10X/0.25, WD 10mm - 20X/0.4, WD 5.1mm - 40X/0.6, WD 2.6mm o Infinite plan phase objective - PH10X/0.25, WD 10mm - PH20X/0.4, WD 10mm - PH40X/0.65, WD 0.35mm
| 3 | Nosepiece: Quintuple Nosepiece
| 4 | Condenser: ELWD Condenser NA0.3, LWD72mm(without condenser 150mm)
| 5 | Centering o Centering Telescope (Ø30mm) o 10X-20X,40X Pre-adjustable phase annulus plate o 10X-20X,40X Adjustable phase annulus plate o Plain stage 160*250mm o Glass insert o Attachable mechanical stage,X-Y coaxial control moving range 120*78mm o Auxiliary stage 70*180mm o Ø38mm Petri dish holder o Ø54mm Slide glass holder o Coaxial coarse and fine adjustment,vertical objective movement o Coarse stocke 37.7mm pet rotation,fine strok 0.2mm per rotation
| 6 | Illumination: Halogen lamp 6V/30W
| 7 | Filter o Diameter 45mm, Blue, Green, and Ground glass o Inverted Fluorescent Attachment o Video Adapter with C mount
| 8 | Installation and training included
| 9 | Warranty: 12 months
| 10 | Availability of the spare parts for 10 years.
| 11 | Delivery time: 8 Weeks
| 12 | Response time for failure notification – 24 (twenty-four) hours
| 13 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.11-Microplate reader
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.11 | Microplate reader |  | 1 |  | 0
| 1 | Reader: Elisa microplate reader, equipped with LCD touch screen display, able to monitor timely and print testing results directly. Application: It is widely used in the fields of hospital clinical diagnosis, immune pathology testing, microbial antigen and antibody detection, parasitic diseases diagnosis, blood disease, research. Industrial grade color LCD display, touch screen operation. Eight channel optical fiber measurement system imported detector. Center positioning function, accurate and reliable. Three kinds of linear vibration plate function. Unique open Cut- off judgment formula, think what you think. End point method, two-point method, dynamics, single/dual wavelength test mode. Configure the inhibition rat measurement module, dedicated to the field of food safety.
| 2 | Lamp: 12V/20W, Halogen tungsten lamp
| 3 | Optical Path: 8 channel vertical light path system
| 4 | Wavelength range: 400-800 nm
| 5 | Filter: default configuration 405,450,492,630 nm, up to 10 filters can be installed
| 6 | Reading range: 0~4.000Abs
| 7 | Resolution: 0.001Abs
| 8 | Accuracy: ≤±0.01Abs
| 9 | Stability: ≤±0.003Abs
| 10 | Repeatability: ≤0.3%
| 11 | Vibration plate: three kind of vibration plate function, adjustable 0~255 seconds
| 12 | Display: 8 inch color LCD screen, display the entire ELISA Plate board information, touch screen operation
| 13 | Software: professional software, can store 200 groups program, 100000 sample results, more than 10 kind of curve fitting equations
| 14 | Power supply: AC 110/220V±10%, 50/60Hz
| 15 | Package Size(W*D*H): 550*395*470 mm
| 16 | Gross Weight: 15.5 kg
| 17 | Washer: Elisa microplate washer, with 8 or 12 washing probes suitable with 96-well and 48-well microplates. It can be adapted to the 96- and 48-well Microplates. Regulates rinse position of cleaning head according to user requirements. Chooses to set washing parameters for different effects. External programming/downloading available through RS-232 ports. 8-channel or 12 channels optional.
| 18 | Microplate Types: 48-well, 96-well
| 19 | Washing: Automatic washing, 5 ways
| 20 | Washing the row number of microplate:1 row per time, 1- 12 rows
| 21 | Average Residue: 2ul per well
| 22 | Dispense Precision: ≤3% CV
| 23 | Shaking: 5 ways
| 24 | Memory: 100 programs
| 25 | Wash Cycles: at least 40000
| 26 | Soak Time: Programmable in seconds, up to 20 minutes.
| 27 | Anti-overflow: Excess fluid is automatically discharged.
| 28 | Installation and training included
| 29 | Warranty: 12 months
| 30 | Availability of the spare parts for 10 years.
| 31 | Delivery time: 8 Weeks
| 32 | Response time for failure notification – 24 (twenty-four) hours
| 33 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.12-Set of pipettors
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.12 | Set of pipettors |  | 1 |  | 0
| 1 | MicroPette Pipettor, 0.1-2.5μl (Precision 2.00/3.00/6.00; Increment 0,05 ul; Test Volume 2.5ul/1.25ul/0.25ul; Accuracy 2.50/3.00/12.00) |  | 2
| 2 | MicroPette Pipettor, 0.5-10 μl |  | 2
| 3 | MicroPette Pipettor, 2-20μl (Precision 0.40/1.00/2.00; Increment 0,5 ul; Test Volume 20ul/10ul/2ul; Accuracy 0.90/1.20/3.00) |  | 2
| 4 | MicroPette Pipettor, 10-100μl (Precision 0.15/0.40/1.50; Increment 1,0 ul; Test Volume 2.5ul/1.25ul/0.25ul; Accuracy 2.50/3.00/12.00) |  | 2
| 5 | MicroPette Pipettor, 100-1000μl (Precision 2.00/3.00/6.00; Increment 0,05 ul; Test Volume100ul/50ul/10ul; Accuracy 0.80/1.00/3.00) |  | 2
| 6 | MicroPette 12-Channel Pipettor, 5-50μl (Precision 0.50/1.00/2.00; Increment 0,5 ul; Test Volume 50ul/25ul/5ul; Accuracy 1.00/1.50/3.00) |  | 2
| 7 | MicroPette 12-Channel Pipettor, 50-300μl (Precision 0.25/0.50/0.80; Increment 5 ul; Test Volume 300ul/150ul/50ul; Accuracy 0.70/1.00/1.50) |  | 2
| 8 | Linear Pipettor Stand for 6 Pipettors |  | 3
| 9 | Pipette filler: Lightweight ergonomic design; Comfortable concave triggers with fine speed control; 8 hour continuous use per charge; For use with 1-100ml pipets; Can fill a 25ml pipette in 5 seconds; Gravity dispense, with blow-out for TD pipettes; Comes with 2 x 0.45um hydrophobic filters; Autoclavable nose-cone; Powerful lithium-ion battery, no 'memory'; Can be used while charging, no down-time; Comes with a charging base stand; |  | 2
| 10 | Installation and training included
| 11 | Warranty: 12 months
| 12 | Availability of the spare parts for 10 years.
| 13 | Delivery time: 8 Weeks
| 14 | Response time for failure notification – 24 (twenty-four) hours
| 15 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.13-Lab waste vacuum aspiratio
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.13 | Lab waste vacuum aspiration system |  | 1 |  | 0
| 1 | Description; o The Scilogex SafeVac is designed for laboratory waste recovery and quick separation of liquids and solids. It can be widely used in cell culture, DNA extraction, microplate waste and any other liquid separation or recovery. Waste is aspirated into the bottle by vacuum which can be easily adjusted by the vacuum control. o Integrated vacuum pump with quick & easy control o Quick connectors for easy installation o Level sensor detecting preventing liquid overflow o Autoclavable waste collection bottle & aspiration nozzles o Single channel nozzle (supplied) & 8-channel nozzle accessory o cTUVus approved
| 2 | Vacuum range: 0 to 600 mbar adjustable
| 3 | Pumping rate: 1 to 15 mL/S adjustable
| 4 | Bottle capacity: 4 liters
| 5 | Power: 20 W
| 6 | Dimensions: 9.6"L x 7.2"W x 18"H
| 7 | Weight: 7 lbs
| 8 | Voltage: 110/240V, 50/60Hz
| 9 | Installation and training included
| 10 | Warranty: 12 months
| 11 | Availability of the spare parts for 10 years.
| 12 | Delivery time: 8 Weeks
| 13 | Response time for failure notification – 24 (twenty-four) hours
| 14 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.14 Portable mini real-timePCR
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.14 | Portable mini real-time PCR |  | 1 |  | 0
| 1 | Description: OPEN platform with superior intra-run and inter-run analytical precision in DNA quantification and melting point analysis. Users can address biological phenomena with subtle effects on gene expression, or pathogen levels, for example discriminating 10% differences in transcript concentrations. For users performing HRM-based. analysis of genetic variants, the combination of thermal control, optical data quality and HRM data analysis system provide compelling functionality. The system can discriminate all classes of SNP, including Class 4 SNPs via HRM.
| 2 | Dimensions: D12cm x H16cm
| 3 | Weight: 1,8 kg or less
| 4 | Operating noise: Silent
| 5 | Electrical Voltage: 100-240 V (+-10%); Frequency:50-60 Hz; Power 90W
| 6 | Number of reactions: 16
| 7 | Reaction wheels: 0,1 ml tubes
| 8 | Reaction volume: 10-100 ul (20ul recommended)
| 9 | Temperature cycling system: Peltier based heating and cooling from +40°C to +99°C heating rate 3°C/s, cooling rate 1,5°C/s
| 10 | Temperature uniformity to +/- 0,5°C
| 11 | Temperature accuracy to +/- 0,25C
| 12 | Run time: less than 50 min
| 13 | Fluorescence excitation: Blue LED 500nm
| 14 | Fluorescence detection: CMOS
| 15 | Two calibrated dyes: SYBR Green I, ResoLight, FAM, VIC, HEX, CAL540, CAL560
| 16 | Multiplexing: Up to 2 targets
| 17 | Sensitivity: 1,1 fold discrimination, 9-log dynamic range, single copy detection
| 18 | Installation and training included
| 19 | Warranty: 12 months
| 20 | Availability of the spare parts for 10 years.
| 21 | Delivery time: 8 Weeks
| 22 | Response time for failure notification – 24 (twenty-four) hours
| 23 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.15-PCR machine
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.15 | PCR machine |  | 1 |  | 0
| 1 | Block Format: 96-well, 0.2 mL block
| 2 | Max block ramp rate: 4°C/sec for reaction volume of 1µL Max sample ramp rate: 3°C/sec for reaction volume of 1µL
| 3 | Temperature accuracy at least ±0.3°C
| 4 | Temperature range: 0-100°C
| 5 | Temperature uniformity <0.5°C after stabilizing at 95°C
| 6 | Temperature calibration: by standards tracable to National Institute of Standards and Technology
| 7 | PCR volume range: 10-100 µL
| 8 | USB and 2000 MB onboard memory; onboard capacity >1000 protocols
| 9 | 8'' color display
| 10 | Blocks with 3 temperature zones
| 11 | Dimensions not bigger than 25x25x50 cm
| 12 | Startup set - 96-Well Tray/Retainer Set Color: Blue For Use With (Equipment): ProFlex™ PCR System, SimpliAmp™ Thermal Cycler, Veriti® Dx Thermal Cycler, Veriti® Thermal Cycler Plate Compatibility: 96-Well Plates Product Line: MicroAmp® Product Size: 10 trays Shipping Condition: Room Temperature Type: Tube Tray
| 13 | Startup set - Splash-Free 96-Well Base Color: Black Plate Compatibility: 96-Well Plates Product Line: MicroAmp® Product Size: 10 bases Shipping Condition: Room Temperature Type: Tube Tray
| 14 | Installation and training included
| 15 | Warranty: 12 months
| 16 | Availability of the spare parts for 10 years.
| 17 | Delivery time: 8 Weeks
| 18 | Response time for failure notification – 24 (twenty-four) hours
| 19 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts

## 8.16-Electrophoresis system
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | QTY | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Biosystems
| 8.16 | Electrophoresis system | 1 |  | 1 |  | 0
| 1 | Type: Wide
| 2 | Sample Range: 25 to 500
| 3 | Volume (Metric) Buffer: 3l
| 4 | Length Gel: at least 25cm
| 5 | Width Gel: at least 20cm
| 6 | Gel tray allows use of up to 10 evenly spaced combs
| 7 | Using special combs and a multichannel pipetter, samples may be loaded directly from a 96, eight or 12-well plate.
| 8 | Sample loading with multichannel pipetter.
| 9 | Bubble level ensures casting of flat, uniformly thick gels.
| 10 | Height Exterior: not bigger than 15 cm
| 11 | Length Exterior: not bigger than 40 cm
| 12 | Width Exterior: not bigger than 40 cm
| 13 | Certifications/Compliance: CE marked
| 14 | Installation and training included
| 15 | Warranty: 12 months
| 16 | Availability of the spare parts for 10 years.
| 17 | Delivery time: 8 Weeks
| 18 | Response time for failure notification – 24 (twenty-four) hours
| 19 | Maximum period for repair and replacement from the moment of declaring - 7 (seven) days after receiving spare parts