## 10 Sensor Fabrication and Char
Bidder: |  | Offered price:
Sensor Fabrication and Characterisation
ID | Equipment name | QTY | Model, Manufacturer and Country of Origin | Total Price per line item
Tube furnace with gas divider system
10.1 | Tube furnace | 1
10.2 | Mass Flow Controllers | 1
10.3 | Valves and connections | 1
10.4 | Precision Dicing Saw | 1
10.5 | Environmental chamber | 1
10.6 | Centrifuge with rotor | 1
10.7 | Planetary ball mill | 1
10.8 | Rotary evaporator | 1
10.9 | Disperser | 1
10.1 | Overhead stirrer | 1
10.11 | Stirring Hot plate | 1
10.12 | SLA 3D printer | 1
10.13 | Desktop filament extruder | 1
10.14 | PCB prototyping | 1
10.15 | Probes for probe station and micropositioners | 1
Total Price per Lot

## 10.1 Tube furnace
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.1 | Tube furnace |  | 1 |  | 0
| 1 | Tube furnace with working temperature up to 1200 °C (maximum continuous temperature of 1100 °C) and heat up time of less than 75 min. Tube length for use in air atmosphere 800 mm length, with heated length 600 mm. Maximum power consumption below 4 kW. N type thermocouple. Programable heating profile, with 5 programs. Should be appropriate for additional work tubes up to 170 mm in diameter. |  | 1
| 2 | Digital protection of overheating |  | 1
| 3 | Additional worktube for controlled atmosphere application. The tube should be from aluminous porcelain, at least 1 m long, with outer diameter 150 mm. Should be applicable in vacuum up to 1200 °C. |  | 1
| 4 | Insulation plugs appropriate for tube furnace and additional work tube and temperature range. |  | 2
| 5 | End seals with gas inlet/outlet appropriate tube furnace and additional worktube. |  | 2
| 6 | Gas flowmeter for tube furnace, for flow control in range 2-25 l/min O2. With appropriate gas inlet diameter, compatible with end seals of additional worktube. |  | 1
| 7 | Appropriate software for control of furnace over the personal computer. Communication between computer and furnace should be made by RS232 port, supplied with appropriate cables. |  | 1
| 8 | Provided training by supplier
| 9 | Installation by supplier
| 10 | Warranty 2 years
| 11 | Spare parts provided by supplier
| 12 | Delivery time 3 months
| 13 | Response time for failure notification – 48 (fourty-eight) hours
| 14 | Maximum period for repair and replacement from the moment of declaring - 45 (fourty-five) days

## 10.2 Mass flow controllers
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.2 | Mass Flow Controllers |  | 1 |  | 0
| 1 | Mass flow and vacuum control system, as a support for simultaneous control of at least 6 mass flow controllers and vacuum. Wide range of pressure reading, from at least 10-9 to more than 2,5 x 106 Pascal. Gas flow control and monitoring using personal computer, connected to control system by RS232/RS485 serial port. Operating temperature up to at least 35 °C. Device should work in humid conditions up to 80%. Reading unit flow in sccm and slm. TFT LCD display. |  | 1
| 2 | Elastomer sealed digital mass flow controller, equipped with thermal flow sensors. Mass flow controller should be sealed with Viton seals. Proof pressure at least 60 bar. Analog connection output compatible with mass flow control system. 1/4 inch male VCR compatible fittings for gas pipeline system. Operating temperature range form room temperature up to at least 45 °C, in high relative humidity conditions. Flow rate readings full scale 10 sccm of N2. |  | 1
| 3 | Elastomer sealed digital mass flow controller, equipped with hot wire thermal flow sensors. Mass flow controller should be sealed with Viton seals. Proof pressure at least 60 bar. Analog connection output compatible with mass flow control system. 1/4 inch male VCR compatible fittings for gas pipeline system. Operating temperature range form room temperature up to at least 45 °C, in high relative humidity conditions. Flow rate readings full scale 100 sccm of N2. |  | 2
| 4 | Elastomer sealed digital mass flow controller, equipped with thermal flow sensors. Mass flow controller should be sealed with Viton seals. Proof pressure at least 60 bar. Analog connection output compatible with mass flow control system. 1/4 inch male VCR compatible fittings for gas pipeline system. Operating temperature range form room temperature up to at least 45 °C, in high relative humidity conditions. Flow rate readings full scale 500 sccm of N2. |  | 2
| 5 | Elastomer sealed digital mass flow controller, equipped with hot wire thermal flow sensors. Mass flow controller should be sealed with Viton seals. Proof pressure at least 60 bar. Analog connection output compatible with mass flow control system. 1/4 inch male VCR compatible fittings for gas pipeline system. Operating temperature range form room temperature up to at least 45 °C, in high relative humidity conditions. Flow rate readings full scale 5000 sccm of N2. |  | 1
| 6 | Supporting cables for connection between control system and mass flow controllers. |  | 6
| 7 | Warranty: 12 months
| 8 | Delivery time 3 months

## 10.3 Valves and connections
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.3 | Valves and connections |  | 1 |  | 0
| 1 | Manual ball valve, 1/4 inch tube fitting, with 4.7 mm orifice. Straight two way flow pattern. Valve body made from 316 stainless steel. Tested according to WS-22. |  | 19
| 2 | Union tee connection, 1/4 inch tube fitting, 316 stainless steel body material. |  | 6
| 3 | Union cross connection, 1/4 inch tube fitting, 316 stainless steel body material. |  | 4
| 4 | Delivery time 3 months

## 10.4 Precision Dicing Saw
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.4 | Precision Dicing Saw |  | 1 |  | 0
| 1 | Motor • Brushless magnetic DC motor (110V) with variable speed up to 3,000 rpm. • Power consumption up to 200 W
| 2 | Effective cutting range • X-axis: at least 200 mm • Y-axis: at least 100 mm • Z-axis: at least 100 mm
| 3 | Cutting speed range • X-axis: at least 1 to 50 mm/min adjustible • Y-axis: at least 1 to 50 mm/min adjustible • Z-axis: at least 1 to 20 mm/min adjustible
| 4 | Accuracy • Moving resolution: at least 2.5 µm • Position Accuracy: at least 10 µm
| 5 | Cutting blades kits • One 4" diameter, 0.35 mm thick fully sintered diamond blade • Two pairs of flanges with 62 mm diameter for dicing and 42 mm diameter for deep cutting
| 6 | Software and laptop • 14''-15'' laptop with PC software are provided with the equipment. • Cutting parameters can be set in software as cutting length, width, diameter, time, feed speed, kerf loss, etc • Manual or programmable control mode • Compatible with Win 7/8/8.1
| 7 | Water cooling • Assemblies of water jetting, draining, and splashing protection is included. • One coolant circulating tank is included. • Adjustable nozzles for re-circulating coolant included
| 8 | Dimensions • Length: not more than 600 mm • Width: not more than 600 mm • Height: not more than 700 mm
| 9 | Accesories • 4'' vacuum chuck with 6 mm push-to-connect fitting • Cross mount vise for cutiing sample with precision 90°angle • Controller • Mechanical chuck and T-bolts assemblies set wiht 1° resolution and angle rotation 0-360° • Water splash guard for covering 4'' blade • Aluminum plates, 2 pcs • Wax for gluing samples 80x60x12 mm, melting point 70°C • Resin bonded ceramics for samples holding 50x50x10 mm, 2 pcs • Pair of blade clamps 62 mm and 42 mm • 4'' steel plate for holding samples • Screw drivers • Water and drain hoses • Operational manual and software • Film for wafer bonding to vacuum chuck 6.5'' wide and 0.25 mm thick • Plastic splash-protection cover • Diamond blade dressing stone 20x20x150 mm • Oilless vacuum pump maximum air flow 120 l/min, power ≤ 450 W, CE approval • Coolant circulating tank with pump, 4 liters, with 1/2'' hose barb connector and 8 mm quick connector PU hose pipe
| 10 | CE certified
| 11 | Warranty: 12 months
| 12 | Delivery time 3 months

## 10.5 Environmental chamber
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.5 | Environmental chamber |  | 1 |  | 0
| 1.1 | Minimum inner chamber size (WxHxD): 500x600x500 mm
| 2 | Maximum External Chamber size (WxHxD): 1000*1600*1100 mm
| 3 | Temperature range: -70C ~ +150°C
| 4 | Humidity range: 10%~98%RH
| 5 | 2m fused silica capillary inlet, heated to 110 °C, coupled to sintered leak technology.
| 6 | Temperature and humidity uniformity :≤±1.5℃
| 7 | Heat up speed :≤5℃/min for heating
| 8 | Cool down speed: ≤ 1.2℃/min for cooling
| 9 | Warranty: 12 months
| 10 | Delivery time 3 months

## 10.6 Centrifuge with rotor
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.6 | Centrifuge with rotor |  | 1 |  | 0
| 1 | Bench top centrifuge for routine work in laboratory; sound volume with adequate rotor <55 db (A); with induction motor without brushes. Rotation speed from 200 up to 14000 rpm. Rotation speed self-adjustment according to thrust force (RCF). Min RCF for centrifuge with requested rotor should be 16g. Option of selecting the centrifuging time 1-99 minutes, with short-spin for shorter centrifuging. Possibility of defining and saving up to 35 user programs. Possibility of adjusting acceleration force and inhibition in maximum 10 steps. Electronical detection of disbalance and rotor recognition. Possibility of mechanical lid opening in the case of power failure. Power supply 220 V/50 Hz. Maximum power 900 W. Weight up to 60 kg. |  | 1
| 2 | Fixed rotor with capacity of 6 conical 50 ml tubes or 6 conical 15 ml tubes; Aerosol-tight certified, for centrifuging potentially dangerous substances. Rotor weight up to 3,5 kg. Easy rotor lid closing with 90° turn. Rotation speed up to 11000 rpm. |  | 1
| 3 | Conical 50 ml tubes |  | 6
| 4 | Conical 15 ml tubes |  | 6
| 5 | Adapter for 15 ml conical tubes |  | 1
| 6 | Training and installation included in the price
| 7 | 2 years warranty
| 8 | Availability of spare parts 7 years after the production stops
| 9 | Delivery time 3 months
| 10 | Response time for failure notification – 48 (fourty-eight) hours
| 11 | Maximum period for repair and replacement from the moment of declaring - 1 (one) month

## 10.7 Planetary ball mill
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.7 | Planetary Ball Mill |  | 1 |  | 0
| 1 | Mill should be ball planetary with pulverizing, mixing, homogenizing, colloidal milling; Size reduction principle impact, friction; Material feed size < 10 mm; Final fineness < 1 µm, for colloidal grinding < 0.1 µm; Batch size / feed quantity : max. 220 ml; Number of grinding stations : 1; Speed ratio 1:-1; Sun wheel speed 100 - 650 min-1; Effective sun wheel diameter 141 mm; G-force 33.3 g; Setting of grinding time 00:00:01 to 99:59:59; Interval operation : must be included with direction reversal; Interval time 00:00:01 to 99:59:59; Pause time 00:00:01 to 99:59:59; Storable standard operating procedures: 10 should be possible; Measurement of input energy : possible; Drive 3-phase asynchronous motor with frequency converter; Drive power : 750 W Electrical supply data different voltages; Power connection : 1-phase; Protection code IP 30; Standard: CE
| 2 | Zirconium oxide grinding jar 125 ml
| 3 | Grinding balls zirconium oxide 2 mm, 0.5 kg, aprox 135 ml
| 4 | Grinding balls zirconium oxide 5 mm, 200 pieces, aprox. 135 ml |  | 3
| 5 | Training and installation included in the price
| 6 | 2 years warranty
| 7 | Availability of spare parts 10 years after the production stops
| 8 | Delivery time 3 months
| 9 | Response time for failure notification – 48 (fourty-eight) hours
| 10 | Maximum period for repair and replacement from the moment of declaring - 1 (one) month

## 10.8 Rotary evaporator
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.8 | Rotary evaporator |  | 1 |  | 0
| 1 | Rotary evaporator with vertical type of cooling and cooling surface of 1500 cm2. DC motor principle. Speed range from 20 to 280 rpm. Possibility of reversible direction of rotation. Stroke of 140 mm. Heating range from room temperature up to 180 0C with resolution of ± 1 K. Heat output of 1300 W. Maximum filling volume of 3 l. Weight of device 20 kg. Permissible ambient temperature range is from 5 up to 40 0C. Permissible relative humidity is up to 80%. Voltage 220-240, frequency 50/60 Hz and power input 1400 W. Protection class according to DIN EN 60529 IP 20. With timer.
| 2 | Training and installation included in the price
| 3 | 1 year warranty
| 4 | Availability of spare parts 7 years after the production stops
| 5 | Delivery time 3 months
| 6 | Response time for failure notification – 72 (seventy-two) hours
| 7 | Maximum period for repair and replacement from the moment of declaring - 45 (fourty-five) days after receiving spare parts

## 10.9 Disperser
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.9 | Disperser |  | 1 |  | 0
| 1 | Motor rating input should be 800 W and output 500W. Volume range of the device: 0,001-2 l. Maximum viscosity 5000 mPas. Speed range of the motor: 3000-25000 rpm with 1% deviation. LED speed display and stepless speed control. Noise produced without dispersing element 75 dB (A). Extension arm diameter of 13 mm, and extension arm length 160 mm. Process type must be batch. Weight must be up to 3 kg. Permissible ambient temperature range is from 5 up to 40 0C. Permissible relative humidity is up to 80%. Voltage 220-240, frequency 50/60 Hz and power input 800 W. Protection class according to DIN EN 60529 IP 20.
| 2 | Dispersing element with volume range 0,001 - 0,1 l, 10 mm stator and 7,5 mm rotor. Maximum rotation speed up to 25000 rpm and maximum circumferential speed of up to 9,8 m/s. Material in contact with medium is PTFE. pH range 2-13. Dispersing element should be suitable for solvents and abrasives. Maximum working temperature up to 180 0C. Suspensions of fineness from 10-50 µm and ultimate fineness of emulsions 10-50 µm.
| 3 | Boss head clamp from aluminium with clamping range 6-16 mm.
| 4 | Strap clamp that has for stand diameter 8-16 mm and for vessel diameter 40-300 mm.
| 5 | Plate stand 1000 mm high with maximum load of 5 kg. Diameter of support rod 16 mm.
| 6 | Training and installation included in the price
| 7 | 1 year warranty
| 8 | Availability of spare parts 7 years after the production stops
| 9 | Delivery time 3 months
| 10 | Response time for failure notification – 72 (seventy-two) hours
| 11 | Maximum period for repair and replacement from the moment of declaring - 45 (fourty-five) days after receiving spare parts

## 10.10 Overhead stirrer
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.1 | Overhead stirrer |  | 1 |  | 0
| 1 | Motor rating input of 70 W and motor rating output of 35 W. Maximum stirring of 20 l. Maximum viscosity 10000 mPas. Speed range is from 60 to 2000 rpm with deviation of ±30 rpm. LED speed display and mechanical, stepless speed control. Extension arm diameter of 13 mm, and extension arm length 160 mm. Weight must be up to 3,2 kg. Permissible ambient temperature range is from 5 up to 40 0C. Permissible relative humidity is up to 80%. Voltage 220-240, frequency 50/60 Hz and power input 72 W. Fastening of the stirring element with chuck (diameter range 0,5 - 10 mm). Normal torque of 1,5 Nm and maximum torque at steering shaft of 150 Ncm. With hollow shaft (inner diameter 10,5 mm). Maximum output of stirring shaft 26 W. Protection class according to DIN EN 60529 IP 20.
| 2 | Boss head clamp from aluminium with clamping range 6-16 mm.
| 3 | Propeller stirrer 3-bladed with maximum speed of 2000 rpm. Stirrer diameter of 45 mm, shaft diameter and length of 8 mm and 350 mm respectively.
| 4 | Dissolver stirrer with maximum speed of 2000 rpm. Stirrer diameter of 80 mm, shaft diameter and length of 8 mm and 350 mm respectively.
| 5 | Anchor stirrer with maximum speed of 1000 rpm. Stirrer diameter of 45 mm, shaft diameter and length of 8 mm and 350 mm respectively.
| 6 | Strap clamp that has for stand diameter 8-16 mm and for vessel diameter 40-300 mm.
| 7 | Plate stand 1000 mm high with maximum load of 5 kg. Diameter of support rod 16 mm.
| 8 | Training and installation included in the price
| 9 | 1 years warranty
| 10 | Availability of spare parts 7 years after the production stops
| 11 | Delivery time 3 months
| 12 | Response time for failure notification – 72 (seventy-two) hours
| 13 | Maximum period for repair and replacement from the moment of declaring - 45 (fourty-five) days after receiving spare parts

## 10.11 Stirring hot plate
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.11 | Stirring Hot Plate |  | 1 |  | 0
| 1 | Top plate • width at least 150 mm • length at least 150 mm • made of aluminum
| 2 | Maximum temperature 450°C
| 3 | Temperature accuracy ±1°C
| 4 | Stirring speeds should cover 100 and 1500 rpm
| 5 | Dimensions not larger than (width x length x height): 200mm x 270mm x 120mm
| 6 | Timer up to 99 hours, 59 minutes with auto-off option
| 7 | Storing up to 10 programs
| 8 | Possibility to define multiple temperatures, temperature rates, stirring speeds and timed events for each program
| 9 | Possibility to repeat any program
| 10 | Offers RS232 interface for control or record data via computer
| 11 | Cord with plug at least 1.5 meter long included
| 12 | CE marked
| 13 | 1 year warranty
| 14 | Delivery time 3 months

## 10.12 SLA 3D printer
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.12 | SLA 3D Printer |  | 1 |  | 0
| 1 | Technology: Stereolithography (SLA) based 3D printer
| 2 | Building Volume: minimum 140x140x170 mm
| 3 | Software • supports uploading of a digital model • automatic positions of the digital model for optimal printing • enables importing, duplicating, aligning and resizing digital models • enables generation of support structures
| 4 | Factory set device with no need for additional calibration and settings
| 5 | Automatic resin fill system
| 6 | Automatic pre-heating and maintaing a constant temperature of resin at 35°C
| 7 | Temperature control is done by self-heating tank
| 8 | Automatically generated supports that are easily removable
| 9 | Laser Specifications EN 60825-1:2007 certified Class 1 Laser Product 405nm violet laser minimum 200mW laser
| 10 | Peel Mechanism Sliding Peel Process with wiper
| 11 | Multiple Layer Thickness: minimum 25 µm
| 12 | Laser Spot Size defined by full width at half maximum not larger than 150 µm
| 13 | Connectivity Wifi, Ethernet and USB
| 14 | Power Requirements 100–240 V 1.5 A 50/60 Hz 65 W
| 15 | Dimensions Not larger than 350x350x550 mm
| 16 | Printer control by touch-screen and push-button
| 17 | System spares kit must include the following: • 5 resin tanks made from shatter-resistant polycarbonate • long term use resin tank with at least 10 times longer life time than conventional resin tanks • Clear Resin Cartridge 1L • Black Resin Cartridge 1L • White Resin Cartridge 1L • Flexible Resin Cartridge 1L • Rigid Resin Cartridge 1L • Color Resin Cartridge 1L
| 18 | Training and installation included in the price
| 19 | Warranty: 24 months
| 20 | Delivery time 3 months

## 10.13 Desktop filament extruder
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
| 10.13 | Sensors fabrication and characterisation |  | 1 |  | 0
| 1 | Maximum extruder operating temperature up to 450°C with ceramic heating
| 2 | 4 controllable heating zones
| 3 | Speed: 1 kg PLA/hour
| 4 | Extruder RPM range: 2 rpm – 15 rpm
| 5 | Output filament diameter range: 0.5 mm – 3 mm
| 6 | Optical sensor accuracy: at least 45 µm or better
| 7 | Nozzle extruder:Diameter 4 mm, replaceable
| 8 | Nitride hardened steel extruder screw
| 9 | 3 stage compression screw
| 10 | Hopper Volume: 2 liters
| 11 | Includes: 2 spools, 1 pouch PLA - 1200g, 1 pouch cleaning - 350g, 2 spool holders, USB cable, power cable, high volume hopper |  | 1
| 12 | Nozzle package: Nozzle 4mm, compression fitting ferrule, compression fitting nut |  | 4
| 13 | Puller wheel |  | 4
| 14 | Ceramic band heater barrel 220V |  | 1
| 15 | Ceramic band heater die head 220V |  | 1
| 16 | Thermocouple 0.9 m |  | 1
| 17 | Thermocouple 1.1 m |  | 1
| 18 | Power: 230 V
| 19 | Power cable |  | 1
| 20 | Warranty: 12 months
| 21 | Delivery time 3 months

## 10.14 PCB prototyping
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.14 | PCB prototyping |  | 1 |  | 0
| 1 | Main unit • Maximum PCB size at least 100x150mm • Mechanical resolution at least 10 µm • Software resolution minimum 1.25 µm • X axis travel range at least 100 mm • Y axis travel range at least 200 mm • Z axis travel range at least 20 mm • Max travel speed, X axis at least 1000 mm/min • Max travel speed, Y axis at least 1000 mm/min • Max travel speed, Z axis at least 1000 mm/min • Maximum spindle speed 30,000 rpm or more • Dimensions not bigger than 300x500x400 mm • Weight less than 15 kg
| 2 | Milling and drilling • Support for double sided boards • Trace width and clearance down to 0.2mm • Milling feed up to 1000mm/min • Industry standard 1/8” shank tools
| 3 | Solder paste dispensing • 10ml solder paste syringes • Dispensing needle diameter down to 0.25mm • No need for compressed air supply
| 4 | SMD pick-and-place • Semi-automatic placement • Built-in vacuum generator • Support for component tapes of various sizes or individual components
| 5 | Dust extraction add-on • Enables connection to dust extractor • Supports attachment to 36 mm outer diameter hose
| 6 | Microscope add-on • USB microscope • Adapter for mounting microscope on SMD pick-and-place • LED illumination
| 7 | Starter pack that includes PCB laminates, drilling/milling spacers, PCB tools insulation milling V tool 0.2 to 0.5 mm, contour end mill 1 mm, drills from 0.6 to 3 mm diameter), dispensing needles
| 8 | Warranty 2 years
| 9 | Delivery time 3 months

## 10.15 Micropositioning probes
Bidder:
Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | DAP Unit price | Total Price per line item | Insert page no. in techical documentation
Sensors fabrication and characterisation
| 10.15 | Micropositioning probes |  | 1 |  | 0
| 1 | Left micropositioner with motion in 3 axis, magnetic base and spring head. The minimum accuracy of motion at least 5 μm. Device connection cable with a male pin jack. Micropositioner is used as holder and moving element of probe tip system (probe tip holder and probe tip) for precise probing of microelectronic components and integrated circuits. |  | 1
| 2 | Right micropositioner with motion in 3 axis, magnetic base and spring head. The minimum accuracy of motion at least 5 μm. Device connection cable with a male pin jack. Micropositioner is used as holder and moving element of probe tip system (probe tip holder and probe tip) for precise probing of microelectronic components and integrated circuits. |  | 1
| 3 | Straight probe tip holder at least 4 inch long, below 0.1 inch in diameter, without insulation. Probe tip holder should have spring head for probe tip. |  | 2
| 4 | Bendable straight tungsten probe tip, with diameter not more than 0.02 inch, length above 1 inch. The probe should be intended for probing of geometries above 5 μm. Tip of the probe tip should be below 5 μm in diameter. |  | 5
| 5 | Bendable straight tungsten probe tip, with diameter not more than 0.02 inch, length above 1 inch. The probe should be intended for probing of geometries in submicrometer scale. Tip of the probe tip should be below 1 μm in diameter. |  | 5
| 6 | Warranty 1 year
| 7 | Delivery time 3 months