## Price Schedule
Bidder: |  |  |  | Offered price (currency):
LOT 5 | Radiology Equipment
ID | Technical Specification Requested | QTY | Technical Specification Offered (model) | Unit price | Total price per line item
4.1 | Volume CT Scanner | 1 |  |  | 0
4.2 | Premium ultrasound machine for radiology department | 1 |  |  | 0
4.3 | Fixed Radiography system with wireless flat panel | 1 |  |  | 0
4.4.1 | RF System for MSK procedures | 1 |  |  | 0
4.4.2 | RF system for abdominal and urology procedures | 1 |  |  | 0
4.5 | Mobile radiography digital system with 2 FPD | 1 |  |  | 0
4.6 | PACS | 1 |  |  | 0
4.7 | Infusion pump, dual channel, compatible with magnetic resonance | 1 |  |  | 0
|  |  |  | Total Price per Lot | 0

## Technical specifications
Bidder:
LOT 4 |  | Radiology Equipment Specification |  |  | Date:
Total DAP price:
Line item No. | ID | Technical Specification Requested | Technical Specification Offered | QTY | Unit Price DAP | Total Price per line item | Insert page no. in technical documentation
4.1 |  | Volume CT Scanner |  | 1
| 1 | GANTRY
| 1.1. | Gentry aperture at least 78 cm
| 1.2. | Physical or digital gentry tilt at least -24⁰/+30⁰
| 2 | PATIENT COUCH
| 2.1. | Vertical movement of patient couch in range of at least 40 cm with the lowest height, maximum 55 cm
| 2.2. | Patient couch movement controls on gentry, on operators console and at patient couch
| 2.3. | Scanning range in horizontal direction with extensions at least 185 cm
| 2.4. | Patient couch maximum load capacity at least 200 kg
| 3 | GENERATOR
| 3.1. | Highest nominal generator X-ray tube power at least 70 kW
| 3.2. | Highest nominal X-ray tube current at least 600 mA
| 3.3. | Highest voltage at least 135 kV
| 4 | X-RAY TUBE
| 4.1. | X-ray tube has heat storage capacity at least 6.8 MHU (heat capacity refers to anode heat capacity or housing of anode depending on manufacturer's technology)
| 4.2. | At least two focal spots
| 5 | DETECTOR SYSTEM
| 5.1. | Number of detector elements in one detector row excluding reference detector elements at least 800
| 5.2. | /
| 6 | ACQUISITION PARAMETERS
| 6.1. | Ability to choose from at least 2 different scan fields
| 6.2. | Greatest scan field at least 50 cm and smallest maximum 25 cm
| 6.3. | Shortest time for one full rotation (360°) maximum 0.28 seconds
| 6.4. | Longest continuous time for spiral scanning at least 60 seconds
| 6.5. | Dynamic study acquisition without patient couch movement at least 160 mm
| 6.6. | System has to have active patient dose reduction option with dose reduction of at least 50 % (Safire, Iris, Aidr3D, iDose, VEO, ASIR)
| 6.7. | System has to have package for acquisition synchronized with ECG signal (hardware and software)
| 6.8. | Software package that enables possibility to acquire image of three body regions in one run with different pitch settings, where one body part is possibly scanned with ECG triggering, to obtain image without artefact related to movement of heart and attached structures
| 7 | CONSOLE
| 7.1. | Flat-screen color monitor
| 7.2. | „DICOM“ protocols at least: DICOM Print, DICOM Storage SCU, DICOM MWM SCU
| 7.3. | Program for dynamic bolus tracking for precise scan time determination and for activation at automatic bolus detection
| 7.4. | Ability to define in total at least 4 different regions of interest during scanning with automatic bolus detection
| 7.5. | On-line storage capacity at least 700,000 non-compressed images in 512x512 pixel matrix
| 7.6. | Fastest image reconstruction time at least 55 images per second with all dose reduction options active
| 7.7. | Advance reconstruction algorithm for reduction of artifacts caused by metal presence in patient’s body (not Dual Energy)
| 7.8. | Software to predict movement during image acquisition triggered with ECG signal in cardiac CT studies in volumetric image acquisition and use of anatomically correct reconstruction in order to minimize artifacts caused by movement
| 7.9. | MPR real time and 3D display
| 8 | DIAGNOSTIC TOOLS
| 8.1 | Integrated data base for all clients
| 8.2 | 3 clients with following hardware specification:
- Clinical, DICOM calibrated color “LCD” display with resolution at least 1920x1080 and display diagonal at least 24”
- Windows® 10 Professional 64-bit brand name PC with at least: memory 500Gb, Intel® Core™ i3 processor, 2Gb RAM, or equivalent
| 8.3 | Ability for system user to access system from any client
| 8.4 | At least 3 advance client licenses for multidimensional digital images post-processing, analysis, review and distribution with following advance tools:
- 2D and 3D, Volume Rendering,
- MIP, MinIP, Average,
- Automatic Segmentation of visible blood vessels with Curved MPR for Vascular analysis, with automatic removal of surrounding anatomical structures and automatized stenosis measurement and detection of internal and external blood vessel lumen,
- CT Virtual Endoscopy of larynx and bronchi,
- CT Brain Perfusion with display of blood flow, blood volume and Mean Transit Time (MTT), to define presence of acute cerebral infarcts. System has automatic correction of artefacts from patient movement,
- 4D CT Brain Perfusion with display of blood flow, blood volume and Mean Transit Time (MTT), to define presence of acute cerebral infarcts. System has automatic correction of artefacts from patient movement. Software package makes calculation in all three orthogonal planes of volume on any slice. Ability to display 4D DSA contrast flow through blood vessels,
- Segmentation of all musculoskeletal structures with ability of easy implant separation,
- Cardiac CT analysis: automatic segmentation of the heart in one or more phases, an automatic segmentation of coronary arteries, curved MPR and blood vessels labeling, the automatic determination of the center line of the blood vessel and its edges,
- CT analysis of coronary plaques, with automatic segmentation of blood vessel and automatic determining the contour of the lumen and the outer wall of the blood vessel, with measurement and display of the surface and the lumen diameter, the volume of the plaque and plaque index,
- CT calcium scoring application for visualization, measurement and reporting on the calcification of coronary blood vessels,
- CT Functional analysis of the heart for the assessment of left ventricular function: automatic segmentation of the heart, the left ventricle in multiple phases, an automatic calculation of ESV, EDV, SV, EF, CO and myocardial mass, display of the heart in the short axis, in the long axis and four-chamber display, calculation of motion and wall thickening with live 3D display of the heart,
- CT Functional analysis of the heart for the assessment of right ventricle and left atrium function.
- The software package for analysis of myocardial perfusion with automatic segmentation of the myocardium and the detection of parts with a decreased perfusion. Ability to load two series, one made at rest and the other after stress and to compare results. Coloring of the walls by myocardial perfusion on MPR images and the automatic processing of the perfusion bull-eye diagram,
- CT package for Aorta segmentation for easy planning of interventions on abdominal and thoracic aortic aneurysms,
- Automatic segmentation of Circle of Willis for easy malformation detection,
- Automatic Carotids segmentation,
- CT Perfusion 4D body. Package for calculation of perfusion of other organs in any axis. Special calculation method of perfusion of the organs that are fed with one blood vessel and different method for those organs that are fed with two blood vessels,
- Virtual endoscopy/colonoscopy, with the analysis and integration of endoluminal views with conventional MPR views, and an interactive navigation through the cavity (colon),
- A software package for automatic segmentation of the lungs and airways. Nodules segmentation with a single click to measure the size. Automatic tracking of nodules through multiple views to calculate elapsed time in days, to calculate time of doubling and to calculate the changes in percentage,
Ability to accept images in ENHANCED DCSOM format and following DICOM possibilities: DICOM Store SCU/SCP, DICOM Query/Retrieve SCU/SCP, DICOM PRINT, DICOM MWM SCU, DICOM PPS SCU and DICOM Storage Commitment
| 9 | Dual head contrast media injector
| 10 | 5 days on site training for staff (doctors and technicians) who will work on the system. Training has to be done by certified Application specialist.
| 11 | Complete system integration and commissioning
4.2 |  | Premium ultrasound machine for radiology department |  | 1
| 1 | Beamformer Technology: The system shall utilize a last-generation beamformer technology which supports high resolution and contrast
| 2 | Hardware channels: The system shall provide minimum 192 hardware channels
| 4 | Dynamic range: The system shall provide a dynamic range of at least 250 Db.
| 5 | Console: 6 way adjustable control panel
| 6 | The system shall have Touch Screen with at least 13 inch (+/- 1’’) touch panel and with digital TGC with predefined 4 curves
| 7 | The system shall have full-size QWERTY keyboard with backlight
| 8 | Control panel with height adjustment
| 9 | Integrated Gel Warmer
| 10 | B/W thermal printer, available
| 11 | The system shall have four independent swivel wheels to allow for positioning in tight spaces and front wheel central lock
| 12 | The system shall have 23“ (+/- 1’’) LCD or OLED or LED technology, with resolution 1920x1080 pixels. Height, rotate and tilt adjustable
| 13 | Active Port Probes: The system shall have 4 active transducer ports (pinless connectors) in a convenient to access location to support proper user positioning and reduce user reach.
| 14 | Scan modes and imaging features: B-mode Imaging, Frame Rate in B mode: The system shall have minimum of 1900 fps, Color Doppler, PW Doppler, Hypersensitive Color Doppler, Power Doppler, Triplex Mode, ElastoScan, Contrast Imaging, Fusion Imaging
| 17 | Transducers:
| 18 | Convex Transducer: Frequency range: min. 1 – 7 MHz,±2MHz number of elements min. 160, single crystal or matrix technology
| 19 | Micro-convex Transducer: Frequency range: min. 4 – 8MHz±2MHz, number of elements min. 128, FOV min.90°+/-10°
| 20 | Linear Transducer: Frequency range: min. 3 – 10MHz±2MHz, number of elements min.192, min., FOV max.40mm+/-10mm
| 21 | Linear Transducer: Frequency range: min. 4 – 18 MHz±3MHz, number of elements min.280, min., FOV max.50mm+/-10mm
| 22 | Linear Transducer Intraoperative : Frequency range: min. 5 – 15 MHz±2MHz, number of elements min.128, min., FOV max.30mm
| 23 | General features :
| 24 | CINE Function: The system shall perform CINE Function, 12.000 cine images minimum or 1GB Cine memory minimum
| 25 | Hard Drive: The system shall include at least 500GB of dedicated hard drive.
| 26 | Maximum imaging depth, min. 40cm
| 27 | DVD Drive: available
| 28 | Exported Image formats: The system shall export images and clips in PC compatible JPG format, AVI format and DICOM format
| 29 | USB Ports: The system shall have min 3 user-accessible USB 2.0 ports.
| 30 | DICOM Connectivity and DICOM Services
| 31 | Advanced features:
| 32 | Real-time Elastography with quantification for Thyroid exams
| 33 | Shearwave elastography for Liver exams
| 34 | Contrast Imaging Studies available on the system
| 35 | Available active Fusion imaging for realtime imaging in conjunction with other volumetric imaging modalities (CT, MRI)
| 36 | The system provides projection of simulated needle path for accurate interventional procedures
| 37 | Automatic image optimization in B and color mode and the range of velocity and scale in PW Doppler by pressing one button.
| 38 | System supports visualization of slow flow microvascular structure for detailed examination of blood flow in relation to surrounding tissue
4.3 |  | Fixed Radiography system with Wireless Flat Panel Detectors - Ceiling mounted |  | 1
| 1 | HV GENERATOR
| 1.1 | Output power: min. 50 kW
| 1.2 | kV range for radiography: 40-150 kV or more
| 1.3 | mA range: min. 10-600mA
| 1.4 | Time range min. 0.01-5 sec. in steps
| 2 | X-RAY TUBE
| 2.1 | Dual focus: small focal spot max 0.6 mm and big focal spot max. 1.3 mm
| 2.2 | Anode heat capacity: min. 300 kHU
| 2.3 | Automatic collimation
| 2.4 | Collimator with individual blade control for radiation dose reduction
| 3 | CEILING-MOUNTED TUBE STAND&WALL STAND
| 3.1 | Fully automated movement
| 3.2 | Translation X(Longitudinal) min. 3500mm x Y(Lateral) min 3000mm x Z(Vertical) min. 1300 mm
| 3.3 | Motorized Vertical Translation
| 3.4 | Arm Rotation movement min ± 135º
| 3.5 | Tube Rotation movement min ± 120º
| 3.6 | Multi functional Touch Screen Display
| 3.7 | WALL STAND motorized up/down range min.400-1800mm
| 3.8 | Detector tilting motorized min. range: -20/+90
| 3.9 | Vertical movement motorized
| 3.1 | Auto Tracking/Auto Centering
| 4 | PATIENT TABLE
| 4.1 | Patient Table top dimensions, length x width, min. 2300x800 mm
| 4.2 | Max patient Weight (kg): min. 270 kg
| 4.3 | Up and Down range: 550-850mm
| 4.4 | Longitudinal movement range: min 600 mm
| 4.5 | Transverse movement range: min 240 mm
| 4.6 | Longitudinal tracking of detector with longitudinal travel and rotation of tube
| 4.7 | Shared bucky for both type of detectors
| 4.8 | Wireless foot switch
| 4.9 | 2x FPD (FLAT PANEL DETECTOR/ Wireless)
| 4.1 | Flat Panel Detector (Patient Table)
| 4.11 | Effective area dimensions: min. 34x42 cm
| 4.12 | Resolution: min. 2400x3000 pixels
| 4.13 | A/D conversion: min. 14 bits
| 4.14 | Pixel size: max 140 µm
| 4.15 | Spatial resolution: min 3,5 lp/mm
| 4.16 | Scintilator material: CsI (Cesium Iodide)
| 4.17 | Detective Quantum Efficiency (DQE) min. 70%
| 4.18 | Additional spare battery for FPD with Battery charger
| 4.19 | Flat Panel Detector (Wall Stand)
| 4.2 | Effective area dimensions: min. 42.5x42.5 cm+/- 1%
| 4.21 | Resolution: min. 3000x3000 pixels
| 4.22 | A/D conversion: min. 14 bits
| 4.23 | Pixel size: max 140 µm
| 4.24 | Spatial resolution: min. 3,5 lp/mm
| 4.25 | Scintilator material: CsI (Cesium Iodide)
| 4.26 | Detective Quantum Efficiency (DQE) min. 75%
| 4.27 | Additional spare battery for FPD with Battery charger
| 5 | ACQUISITION WORKSTATION FOR OPERATION AND IMAGE POST PROCESSING
| 5.1 | Computer with min. 1 LED HD monitor diagonal size min. 23 inches in control room or 2 LED HD monitors diagonal size of min. 19 inches in control room¦
| 5.2 | HDD capacity: min. 500 MB
| 5.3 | DVD/ CD export in DICOM format with automatic export of DICOM viewer on DVD/ CD
| 5.4 | DICOM functionality: min.. print, send, query/retrieve, MW, MPPS
| 5.5 | Tools for image post-procesing: contrast and brightness adjustment, „Edge enhancement ", zoom, distance and angle measurement, negative/positive inversion, annotation, R/L markers
| 5.6 | Anatomical Programmed Radiography (APR matching, x-ray conditions, image processing parameters, markers) enabling dose optimization
| 5.7 | Image acquisition times, final image max 10 sec.
| 5.8 | Dose Area Management available for cumulative patient dose tracking
| 6 | Advanced technology and additional features:
| 6.1 | /
| 6.2 | Software grid available for scatter radiation reduction
| 6.3 | Technic for suppressing the appearance of bones in chest images available
| 6.4 | Stitching imaging: single image of complete anathomy (spine and long bones) is provided by synchronized automated movement of tube and detector
| 6.5 | Pediatrics exposure management dependent on patient weight
4.4.1 |  | Radiography Fluoroscopy system with dynamic Flat Panel Detector for diagnostic MSK procedures |  | 1
| 1 | HV GENERATOR
| 2 | Output power: min. 80 kW
| 3 | Frequency: min. 50 kHz
| 4 | mAs range for radiography: 0.5 - 800 mAs or more
| 5 | kV range for radiography: 40-150 kV or more
| 6 | Tube current value for radiography: min. 800mA
| 7 | kV range for pulsed fluoroscopy: 50-110 kV or more
mA range for pulsed fluoroscopy: 0,5-20 mA or more
| 8 | Anatomic programs: min. 200
| 9 | Pulsed fluoroscopy with ability of 30 fps
| 10 | Automatic Exposure Control (AEC)
| 11 | Automatic kV and mA regulation in fluoroscopy mode
| 12 | X-RAY TUBE
| 13 | High speed rotating anode with min. 8000 rot/min
| 14 | Dual focus: small focal spot max 0.7 mm and big focal spot min. 1.0 mm
| 16 | Anode heat capacity: min. 750 kHU
| 17 | REMOTE CONTROLLED PATIENT TABLE
| 18 | Table tilt: min. +90º/-45°
| 19 | Table top dimensions, length x width, min. 2100x700 mm
| 20 | Max table load in any position and table tilt: min. 180 kg
| 21 | Height adjustable patient table with lowest position (distance to floor) of 50 cm
| 22 | Lateral motorized movement of min. 24 cm
| 23 | Longitudinal motorized movement of X-ray tube and Flat Panel Detector (FPD) with variable speed in range of 100cm
Motorized change of SID. Min SID of 1150 mm. Max SID of 1500 mm.
| 24 | Motorized angulation of X-ray tube assembly: min. ± 40º
| 25 | X-ray tube rotation when patient table is in vertical position: min. 90˚
| 26 | Automatic collimator with permanent filtration min. 1 mm Al and with additional filtration min. 0.1, 0.2 and 0.3 mm Cu
| 27 | Integrated control panel on front side of patient table or on front side of X-ray tube housing, with minimum of following functions: table tilt, SID adjustment, X-ray tube angulation
| 28 | Remote controlled control console in control room with following functions: Patient table movement, X-ray tube and FPD movement, motorized compression adjustment. Control panel has touch screen user interface for generator exposure parameters adjustment, anatomic program selection, field of view selection, collimator selection, imaging mode selection (digital pulsed fluoroscopy, digital serial acquisition, digital radiography), selection of fps in pulsed fluoroscopy mode and digital serial acquisition mode, tools for image postprocessing and complete exposure in fluoroscopy and radiography modes
| 29 | Additional independent control console in examination room, portable (on wheels) with following functions: Patient table movement, X-ray tube and FPD movement, motorized compression adjustment. Control panel has touch screen user interface for generator exposure parameters adjustment, anatomic program selection, field of view selection, collimator selection, imaging mode selection (digital pulsed fluoroscopy, digital serial acquisition, digital radiography), selection of fps in pulsed fluoroscopy mode and digital serial acquisition mode, tools for image postprocessing and complete exposure in fluoroscopy and radiography modes
| 30 | This control console has to be identical to control console in control room so operator does not need to leave the examination room to do all operations wich are part of all fluoro/radiography patient exams. Only patient data entry functionality should be available on control console in the control room.
| 32 | Foot support-1 pc. Arm rest 2 pcs.
| 33 | AEC chamber with min. 3 fields
| 34 | Program for MSK procedures
| 35 | DYNAMIC FPD (FLAT PANEL DETECTOR)
| 36 | Active field of view dimensions: min. 42x42 cm with possibility of 4 different field of view selection
| 37 | Resolution: min. 2800x2800 pixels
| 38 | Dynamic range: min. 14 bits
| 39 | Pixel size: max 150 µm
| 40 | Spatial resolution: min. 3,4 lp/mm
| 41 | ACQUISITION WORKSTATION FOR OPERATION AND IMAGE POST PROCESSING
| 42 | Computer with min. 1 LCD/TFT monitor diagonal size min. 19 inches in control room
| 43 | Additional TFT/ LCD monitor in examination room diagonal size min. 19 inches, resolution min. 1MP
| 44 | Trolley with wheels for TFT/LCD monitor in the examination room
| 45 | Digital pulsed fluoroscopy: min. 1024 x 1024, 12 bits with min. 30 fps, with ability of fps selection
| 46 | LIH (Last Image Hold)
| 47 | Virtual collimation function (non exposure collimation on LIH)
| 48 | Digital serial acquisition: min. 1024x1024, 12 bits, with min. 7.5 fps
| 49 | HDD capacity: min. 10.000 images in 1024x1024 matrix, 12 bits
| 50 | DVD/ CD export in DICOM format with automatic export of DICOM viewer on DVD/ CD
| 51 | DICOM functionality: min.. print, send
| 52 | Dedicated acquisition and postprocessing parameters optimized for children of different age and size
53
| 54 | Installation of power distribution box for new system
| 55 | Bidder has to adopt floor after system cabling (antistatic floor)
4.4.2 |  | Radiography Fluoroscopy system with dynamic Flat Panel Detector for interventional and diagnostic abdominal and urology procedures |  | 1
| 1 | HIGH-VOLTAGE, HIGH-FREQUENCY X-RAY GENERATOR
| 2 | Maximum output power, at least 80 kW
| 3 | Maximum inverter frequency: min. 50 kHz
| 4 | mAs range for radiography: 0,5 – 800 mAs or wider
| 5 | kV range for radiography: 40 – 150 kV or wider
| 6 | mA range for radiography: min. 10 - 800 mA
| 7 | mA range for fluoroscopy: min. 0,3 – 20 mA
| 8 | Anatomical programs that could be registered, associated with examination area and conditions: min. 200
| 9 | Pulse fluoroscopy
| 10 | Automatic exposure control AEC
| 11 | X-RAY TUBE
| 12 | Fast rotating anode: min. 8000 rpm
| 13 | Dual focus: max. 0.7/1.2 mm
| 14 | Anode heat capacity: min. 750 kHU
| 15 | Automatic Iris collimation
| 16 | REMOTE CONTROLLED PATIENT TABLE
| 17 | The table can be tilted without changing the observation position
| 18 | Table tilt: min. +90°/-45°
| 19 | Table plate dimensions : min. 2100x700 mm
| 20 | Maximum table load in horizontal position: min. 300 kg or Maximum table load in any position: min. 280 kg
| 21 | Motorized vertical movement
| 22 | Lowest table plate height: min. 50 cm
| 23 | Motorized lateral movement: 24 cm
| 24 | Longitudinal motorized synchronized movement of X-Ray tube and FPD (imaging chain) without patient table movement.
| 25 | Progressive variable imaging chain movement speed
| 26 | min. 150 cm/sec
| 27 | Support for legs, hand grips 2 pcs.
| 28 | Motorized source to FPD distance (SID) adjustment.
| 29 | SID range: min 115 – 150cm or wider
| 30 | Motorized angulation of X-Ray tube holder: min. ±35°
| 31 | Motorized rotation of X-Ray tube in patient table vertical position: 90°
| 32 | Automatic collimator with optional asymmetric collimation
| 33 | Virtual collimation (without radiation)
| 34 | Digital motorized tomography in selected table tilt angle.
| 35 | Minimum 3 angles selectable.
| 36 | Motorized compression for GI procedures
| 37 | Compression mechanism for IVU procedures
| 38 | Additional control console in examination room, movable on wheels, which enables control of all patient table movements, adjustment of generator output (exposure conditions) and complete exposure and acquisition process control in both fluoroscopy and radiography procedures.
| 39 | Additional set of foot commands in examination room, for exposures in both fluoroscopy and radiography mode, independent of exposure commands already existing on control console in examination room.
| 40 | DYNAMIC FPD ( FLAT PANEL DETECTOR )
| 41 | Quote manufacturer of FPD and model
| 42 | Dimensions of FPD: min. 43x43 cm, with possibility to select 4 different size field of view.
| 43 | Resolution matrix: min. 2800 x 2800 pixel
| 44 | Dynamic range: min. 14 bit
| 45 | Pixel pitch: max. 150 μm
| 46 | ACQUISITION WORKSTATION FOR CONTROL AND IMAGE PROCESSING
| 47 | Computer with min. 1 LCD/TFT monochromatic high resolution monitor with min 19" display in operators room and min 1 LCD/TFT monochromatic high resolution monitor with min 19" display in examination room
| 48 | Additional min 1 LCD/TFT monochromatic high resolution monitor with min 19" display in examination room
| 49 | Trolley for monitors in examination room
| 50 | Digitalization depth: min. 12 bit
| 51 | Digital pulse fluoroscopy: min. 15 fps in matrix of min. 1024x1024, 12 bit and the biggest field of view
| 52 | LIH ( Last Image Hold )
| 53 | Digital acquisition: min. matrix 1024x1024, 12 bit at the biggest field of view
| 54 | Digital serial acquisition: min. matrix 1024x1024, 12 bit, 7,5 fps at the biggest field of view
| 55 | HDD capacity: min. 10.000 images in matrix 1024x1024, 12 bit
| 56 | Recording on CD/DVD in DICOM format
| 57 | DICOM functions: min. print, store, Modality Worklist (MWM)
| 58 | IMAGE POST-PROCESSING:
| 59 | Filter
| 60 | Fluoro-loop
| 61 | Edge enhancement
| 62 | Multi-image display
| 63 | Zoom
| 64 | Negative/positive inversion
| 65 | Distance measurements
| 66 | ADVANCED APPLICATIONS:
| 67 | /
| 68 | /
| 69 | DRY LASER PRINTER
| 70 | Resolution minimum 500 dots per inch
| 71 | Two films trays at least
| 72 | Possibility of printing various types of films: 35x43 cm; 35x35 cm; 26x36 cm; 25x30 cm; 20x25 cm;
| 73 | Processing capacity at least 70 sheets/hour size 35x43 cm
| 74 | Daylight film loading
| 75 | DICOM compatibility
| 76 | Recording gradation 14 bits at least
| 77 | Automatic density adjustment
| 78 | Maximum weight 110kg
4.5 |  | Mobile radiography digital system with 2 FPD (Flat Panel Detectors) |  | 1
| 1 | HV GENERATOR
| 2 | Power min. 30 kW
| 3 | Voltage in range of min 40-133 kV
| 4 | Maximum current value, min. 300 mA
| 5 | mAs values in range of min 0.4-320 mAs
| 6 | Minimum exposure time 1 ms
| 7 | Exposure with battery power
| 8 | Battery power indicator
| 9 | 220V input for battery charging
| 10 | Hand switch exposure
| 11 | X-RAY TUBE
| 12 | /
| 13 | /
| 14 | Anode heat capacity, min. 120 kHU
| 15 | Single focus max 0.8 mm or dual focus, small max 0.8 mm, big min 1.2 mm
| 16 | Manual or automatic collimation
| 17 | Collimator rotation min +/- 90º
| 18 | Auto or manual filter included
| 19 | WIRELESS PORTABLE FLAT PANEL DETECTOR
| 20 | Wireless (WLAN) communication between detector and acquisition workstation integrated in the system
| 21 | Scintillator material: CsI (Cesium Iodide)
| 22 | Detective Quantum Efficiency (DQE) min. 70%
| 23 | Effective area dimensions: min. 34x42 cm
| 24 | Acquisition depth min. 14 bits
| 25 | Pixel size max. 150 μm
| 26 | Detector weight: max. 4 kg
| 27 | Detector charged automatically in the cabinet
| 28 | Additional spare battery for FPD with Battery charger
| 29 | WIRELESS PORTABLE FLAT PANEL DETECTOR
| 30 | Wireless (WLAN) communication between detector and acquisition workstation integrated in the system
| 31 | Scintillator material: CsI (Cesium Iodide)
| 32 | Detective Quantum Efficiency (DQE) min. 66%
| 33 | Effective area dimensions: max. 25x30 cm
| 34 | Acquisition depth min. 16 bits
| 35 | Pixel size max. 150 μm
| 36 | Detector weight: max. 1.80 kg
| 37 | /
| 38 | Additional spare battery for FPD with Battery charger
| 39 | ACQUISITION WORKSTATION INTEGRATED IN SYSTEM
| 40 | LCD touch screen monitor diagonal size min. 15"
| 41 | Disk storage capacity of min 10.000 images
| 42 | Image preview in max 4 seconds after exposure
| 43 | Gray scale inversion
| 44 | Zoom
| 45 | Image Rotation
| 46 | L/R markers
| 47 | Annotations
| 48 | Wireless interface for exporting images in DICOM format
| 49 | DICOM functionality min: Send, Storage Commitment Print, MWM
| 50 | Image archiving on CD/DVD and USB in DICOM format
| 51 | Exposure Index (EI) display
| 52 | Pediatric Exposure Management provided
| 53 | Dose Area Management available for cumulative patient dose tracking
| 54 | MECHANICAL PARAMETERS
| 55 | Motorized back/forward movement on batteries
| 56 | Buttons for fine movement of the system back/forward on front side of collimator or on arm, next to X-ray tube, for easy and faster tube positioning, without need to go in front of the system
| 57 | Distance from focal spot to floor in vertical direction, in range of min 500 – 2100 mm +/-20%
| 58 | Stand rotation min +/- 90°
| 59 | X-ray tube rotation around arm min range: 300º
| 60 | X-ray tube rotation around vertical axes min 90°/-10°
| 61 | System weight max 450 kg
| 62 | System width max 600 mm
| 63 | System height in transport position, max 1600 mm
| 64 | Advanced technology and additional features
| 65 | /
| 66 | Software grid available for scatter radiation reduction
| 67 | Software for enhancing visibility of tube and lines without additional exposure provided
| 68 | Pediatrics exposure management dependent on patient weight
4.6 |  | PACS |  | 1
| 1 | HARDWARE
| 2 | SERVER HARDWARE
| 3 | Rack mount
| 4 | At least 2 x Processor 2.0GHz/8-core
| 5 | At least 64GB RAM
| 6 | At least 2 x SAS 6G 1.2TB 10Krpm HOT PL hard drives in RAID1 configuration
| 7 | At least 6TB archive storage in RAID5 configuration, disk type SAS 6G 7,2Krpm HOT PL
| 8 | WORKSTATION HARDWARE – 2 units
| 9 | AMD or Intel processor at least 3GHz, 4 Core
| 10 | At least 16GB RAM
| 11 | Hard drive at least 500GB SATA 5,4KRpm
| 12 | CD/DVD unit
| 13 | USB mouse, USB keyboard with Serbian characteristics
| 14 | Administrative RIS monitor: at least 20", full HD
| 15 | Medical diagnostic colour monitor:
| 16 | at least 21", LED, Maximum luminance at least 800cd/m2, DICOM calibrated luminance of at least 500cd/m2, contrast ratio of at least 1400:1, Response time: maximum 12ms
| 17 | Monitor and calibration software from the same manufacturer; Built-in ambient light sensor on the front of the screen, for calibration of the monitor with calibration software; Displays monochrome images in accordance with DICOM, as well as color images in accordance with DICOM.
| 18 | 2 UNITS
| 19 | PACS SOFTWARE
| 20 | Work on the "Microsoft Windows" platform (supported 32-bit and 64-bit versions of operating systems: Minimum MS Windows 7
| 21 | The use of SSL certificates for security
| 22 | Supports an unlimited number of named users
| 23 | The system supports different user rights depending on the function: a radiologist, technician, clinician
| 24 | Setting the period of time for logout in case of inactive users
| 25 | The integration of all DICOM compliant modalities in the system (CT, MR, CR, US, Angio room, Mammography)
| 26 | Acceptance and archiving unlimited number generated procedures (studies) per year in accordance with archive size
| 27 | The system supports DICOM standards (DMWL, MPPS, Query / Retrieve)
| 28 | The choice of image compression for individual modalities (lossy or lossless)
| 29 | Automatic backup of database
| 30 | Web based system administration
| 31 | Scalability (ability to increase archive capacity)
| 32 | Desk-top and Web client
| 33 | Modularity (possibility of adding new application functionality - modules)
| 34 | Support for two language for user interface English or Serbian
| 35 | System support access to other pacs archives and can pull images (dicom Query/Retrieve)
| 36 | System can view images from any dicom modality (CT, RX, MR, MG, US, XA, CR...)
| 37 | STANDARD TOOLS AVAILABLE ON ALL WORKPLACES – UNLIMITED USERS
| 38 | A comparison of two or more series of images from the same or different studies of the same patient, the possibility of automatic comparison of older and current picture of the patient
| 39 | Search by multiple criteria: patient name, patient ID, accession number, modality, and so on.
| 40 | Creating work lists for easy search
| 41 | Variable user interface (graphic) for each user
| 42 | Worklist functionality
| 43 | Variable user interface (graphic) for individual users
| 44 | A graphical editor (menu) that allows each user to define and customize all the tools to ownway and according to their needs
| 45 | The possibility of "recording the current screen" for the studies, so that the radiologist can continue later diagnostic with the same look of the screen that was before ("in the same place where he left")
| 46 | CINE- mode – automatically image changing with number of frames per second
| 47 | CINE-mode, Image rotation (90 °, -90 °), Image flip (vertical, horizontal)
| 48 | Image flip (vertical, horizontal)
| 49 | Contrast/Brightness change by moving of mouse
| 50 | Automatic selection of predefined contrast "window level" for the lung, head, liver, bone, soft tissue
| 51 | „Hounsfield" value (next to the cursor)
| 52 | Zoom, Magnifying glass, Ruler, Spine Labelling, Angle measurement, Invert (negative/positive)
| 53 | Magnifying glass for the specific part of image
| 54 | (Ruler)
| 55 | „Spine Labeling“
| 56 | Angle measurement
| 57 | Invert (negative/positive)
| 58 | Annotations and arrows to highlight the details
| 59 | The ability to mark key images
| 60 | Possibility to reset the zoom, contrast, orientation, measurements and annotations
| 61 | Print selected images on the film, along with all tags added on any medical printer of dry technology
| 62 | Choosing the print format (size film, portrait / landscape, number and distribution of images on a single film)
| 63 | Export of images in „True Size“
| 64 | Export images in JPG format
| 65 | Archiving on CD/DVD with Viewer program
| 66 | ADVANCED TOOLS AVAILABLE ON ALL WORKPLACES WITH CONCURRENT ACCESS
| 67 | Multiplanar reconstruction (MPR) with following functionalities – 3 USERS
| 68 | Adjusting the thickness and spacing of the study
| 69 | Magnification, moving (Pan) Change the contrast (Window Level)
| 70 | The default contrast for the lung, soft tissue, bone, head, liver
| 71 | Possibility of making curved cross-section (Oblique MPR)
| 72 | The possibility of making a double curved section (Double-oblique MPR)
| 73 | Curved reformation (CPR)
| 74 | MIP, MinIP, AvgIP,
| 75 | Printing MPR reformatted images
| 76 | Creating annotations and regions of interest
| 77 | Reformation with a new number and thickness of slices
| 78 | Export of processed images as a new DICOM series
| 79 | 3D RECONSTRUCTIONS FOR 3 CONCURRENT USERS:
| 80 | Automated segmentation
| 81 | MIP processing
| 82 | /
| 83 | Removing the patient table with touch of a button
| 84 | Remove bones with touch of a button
| 85 | Removing fragments
| 86 | Volume measurement
| 87 | Distance measurement in 3 projections
| 88 | View in multi window
| 89 | Export 3D images for archiving
| 90 | FUNCTIONALITY OF BLOOD VESSEL ANALYSIS FOR 3 CONCURRENT USERS:
| 91 | Calculation of stenosis
| 92 | Cross section measurement
| 93 | Measurement of the maximum and minimum diameters
| 94 | Automated tracking the centerline of the vessel
| 95 | 3D view of the chosen vessel
| 96 | Manual changes the contour of the vessel
| 97 | The graph display of vessel diameter in length and the straitened blood vessel
| 98 | Zero footprint viewer available on all workplaces (tablet, pc, mobile phone) – Unlimited users
| 99 | Access via user name and password
| 100 | Unlimited access within the hospital network
| 101 | Unlimited access outside the hospital network through a VPN, static IP addresses, or other ways
| 102 | Zero footprint viewer (standard tools)for Unlimited users
| 103 | Search by multiple criteria: patient name, patient ID, date
| 104 | Window Level,
| 105 | Zoom,
| 106 | Flip,
| 107 | Rotate
| 108 | Annotations in the image
| 109 | Ruler
| 110 | Arrow-indicator
| 111 | Cine mod
| 112 | Tools for manipulating CT studies in zero footprint viewer – Unlimited users
| 113 | MPR, MPI, MINIP, AVGIP
| 114 | /
| 115 | 3D reconstruction
| 116 | Window Level,
| 117 | The default contrast of the lung, soft tissue, bone, head, liver
| 118 | Magnify,
| 119 | Measurement of distances on obtained cross sections
| 120 | Surface measurement,
| 121 | ROI – region of interest
| 122 | Inversion
| 123 | Housfield unit
| 124 | /
| 125 | /
| 126 | Exporting images
| 127 | RADIOLOGY INFORMATION SYSTEM – 10 CONCURRENT USERS
| 128 | User interface in Serbian language;
| 129 | The application supports the business management function of connecting radiological equipment DMWL (DICOM Modality Worklist);
| 130 | The application supports the business management function of worklist radiological modalities; support in radiographer’s work;
| 131 | The application supports the business management function of changing patient status; integration with PACS system;
| 132 | The application supports the business reporting function by number of completed studies at daily, monthly and annual levels;
| 133 | The application supports the business reporting function by radiologist’s comparison according to the number of studies made at daily, monthly and annual levels;
| 134 | The application supports a business management function of automatic opening radiologic images on the diagnostic radiologist station. When selecting a patient from the worklist, radiological images and reports shows in the PACS application;
| 135 | The application supports the automatic creation of basic medical documentation based on the law of the Republic of Serbia;
| 136 | The application supports the automatic creation and sending of predefined reports (electronic format) for the purposes of health statistics;
| 137 | The application supports the automatic creation and management of an electronic invoice (account);
| 138 | The application supports central automatic backup of non DICOM data;
| 139 | The application supports the use official nomenclatures and classifications in the Republic of Serbia (MKB 10, JKL, municipalities, settlements, professions, specialization, medical equipment...);
| 140 | Records of consumable material and automatic connection with procedures;
| 141 | Integration with existing BIS (Hospital Information System) system (communication - exchange of HL7 messages);
| 142 | Electronic signing of the consent by the patient;
| 143 | Management of additional medical data and questionnaires;
| 144 | Functionality for entering laboratory findings of interest (urea, creatinine, etc );
| 145 | Functionality for connecting with LDAP systems with user names and passwords;
| 146 | Functionality for writing all actions and changes within the system, at least those related to system users and their actions;
| 147 | Functionality for entering a template with the functionality for defining and dividing patterns among users;
| 148 | Functionality for reports archive for all patients;
| 149 | Input of glomerular filtration rate or determination of eGFR based on the level of kretinin;
| 150 | Modular design, any user in the system depending on the duty has an appropriate module available (eg reception, radiology engineer, specialist radiologist, - ultrasound diagnostic, etc.). The system allows several modules to be opened simultaneously;
| 151 | View of demographics data of patient;
| 152 | Print functionality of anonymized findings to protect the patient's privacy;
| 153 | Functionality of creating individual forms of reports for different procedures and users
| 154 | Functionality of signing documents by entering scanned signatures;
| 155 | Authorization functionality via PIN;
| 156 | Input, display, referral diagnosis according to MKD10 catalogue;
| 157 | Arrangement of examinations for diagnostic units; (CT worklist contain only CT patients, ie precisely defined modality );
| 158 | Functionality of defining certain types of examinations in certain days, determined by days / weeks / months;
| 159 | Automatically deploy search to the first possible free term, separate work from non-working days, scheduling function depending on the degree of exam urgency;
| 160 | Functionality for direct input of urgent examinations in schedule;
| 161 | Functionality for defining the duration of the exam for each study;
| 162 | Functionality for making personal schedules and worksheets (for radiologist’s engineers);
| 163 | Functionality for scheduling by workplaces;
| 164 | Functionality for scheduling by diagnostic disciplines and worksheets by diagnostic disciplines;
| 165 | Presentation of the associated data (images, PDF, scanned documents) and attaching to radiological card of the patient;
| 166 | Supported system for dictation of reports and audio file storage;
| 167 | Functionality of management of predefined radiological reports by user, diagnostics or procedure;
| 168 | Possibility of radiological consultation via the Web zero footprint browser or via email;
| 169 | Functionality of inserting a radiological image in a radiological report.
4.7 |  | Infusion pump, dual channel, compatible with magnetic resonance |  | 1
| 1 | Compact peristaltic linear pump designed for use in the MRI environment (minimum 10,000 Gauss) and transport of patients
| 2 | For use in the MRI environment from 0.2 to 3T
| 3 | For use in a magnetic field, with or without the use of an additional shield
| 4 | The infusion pump has to have two channels as well as the possibility of using both channels and the possibility of functioning as two volumetric, two syringes or combined syringe volumetric depending, depending on needs of the user
| 5 | /
| 6 | The infusion pump must be delivered with a mobile stand which is made of non-magnetic material
| 7 | The infusion pump must have a built-in battery capacity of more than 4 hours
| 8 | Flow range 0.1-1200 ml / hr I to 0-100 ml / hr at intervals of 0.1 ml / hr, 100-1200 ml / hr at intervals of 1 ml / hr
| 9 | Accuracy - up to 5% (for flow 0.1 to 0.9 within up to 10%)
| 10 | Primary VTBI range - 0.1 to 99.9, 100 to 999 ml
| 11 | Secondary VTBI range 0.1 to 99.9, 100 to 999 ml
| 12 | The total volume volume (VI) of 0.1 to 99.9, 100 to 999 ml
| 13 | KVO range adjustable 1 to 5 ml / hr, or adjusted speed (lower value)
14
| 15 | Patient line (downstream) Return pressure range +300 to -100 mmHg
| 16 | Occlusion proximal detection range - 1 to 10 PSI (6.9 to 68.8 kPa), adjustable by the user
| 17 | Occlusion detection mechanism - two independent sensors for force
| 18 | The range of occlusion measurement is 1 to 10 PSI (6.9 to 68.8 kPa) accuracy <2 PSI (13.8 kPa), or 10% of set values, (higher value)
| 19 | Occlusion detection time, less than 30 seconds
| 20 | Bolus range 0.1 to 99.9, 100 to 1400 ml / hr at a distance of 1 ml / hr
| 21 | Air Detection Method in Line - Ultrasonic Bubble Detector
| 22 | Automatic free-flow protection for I.V. line
| 23 | The LCD display of the device has to show the following: Speed, VTBI, VI for each channel, Battery charge capacity, Indication that the device is plugged in to the network, an indicator that is channel acoustic, pulse, saturation, perfusion index, indicator whether the infusion is in stream or stopped
| 24 | Soft keys for programming different functions as well as numeric keypad for quick and easy changing of numerical values on the front panel